## Summary
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
230427_23-097_23-074_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -11.1752 | -83.4786 |  |  |  | 0.193 | 2.3858
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25939 | 3.729
3 | * | UT2 | -15.84 | -119 |  |  |  | 4.3532 | 33.1307
4 |  | EV | -11.2103 | -82.5501 |  |  |  | 0.10867 | 1.6505
5 |  | A00000322632 | -8.6928 | -54.5176 |  |  |  | 0.067823 | 1.3065
6 |  | A00000322644 | -8.8122 | -54.5801 |  |  |  | 0.028414 | 0.14737
7 |  | A00000322660 | -8.4114 | -52.1194 |  |  |  | 0.053009 | 0.38179
8 |  | A00000322734 | -8.4237 | -51.1711 |  |  |  | 0.024913 | 0.14703
9 |  | A00000322666 | -8.6358 | -52.8091 |  |  |  | 0.020412 | 0.15632
10 |  | A00000322667 | -8.0703 | -48.5709 |  |  |  | 0.061809 | 0.50714
11 |  | A00000322665 | -8.7634 | -53.7146 |  |  |  | 0.030049 | 0.052696
12 |  | A00000322668 | -8.9189 | -55.2545 |  |  |  | 0.038232 | 0.12976
13 |  | EV | -11.1823 | -79.5383 |  |  |  | 0.074106 | 1.1096
14 |  | A00000322423 | -6.0835 | -43.9786 |  |  |  | 0.17439 | 1.9376
15 |  | A00000322427 | -9.6366 | -65.4461 |  |  |  | 0.080872 | 1.0605
16 |  | A00000170752 | -16.4443 | -132.2894 |  |  |  | 0.17518 | 3.2295
17 |  | A00000170760 | -4.9647 | -30.1511 |  |  |  | 0.28734 | 4.8194
18 |  | A00000322421 | -7.6965 | -49.0601 |  |  |  | 0.065168 | 1.0064
19 |  | A00000170829 | -9.8501 | -60.4371 |  |  |  | 0.073096 | 0.65647
20 |  | A00000170754 | -17.5185 | -132.8334 |  |  |  | 0.2008 | 3.6009
21 |  | A00000170758 | -6.3437 | -41.1487 |  |  |  | 0.27578 | 4.3514
22 |  | EV | -11.134 | -79.0863 |  |  |  | 0.11449 | 1.9194
23 |  | A00000322425 | -7.4288 | -43.768 |  |  |  | 0.094435 | 1.7802
24 |  | A00000170750 | -6.5707 | -27.7686 |  |  |  | 0.037385 | 0.798
25 |  | A00000170756 | -10.1985 | -69.0825 |  |  |  | 0.066139 | 1.8205
26 |  | A00000322705 | -8.3737 | -53.326 |  |  |  | 0.032755 | 0.74118
27 |  | A00000322627 | -8.3857 | -52.8141 |  |  |  | 0.0093452 | 0.077672
28 |  | A00000322639 | -8.3994 | -52.8353 |  |  |  | 0.040145 | 0.13911
29 |  | A00000322654 | -8.2874 | -52.636 |  |  |  | 0.0060553 | 0.078999
30 |  | A00000322729 | -8.2297 | -52.2568 |  |  |  | 0.041716 | 0.34678
31 |  | EV | -11.1214 | -79.0754 |  |  |  | 0.089454 | 1.2656
32 |  | A00000322757 | -8.3692 | -53.6418 |  |  |  | 0.064038 | 1.2288
33 |  | A00000322747 | -8.3868 | -52.9462 |  |  |  | 0.018184 | 0.097367
34 |  | A00000322748 | -7.9656 | -49.7578 |  |  |  | 0.048093 | 0.34848
35 |  | A00000322749 | -8.3477 | -52.1524 |  |  |  | 0.052058 | 0.28595
36 |  | A00000322746 | -8.5122 | -53.7208 |  |  |  | 0.03249 | 0.21055
37 |  | A00000322687 | -8.3215 | -50.6813 |  |  |  | 0.019536 | 0.18626
38 |  | A00000322622 | -8.5392 | -51.8493 |  |  |  | 0.03934 | 0.25403
39 |  | WOOD4-D-5 | 10.5495 | -56.2794 | -9.5995 | 351.7602 | 133.8074 | 0.084093 | 1.4706 |  | 2.7459
40 |  | EV | -10.9637 | -80.4713 |  |  |  | 0.13234 | 0.07816

## Detail
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.435 | -75.152 | -11.7467 | -90.9248 | -1.404 | 78.99 | 2023/04/27 11:36:38
1 |  | EV | -9.045 | -70.259 | -11.3112 | -85.1536 | -1.085 | 52.747 | 2023/04/27 11:45:34
1 |  | EV | -8.902 | -68.826 | -11.1516 | -83.4634 | -1.096 | 51.443 | 2023/04/27 11:54:31
1 |  | EV | -8.865 | -68.381 | -11.1102 | -82.9385 | -1.003 | 51.412 | 2023/04/27 12:03:26
1 |  | EV | -8.872 | -68.147 | -11.1181 | -82.6625 | -1.083 | 50.174 | 2023/04/27 12:12:23
1 |  | EV | -8.853 | -67.671 | -11.0968 | -82.101 | -1.111 | 47.591 | 2023/04/27 12:21:19
1 |  | EV | -8.821 | -67.553 | -11.0611 | -81.9619 | -1.223 | 47.434 | 2023/04/27 12:30:15
1 |  | EV | -8.802 | -67.485 | -11.0399 | -81.8817 | -1.254 | 46.755 | 2023/04/27 12:39:11
1 |  | EV | -8.815 | -67.428 | -11.0544 | -81.8144 | -1.119 | 46.355 | 2023/04/27 12:48:07
1 |  | EV | -8.822 | -67.487 | -11.0622 | -81.884 | -0.851 | 45.77 | 2023/04/27 12:57:04
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -11.1752 | -83.4786 |  |  |  | 0.193 | 2.3858
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.13 | 7.531 | 1.1669 | 6.5987 | -1.129 | 43.879 | 2023/04/27 13:06:53
2 |  | PZ | 2.667 | 15.006 | 1.7665 | 15.4154 | -1.22 | 44.686 | 2023/04/27 13:15:49
2 |  | PZ | 2.79 | 16.88 | 1.9039 | 17.6257 | -1.22 | 44.58 | 2023/04/27 13:24:45
2 |  | PZ | 2.854 | 18.014 | 1.9753 | 18.9633 | -1.207 | 44.506 | 2023/04/27 13:33:42
2 |  | PZ | 2.91 | 18.437 | 2.0379 | 19.4622 | -1.221 | 44.619 | 2023/04/27 13:42:38
2 |  | PZ | 2.953 | 18.769 | 2.0859 | 19.8538 | -1.116 | 44.432 | 2023/04/27 13:51:36
2 |  | PZ | 2.943 | 19.42 | 2.0747 | 20.6216 | -1.179 | 43.847 | 2023/04/27 14:00:31
2 |  | PZ | 2.923 | 19.303 | 2.0524 | 20.4836 | -1.043 | 44.006 | 2023/04/27 14:09:28
2 |  | PZ | 2.968 | 19.749 | 2.1026 | 21.0097 | -0.938 | 43.963 | 2023/04/27 14:18:25
2 |  | PZ | 2.996 | 19.712 | 2.1339 | 20.966 | -1.018 | 43.993 | 2023/04/27 14:27:22
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25939 | 3.729
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 |  |  | -1.2115 | -2.284 |  |  | 2023/04/27 14:36:47
3 |  | UT2 | -13.658 | -96.212 | -16.4622 | -115.7648 | -0.944 | 44.551 | 2023/04/27 14:46:46
3 |  | UT2 | -14.294 | -106.125 | -17.1723 | -127.457 | -1.055 | 44.184 | 2023/04/27 14:55:43
3 |  | UT2 | -14.392 | -108.58 | -17.2817 | -130.3527 | -0.815 | 43.91 | 2023/04/27 15:04:41
3 |  | UT2 | -14.493 | -109.969 | -17.3945 | -131.991 | -0.934 | 43.773 | 2023/04/27 15:13:37
3 |  | UT2 | -14.449 | -110.298 | -17.3454 | -132.379 | -0.981 | 43.352 | 2023/04/27 15:22:35
3 |  | UT2 | -14.589 | -111.161 | -17.5017 | -133.3969 | -0.977 | 43.986 | 2023/04/27 15:31:31
3 |  | UT2 | -14.563 | -111.23 | -17.4727 | -133.4783 | -0.809 | 43.772 | 2023/04/27 15:40:27
3 |  | UT2 | -14.521 | -111.477 | -17.4258 | -133.7696 | -0.927 | 43.288 | 2023/04/27 15:49:24
3 |  | UT2 | -14.544 | -111.474 | -17.4515 | -133.7661 | -0.976 | 43.252 | 2023/04/27 15:58:20
3 |  | UT2 | -14.606 | -111.978 | -17.5207 | -134.3606 | -0.958 | 43.454 | 2023/04/27 16:07:16
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 | * | UT2 | -15.84 | -119 |  |  |  | 4.3532 | 33.1307
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.247 | -72.426 | -11.5368 | -87.7095 | -0.97 | 43.298 | 2023/04/27 16:17:05
4 |  | EV | -9 | -69.15 | -11.261 | -83.8455 | -0.97 | 43.394 | 2023/04/27 16:26:02
4 |  | EV | -8.869 | -67.662 | -11.1147 | -82.0904 | -0.872 | 43.212 | 2023/04/27 16:34:59
4 |  | EV | -8.934 | -67.707 | -11.1873 | -82.1435 | -0.87 | 43.13 | 2023/04/27 16:43:56
4 |  | EV | -8.89 | -67.432 | -11.1382 | -81.8191 | -0.888 | 43.153 | 2023/04/27 16:52:52
4 |  | EV | -8.932 | -67.399 | -11.185 | -81.7802 | -1.01 | 43.081 | 2023/04/27 17:01:48
4 |  | EV | -8.904 | -67.133 | -11.1538 | -81.4665 | -0.878 | 42.588 | 2023/04/27 17:10:44
4 |  | EV | -8.946 | -67.447 | -11.2007 | -81.8368 | -0.956 | 42.59 | 2023/04/27 17:19:40
4 |  | EV | -8.909 | -67.086 | -11.1594 | -81.411 | -1.051 | 42.658 | 2023/04/27 17:28:36
4 |  | EV | -8.915 | -67.075 | -11.1661 | -81.3981 | -1.06 | 42.695 | 2023/04/27 17:37:32
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -11.2103 | -82.5501 |  |  |  | 0.10867 | 1.6505
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | A00000322632 | -6.794 | -46.184 | -8.7977 | -56.7574 | -0.98 | 43.055 | 2023/04/27 17:47:22
5 |  | A00000322632 | -6.688 | -44.054 | -8.6794 | -54.2451 | -0.978 | 42.686 | 2023/04/27 17:56:18
5 |  | A00000322632 | -6.686 | -43.638 | -8.6771 | -53.7544 | -1.052 | 43.216 | 2023/04/27 18:05:15
5 |  | A00000322632 | -6.632 | -43.264 | -8.6168 | -53.3133 | -0.998 | 42.918 | 2023/04/27 18:14:11
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | A00000322632 | -8.6928 | -54.5176 |  |  |  | 0.067823 | 1.3065
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | A00000322644 | -6.811 | -44.26 | -8.8167 | -54.4881 | -0.865 | 43.036 | 2023/04/27 18:24:02
6 |  | A00000322644 | -6.821 | -44.554 | -8.8279 | -54.8349 | -0.807 | 42.705 | 2023/04/27 18:32:59
6 |  | A00000322644 | -6.766 | -44.307 | -8.7665 | -54.5435 | -1.133 | 43.771 | 2023/04/27 18:41:55
6 |  | A00000322644 | -6.83 | -44.231 | -8.8379 | -54.4539 | -1.038 | 43.256 | 2023/04/27 18:50:51
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | A00000322644 | -8.8122 | -54.5801 |  |  |  | 0.028414 | 0.14737
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | A00000322660 | -6.515 | -42.744 | -8.4862 | -52.7 | -1.062 | 43.365 | 2023/04/27 19:00:41
7 |  | A00000322660 | -6.466 | -42.36 | -8.4315 | -52.2471 | -1.126 | 43.597 | 2023/04/27 19:09:37
7 |  | A00000322660 | -6.404 | -41.913 | -8.3623 | -51.7198 | -1.027 | 42.983 | 2023/04/27 19:18:33
7 |  | A00000322660 | -6.407 | -41.99 | -8.3656 | -51.8107 | -0.995 | 42.993 | 2023/04/27 19:27:29
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | A00000322660 | -8.4114 | -52.1194 |  |  |  | 0.053009 | 0.38179
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | A00000322734 | -6.435 | -41.317 | -8.3969 | -51.0169 | -1.095 | 42.451 | 2023/04/27 19:37:17
8 |  | A00000322734 | -6.494 | -41.635 | -8.4628 | -51.3919 | -0.969 | 42.282 | 2023/04/27 19:46:14
8 |  | A00000322734 | -6.453 | -41.345 | -8.417 | -51.0499 | -0.873 | 42.533 | 2023/04/27 19:55:11
8 |  | A00000322734 | -6.454 | -41.494 | -8.4181 | -51.2256 | -1.135 | 42.856 | 2023/04/27 20:04:07
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | A00000322734 | -8.4237 | -51.1711 |  |  |  | 0.024913 | 0.14703
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | A00000322666 | -6.676 | -43.023 | -8.666 | -53.0291 | -0.792 | 43.256 | 2023/04/27 20:13:55
9 |  | A00000322666 | -6.63 | -42.817 | -8.6146 | -52.7861 | -0.941 | 42.721 | 2023/04/27 20:22:51
9 |  | A00000322666 | -6.653 | -42.643 | -8.6403 | -52.5809 | -1.086 | 42.768 | 2023/04/27 20:31:47
9 |  | A00000322666 | -6.637 | -42.863 | -8.6224 | -52.8403 | -1.124 | 42.662 | 2023/04/27 20:40:43
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | A00000322666 | -8.6358 | -52.8091 |  |  |  | 0.020412 | 0.15632
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | A00000322667 | -6.208 | -39.934 | -8.1434 | -49.3856 | -1.008 | 43.07 | 2023/04/27 20:50:31
10 |  | A00000322667 | -6.182 | -39.303 | -8.1144 | -48.6414 | -0.886 | 42.481 | 2023/04/27 20:59:26
10 |  | A00000322667 | -6.097 | -38.8 | -8.0195 | -48.0481 | -0.769 | 42.371 | 2023/04/27 21:08:23
10 |  | A00000322667 | -6.083 | -38.936 | -8.0038 | -48.2085 | -1.119 | 42.563 | 2023/04/27 21:17:20
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | A00000322667 | -8.0703 | -48.5709 |  |  |  | 0.061809 | 0.50714
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | A00000322665 | -6.805 | -43.551 | -8.81 | -53.6518 | -0.993 | 43.425 | 2023/04/27 21:27:09
11 |  | A00000322665 | -6.753 | -43.592 | -8.752 | -53.7002 | -0.97 | 42.537 | 2023/04/27 21:36:06
11 |  | A00000322665 | -6.734 | -43.597 | -8.7307 | -53.7061 | -0.97 | 42.731 | 2023/04/27 21:45:02
11 |  | A00000322665 | -6.761 | -43.677 | -8.7609 | -53.8004 | -0.876 | 42.571 | 2023/04/27 21:53:59
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | A00000322665 | -8.7634 | -53.7146 |  |  |  | 0.030049 | 0.052696
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | A00000322668 | -6.887 | -44.734 | -8.9016 | -55.0472 | -0.936 | 43.176 | 2023/04/27 22:03:49
12 |  | A00000322668 | -6.871 | -44.909 | -8.8837 | -55.2536 | -0.95 | 42.608 | 2023/04/27 22:12:45
12 |  | A00000322668 | -6.894 | -44.953 | -8.9094 | -55.3055 | -0.883 | 42.478 | 2023/04/27 22:21:42
12 |  | A00000322668 | -6.958 | -45.043 | -8.9809 | -55.4116 | -1.052 | 42.608 | 2023/04/27 22:30:38
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | A00000322668 | -8.9189 | -55.2545 |  |  |  | 0.038232 | 0.12976
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.821 | -63.839 | -11.0611 | -77.5813 | -1.057 | 42.906 | 2023/04/27 22:40:28
13 |  | EV | -8.979 | -65.983 | -11.2375 | -80.1101 | -1.048 | 42.915 | 2023/04/27 22:49:24
13 |  | EV | -8.943 | -65.997 | -11.1973 | -80.1266 | -1.017 | 42.619 | 2023/04/27 22:58:20
13 |  | EV | -8.975 | -66.174 | -11.2331 | -80.3354 | -0.905 | 42.578 | 2023/04/27 23:07:17
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -11.1823 | -79.5383 |  |  |  | 0.074106 | 1.1096
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | A00000322423 | -4.618 | -38.126 | -6.368 | -47.2531 | -0.995 | 42.876 | 2023/04/27 23:17:07
14 |  | A00000322423 | -4.327 | -35.175 | -6.0431 | -43.7724 | -1.053 | 42.444 | 2023/04/27 23:26:03
14 |  | A00000322423 | -4.278 | -34.314 | -5.9883 | -42.7569 | -0.914 | 42.361 | 2023/04/27 23:35:00
14 |  | A00000322423 | -4.23 | -33.784 | -5.9347 | -42.1318 | -1.015 | 42.426 | 2023/04/27 23:43:57
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | A00000322423 | -6.0835 | -43.9786 |  |  |  | 0.17439 | 1.9376
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | A00000322427 | -7.433 | -51.963 | -9.5112 | -63.5737 | -1.098 | 42.93 | 2023/04/27 23:53:44
15 |  | A00000322427 | -7.619 | -54.002 | -9.7189 | -65.9786 | -1.015 | 42.629 | 2023/04/28 00:02:39
15 |  | A00000322427 | -7.544 | -54.07 | -9.6352 | -66.0588 | -1.055 | 42.646 | 2023/04/28 00:11:36
15 |  | A00000322427 | -7.585 | -54.167 | -9.681 | -66.1733 | -1.001 | 42.494 | 2023/04/28 00:20:31
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | A00000322427 | -9.6366 | -65.4461 |  |  |  | 0.080872 | 1.0605
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | A00000170752 | -13.382 | -105.572 | -16.154 | -126.8048 | -0.968 | 42.866 | 2023/04/28 00:30:19
16 |  | A00000170752 | -13.694 | -110.597 | -16.5024 | -132.7317 | -0.987 | 42.336 | 2023/04/28 00:39:15
16 |  | A00000170752 | -13.738 | -111.919 | -16.5515 | -134.291 | -1.02 | 42.407 | 2023/04/28 00:48:11
16 |  | A00000170752 | -13.754 | -112.8 | -16.5694 | -135.3301 | -0.899 | 42.322 | 2023/04/28 00:57:08
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | A00000170752 | -16.4443 | -132.2894 |  |  |  | 0.17518 | 3.2295
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | A00000170760 | -3.782 | -30.628 | -5.4345 | -38.4093 | -1.045 | 43.036 | 2023/04/28 01:06:57
17 |  | A00000170760 | -3.301 | -22.916 | -4.8974 | -29.3131 | -1.14 | 42.475 | 2023/04/28 01:15:53
17 |  | A00000170760 | -3.213 | -20.909 | -4.7991 | -26.9459 | -1.158 | 42.36 | 2023/04/28 01:24:48
17 |  | A00000170760 | -3.149 | -20.053 | -4.7277 | -25.9363 | -1.114 | 42.117 | 2023/04/28 01:33:43
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | A00000170760 | -4.9647 | -30.1511 |  |  |  | 0.28734 | 4.8194
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | A00000322421 | -5.712 | -38.15 | -7.5896 | -47.2814 | -1.17 | 42.878 | 2023/04/28 01:43:32
18 |  | A00000322421 | -5.833 | -40.129 | -7.7247 | -49.6156 | -1.08 | 42.506 | 2023/04/28 01:52:27
18 |  | A00000322421 | -5.828 | -40.126 | -7.7191 | -49.6121 | -0.912 | 42.488 | 2023/04/28 02:01:24
18 |  | A00000322421 | -5.858 | -40.227 | -7.7526 | -49.7312 | -0.965 | 42.502 | 2023/04/28 02:10:20
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | A00000322421 | -7.6965 | -49.0601 |  |  |  | 0.065168 | 1.0064
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | A00000170829 | -7.643 | -48.355 | -9.7457 | -59.3181 | -1.01 | 42.801 | 2023/04/28 02:20:08
19 |  | A00000170829 | -7.762 | -49.481 | -9.8786 | -60.6462 | -0.995 | 42.274 | 2023/04/28 02:29:05
19 |  | A00000170829 | -7.724 | -49.511 | -9.8362 | -60.6816 | -0.963 | 42.366 | 2023/04/28 02:38:01
19 |  | A00000170829 | -7.817 | -49.868 | -9.94 | -61.1026 | -0.915 | 42.304 | 2023/04/28 02:46:56
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | A00000170829 | -9.8501 | -60.4371 |  |  |  | 0.073096 | 0.65647
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | A00000170754 | -14.314 | -105.474 | -17.1947 | -126.6892 | -1.071 | 42.629 | 2023/04/28 02:56:45
20 |  | A00000170754 | -14.624 | -111.152 | -17.5408 | -133.3863 | -1.042 | 42.059 | 2023/04/28 03:05:39
20 |  | A00000170754 | -14.737 | -112.662 | -17.667 | -135.1673 | -0.883 | 42.051 | 2023/04/28 03:14:36
20 |  | A00000170754 | -14.741 | -113.445 | -17.6714 | -136.0909 | -1.018 | 42.284 | 2023/04/28 03:23:31
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | A00000170754 | -17.5185 | -132.8334 |  |  |  | 0.2008 | 3.6009
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | A00000170758 | -5.007 | -39.298 | -6.8023 | -48.6355 | -1.004 | 42.78 | 2023/04/28 03:33:20
21 |  | A00000170758 | -4.505 | -32.211 | -6.2418 | -40.2765 | -1.081 | 42.388 | 2023/04/28 03:42:16
21 |  | A00000170758 | -4.443 | -30.494 | -6.1726 | -38.2513 | -1.001 | 42.347 | 2023/04/28 03:51:13
21 |  | A00000170758 | -4.43 | -29.799 | -6.1581 | -37.4315 | -0.84 | 42.277 | 2023/04/28 04:00:09
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | A00000170758 | -6.3437 | -41.1487 |  |  |  | 0.27578 | 4.3514
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.717 | -62.306 | -10.945 | -75.7731 | -0.963 | 42.957 | 2023/04/28 04:09:54
22 |  | EV | -8.917 | -65.497 | -11.1683 | -79.5368 | -1.074 | 42.36 | 2023/04/28 04:18:51
22 |  | EV | -8.947 | -66.14 | -11.2018 | -80.2953 | -1.047 | 42.402 | 2023/04/28 04:27:47
22 |  | EV | -8.964 | -66.517 | -11.2208 | -80.7399 | -1.183 | 42.442 | 2023/04/28 04:36:43
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -11.134 | -79.0863 |  |  |  | 0.11449 | 1.9194
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | A00000322425 | -5.704 | -37.664 | -7.5806 | -46.7082 | -0.974 | 42.817 | 2023/04/28 04:46:29
23 |  | A00000322425 | -5.56 | -35.224 | -7.4198 | -43.8302 | -1.168 | 42.387 | 2023/04/28 04:55:25
23 |  | A00000322425 | -5.505 | -34.058 | -7.3584 | -42.455 | -0.965 | 42.266 | 2023/04/28 05:04:21
23 |  | A00000322425 | -5.503 | -33.739 | -7.3562 | -42.0787 | -1.047 | 42.281 | 2023/04/28 05:13:17
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | A00000322425 | -7.4288 | -43.768 |  |  |  | 0.094435 | 1.7802
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | A00000170750 | -4.852 | -22.734 | -6.6293 | -29.0985 | -1.056 | 42.876 | 2023/04/28 05:23:04
24 |  | A00000170750 | -4.784 | -21.601 | -6.5533 | -27.7621 | -0.975 | 42.525 | 2023/04/28 05:32:00
24 |  | A00000170750 | -4.765 | -21.115 | -6.5321 | -27.1889 | -1.282 | 42.589 | 2023/04/28 05:40:56
24 |  | A00000170750 | -4.797 | -20.976 | -6.5679 | -27.0249 | -0.82 | 42.462 | 2023/04/28 05:49:51
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | A00000170750 | -6.5707 | -27.7686 |  |  |  | 0.037385 | 0.798
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | A00000170756 | -7.961 | -54.032 | -10.1008 | -66.014 | -1.384 | 43.335 | 2023/04/28 05:59:38
25 |  | A00000170756 | -8.049 | -56.718 | -10.1991 | -69.1821 | -1.276 | 42.51 | 2023/04/28 06:08:34
25 |  | A00000170756 | -8.121 | -57.895 | -10.2795 | -70.5704 | -1.14 | 42.69 | 2023/04/28 06:17:29
25 |  | A00000170756 | -8.063 | -57.889 | -10.2147 | -70.5633 | -1.256 | 42.502 | 2023/04/28 06:26:25
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | A00000170756 | -10.1985 | -69.0825 |  |  |  | 0.066139 | 1.8205
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | A00000322705 | -6.459 | -44.371 | -8.4237 | -54.619 | -1.186 | 43.075 | 2023/04/28 06:36:12
26 |  | A00000322705 | -6.415 | -43.079 | -8.3745 | -53.0951 | -1.043 | 42.376 | 2023/04/28 06:45:08
26 |  | A00000322705 | -6.401 | -42.86 | -8.3589 | -52.8368 | -1.026 | 42.579 | 2023/04/28 06:54:03
26 |  | A00000322705 | -6.382 | -42.789 | -8.3377 | -52.7531 | -1.027 | 42.514 | 2023/04/28 07:02:59
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | A00000322705 | -8.3737 | -53.326 |  |  |  | 0.032755 | 0.74118
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | A00000322627 | -6.42 | -42.75 | -8.3801 | -52.7071 | -0.879 | 42.711 | 2023/04/28 07:12:47
27 |  | A00000322627 | -6.429 | -42.814 | -8.3902 | -52.7825 | -0.916 | 42.216 | 2023/04/28 07:21:43
27 |  | A00000322627 | -6.415 | -42.866 | -8.3745 | -52.8439 | -0.978 | 42.339 | 2023/04/28 07:30:39
27 |  | A00000322627 | -6.436 | -42.933 | -8.398 | -52.9229 | -0.877 | 42.316 | 2023/04/28 07:39:35
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | A00000322627 | -8.3857 | -52.8141 |  |  |  | 0.0093452 | 0.077672
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | A00000322639 | -6.434 | -42.811 | -8.3958 | -52.779 | -1.041 | 42.512 | 2023/04/28 07:49:23
28 |  | A00000322639 | -6.472 | -42.913 | -8.4382 | -52.8993 | -0.886 | 41.867 | 2023/04/28 07:58:18
28 |  | A00000322639 | -6.461 | -43.018 | -8.4259 | -53.0232 | -0.978 | 41.942 | 2023/04/28 08:07:14
28 |  | A00000322639 | -6.382 | -42.693 | -8.3377 | -52.6398 | -1.099 | 42.228 | 2023/04/28 08:16:11
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | A00000322639 | -8.3994 | -52.8353 |  |  |  | 0.040145 | 0.13911
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | A00000322654 | -6.345 | -42.695 | -8.2964 | -52.6422 | -1.164 | 42.87 | 2023/04/28 08:25:55
29 |  | A00000322654 | -6.331 | -42.581 | -8.2807 | -52.5077 | -0.995 | 42.324 | 2023/04/28 08:34:51
29 |  | A00000322654 | -6.334 | -42.769 | -8.2841 | -52.7295 | -1.22 | 42.499 | 2023/04/28 08:43:46
29 |  | A00000322654 | -6.338 | -42.714 | -8.2886 | -52.6646 | -0.96 | 42.392 | 2023/04/28 08:52:42
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | A00000322654 | -8.2874 | -52.636 |  |  |  | 0.0060553 | 0.078999
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | A00000322729 | -6.287 | -42.555 | -8.2316 | -52.4771 | -0.887 | 42.748 | 2023/04/28 09:02:28
30 |  | A00000322729 | -6.343 | -42.739 | -8.2941 | -52.6941 | -0.934 | 42.163 | 2023/04/28 09:11:24
30 |  | A00000322729 | -6.263 | -41.961 | -8.2048 | -51.7764 | -0.81 | 42.385 | 2023/04/28 09:20:20
30 |  | A00000322729 | -6.248 | -42.218 | -8.1881 | -52.0796 | -0.907 | 42.26 | 2023/04/28 09:29:15
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | A00000322729 | -8.2297 | -52.2568 |  |  |  | 0.041716 | 0.34678
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.742 | -63.273 | -10.9729 | -76.9137 | -1.03 | 42.98 | 2023/04/28 09:38:59
31 |  | EV | -8.905 | -65.25 | -11.1549 | -79.2455 | -0.889 | 42.323 | 2023/04/28 09:47:55
31 |  | EV | -8.919 | -65.947 | -11.1705 | -80.0676 | -1.096 | 42.538 | 2023/04/28 09:56:50
31 |  | EV | -8.934 | -65.953 | -11.1873 | -80.0747 | -1.051 | 42.669 | 2023/04/28 10:05:46
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -11.1214 | -79.0754 |  |  |  | 0.089454 | 1.2656
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | A00000322757 | -6.475 | -45.205 | -8.4415 | -55.6027 | -1.294 | 42.539 | 2023/04/28 10:15:32
32 |  | A00000322757 | -6.455 | -43.726 | -8.4192 | -53.8582 | -1.052 | 42.779 | 2023/04/28 10:24:28
32 |  | A00000322757 | -6.363 | -42.739 | -8.3165 | -52.6941 | -1.179 | 42.942 | 2023/04/28 10:33:24
32 |  | A00000322757 | -6.348 | -42.5 | -8.2997 | -52.4122 | -0.929 | 42.593 | 2023/04/28 10:42:21
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | A00000322757 | -8.3692 | -53.6418 |  |  |  | 0.064038 | 1.2288
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | A00000322747 | -6.446 | -43.024 | -8.4092 | -53.0302 | -1.146 | 43.009 | 2023/04/28 10:52:07
33 |  | A00000322747 | -6.426 | -43.019 | -8.3868 | -53.0243 | -1.11 | 42.485 | 2023/04/28 11:01:03
33 |  | A00000322747 | -6.43 | -42.815 | -8.3913 | -52.7837 | -1.116 | 42.454 | 2023/04/28 11:09:59
33 |  | A00000322747 | -6.402 | -42.953 | -8.36 | -52.9465 | -0.964 | 42.388 | 2023/04/28 11:18:55
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | A00000322747 | -8.3868 | -52.9462 |  |  |  | 0.018184 | 0.097367
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | A00000322748 | -6.118 | -40.757 | -8.0429 | -50.3563 | -0.867 | 42.161 | 2023/04/28 11:28:42
34 |  | A00000322748 | -6.044 | -40.156 | -7.9603 | -49.6475 | -1.147 | 42.474 | 2023/04/28 11:37:37
34 |  | A00000322748 | -6.021 | -40.121 | -7.9346 | -49.6062 | -0.702 | 42.35 | 2023/04/28 11:46:33
34 |  | A00000322748 | -6.012 | -39.964 | -7.9245 | -49.421 | -0.943 | 42.199 | 2023/04/28 11:55:29
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | A00000322748 | -7.9656 | -49.7578 |  |  |  | 0.048093 | 0.34848
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | A00000322749 | -6.33 | -42.008 | -8.2796 | -51.8319 | -1.141 | 43.13 | 2023/04/28 12:05:15
35 |  | A00000322749 | -6.457 | -42.65 | -8.4214 | -52.5891 | -0.853 | 42.258 | 2023/04/28 12:14:11
35 |  | A00000322749 | -6.384 | -42.109 | -8.3399 | -51.951 | -1.16 | 42.522 | 2023/04/28 12:23:07
35 |  | A00000322749 | -6.393 | -42.352 | -8.35 | -52.2376 | -1.087 | 42.425 | 2023/04/28 12:32:03
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | A00000322749 | -8.3477 | -52.1524 |  |  |  | 0.052058 | 0.28595
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | A00000322746 | -6.495 | -43.437 | -8.4639 | -53.5174 | -1.155 | 43.233 | 2023/04/28 12:41:46
36 |  | A00000322746 | -6.532 | -43.643 | -8.5052 | -53.7603 | -1.075 | 42.394 | 2023/04/28 12:50:41
36 |  | A00000322746 | -6.559 | -43.464 | -8.5353 | -53.5492 | -0.85 | 42.432 | 2023/04/28 12:59:36
36 |  | A00000322746 | -6.567 | -43.894 | -8.5443 | -54.0564 | -1.047 | 42.468 | 2023/04/28 13:08:32
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | A00000322746 | -8.5122 | -53.7208 |  |  |  | 0.03249 | 0.21055
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | A00000322687 | -6.363 | -41.29 | -8.3165 | -50.985 | -0.97 | 42.689 | 2023/04/28 13:18:17
37 |  | A00000322687 | -6.344 | -40.871 | -8.2953 | -50.4908 | -1.276 | 42.397 | 2023/04/28 13:27:13
37 |  | A00000322687 | -6.372 | -41.044 | -8.3265 | -50.6949 | -1.176 | 42.6 | 2023/04/28 13:36:10
37 |  | A00000322687 | -6.391 | -40.925 | -8.3477 | -50.5545 | -1.07 | 42.281 | 2023/04/28 13:45:05
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | A00000322687 | -8.3215 | -50.6813 |  |  |  | 0.019536 | 0.18626
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | A00000322622 | -6.53 | -41.836 | -8.5029 | -51.629 | -1.041 | 42.891 | 2023/04/28 13:54:50
38 |  | A00000322622 | -6.54 | -41.882 | -8.5141 | -51.6833 | -1.024 | 42.621 | 2023/04/28 14:03:45
38 |  | A00000322622 | -6.562 | -41.98 | -8.5387 | -51.7989 | -1.048 | 42.268 | 2023/04/28 14:12:41
38 |  | A00000322622 | -6.618 | -42.393 | -8.6012 | -52.286 | -0.947 | 42.25 | 2023/04/28 14:21:37
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | A00000322622 | -8.5392 | -51.8493 |  |  |  | 0.03934 | 0.25403
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | WOOD4-D-5 | 10.42 | -43.608 | 10.4236 | -53.7191 | -9.414 | 392.433 | 2023/04/28 14:31:21
39 |  | WOOD4-D-5 | 10.522 | -46.133 | 10.5375 | -56.6973 | -9.584 | 394.005 | 2023/04/28 14:40:17
39 |  | WOOD4-D-5 | 10.614 | -46.672 | 10.6402 | -57.333 | -9.761 | 393.79 | 2023/04/28 14:49:14
39 |  | WOOD4-D-5 | 10.575 | -46.702 | 10.5967 | -57.3684 | -9.639 | 394.438 | 2023/04/28 14:58:10
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | WOOD4-D-5 | 10.5495 | -56.2794 | -9.5995 | 351.7602 | 133.8074 | 0.084093 | 1.4706 |  | 2.7459
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.538 | -66.286 | -10.7451 | -80.4675 | -1.112 | 47.786 | 2023/04/28 15:07:55
40 |  | EV | -8.781 | -66.186 | -11.0164 | -80.3495 | -1.001 | 44.297 | 2023/04/28 15:16:51
40 |  | EV | -8.786 | -66.374 | -11.022 | -80.5713 | -1.156 | 43.76 | 2023/04/28 15:25:47
40 |  | EV | -8.83 | -66.311 | -11.0712 | -80.4969 | -0.886 | 43.189 | 2023/04/28 15:34:41
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.9637 | -80.4713 |  |  |  | 0.13234 | 0.07816

## Source
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-24278 | 2023/04/27 11:36:38 | 1-01 | 1 | -9.435 | -75.152 | 16308 | -1 | 1 | EV | standard | H2O | 1682617304.17 | 0.237 | 0.797 | 94.744 | 0.002 | 0.004 | 2.282 | 78.99 | 38.375 | 2.697 | 77.371 | 0.858 | -1.404 | 0 | 0 | 1 | 40.062 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-24278 | 2023/04/27 11:45:34 | 1-01 | 2 | -9.045 | -70.259 | 12888 | -1 | 1 | EV | standard | H2O | 1682617840.04 | 0.221 | 0.776 | 70.877 | 0.001 | -0.003 | 1.755 | 52.747 | 29.663 | 2.19 | 63.328 | 0.856 | -1.085 | 0 | 0 | 1 | 39.312 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-24278 | 2023/04/27 11:54:31 | 1-01 | 3 | -8.902 | -68.826 | 12751 | -1 | 1 | EV | standard | H2O | 1682618376.79 | 0.206 | 0.734 | 73.561 | 0 | 0.001 | 1.82 | 51.443 | 29.482 | 2.154 | 62.559 | 0.856 | -1.096 | 0 | 0 | 1 | 39.062 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-24278 | 2023/04/27 12:03:26 | 1-01 | 4 | -8.865 | -68.381 | 12751 | 0 | 1 | EV | standard | H2O | 1682618912.04 | 0.194 | 0.731 | 73.023 | 0.002 | -0.003 | 1.805 | 51.412 | 29.018 | 2.16 | 64.449 | 0.856 | -1.003 | 0 | 0 | 1 | 39.125 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-24278 | 2023/04/27 12:12:23 | 1-01 | 5 | -8.872 | -68.147 | 12663 | 0 | 1 | EV | standard | H2O | 1682619447.81 | 0.197 | 0.73 | 71.166 | 0.001 | -0.002 | 1.762 | 50.174 | 28.907 | 2.149 | 63.304 | 0.856 | -1.083 | 0 | 0 | 1 | 39.062 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-24278 | 2023/04/27 12:21:19 | 1-01 | 6 | -8.853 | -67.671 | 12616 | 0 | 1 | EV | standard | H2O | 1682619984.5 | 0.198 | 0.731 | 70.808 | 0.001 | 0 | 1.754 | 47.591 | 29.089 | 2.141 | 62.779 | 0.856 | -1.111 | 0 | 0 | 1 | 39.062 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-24278 | 2023/04/27 12:30:15 | 1-01 | 7 | -8.821 | -67.553 | 12766 | 0 | 1 | EV | standard | H2O | 1682620521.36 | 0.198 | 0.706 | 64.227 | 0.001 | 0 | 1.587 | 47.434 | 29.635 | 2.151 | 62.447 | 0.856 | -1.223 | 0 | 0 | 1 | 39 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-24278 | 2023/04/27 12:39:11 | 1-01 | 8 | -8.802 | -67.485 | 12657 | 0 | 1 | EV | standard | H2O | 1682621056.41 | 0.211 | 0.73 | 68.534 | 0.001 | -0.001 | 1.692 | 46.755 | 29.118 | 2.128 | 61.824 | 0.856 | -1.254 | 0 | 0 | 1 | 39 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-24278 | 2023/04/27 12:48:07 | 1-01 | 9 | -8.815 | -67.428 | 12733 | 0 | 1 | EV | standard | H2O | 1682621593.16 | 0.206 | 0.884 | 68.15 | 0.001 | -0.003 | 1.682 | 46.355 | 29.335 | 2.159 | 63.249 | 0.855 | -1.119 | 0 | 0 | 1 | 39 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-24278 | 2023/04/27 12:57:04 | 1-01 | 10 | -8.822 | -67.487 | 12773 | 0 | 1 | EV | standard | H2O | 1682622129.91 | 0.208 | 0.736 | 69.696 | 0.001 | 0.001 | 1.713 | 45.77 | 29.349 | 2.166 | 64.791 | 0.855 | -0.851 | 0 | 0 | 1 | 39 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-24279 | 2023/04/27 13:06:53 | 1-02 | 1 | 2.13 | 7.531 | 12644 | -1 | 1 | PZ | standard | H2O | 1682622718.26 | 0.184 | 0.793 | 66.444 | 0.001 | 0 | 1.646 | 43.879 | 29.248 | 2.261 | 64.794 | 0.855 | -1.129 | 0 | 0 | 1 | 39.25 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-24279 | 2023/04/27 13:15:49 | 1-02 | 2 | 2.667 | 15.006 | 12740 | -1 | 1 | PZ | standard | H2O | 1682623254.62 | 0.203 | 0.886 | 67.016 | 0 | -0.007 | 1.647 | 44.686 | 29.653 | 2.27 | 65.08 | 0.855 | -1.22 | 0 | 0 | 1 | 39.25 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-24279 | 2023/04/27 13:24:45 | 1-02 | 3 | 2.79 | 16.88 | 12624 | -1 | 1 | PZ | standard | H2O | 1682623790.66 | 0.2 | 0.924 | 66.056 | 0.001 | -0.005 | 1.636 | 44.58 | 29.044 | 2.264 | 65.073 | 0.855 | -1.22 | 0 | 0 | 1 | 39.125 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-24279 | 2023/04/27 13:33:42 | 1-02 | 4 | 2.854 | 18.014 | 12684 | 0 | 1 | PZ | standard | H2O | 1682624327.38 | 0.212 | 0.725 | 68.327 | 0 | -0.004 | 1.68 | 44.506 | 29.698 | 2.269 | 64.762 | 0.856 | -1.207 | 0 | 0 | 1 | 39.062 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-24279 | 2023/04/27 13:42:38 | 1-02 | 5 | 2.91 | 18.437 | 12787 | 0 | 1 | PZ | standard | H2O | 1682624864.31 | 0.207 | 0.868 | 68.428 | 0 | -0.005 | 1.687 | 44.619 | 29.894 | 2.285 | 65.17 | 0.856 | -1.221 | 0 | 0 | 1 | 39 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-24279 | 2023/04/27 13:51:36 | 1-02 | 6 | 2.953 | 18.769 | 12764 | 0 | 1 | PZ | standard | H2O | 1682625401.6 | 0.2 | 0.805 | 71.504 | 0.001 | -0.004 | 1.767 | 44.432 | 29.616 | 2.289 | 66.259 | 0.856 | -1.116 | 0 | 0 | 1 | 39 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-24279 | 2023/04/27 14:00:31 | 1-02 | 7 | 2.943 | 19.42 | 12565 | 0 | 1 | PZ | standard | H2O | 1682625936.65 | 0.215 | 0.824 | 67.829 | 0 | -0.004 | 1.675 | 43.847 | 29.293 | 2.263 | 64.668 | 0.856 | -1.179 | 0 | 0 | 1 | 39 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-24279 | 2023/04/27 14:09:28 | 1-02 | 8 | 2.923 | 19.303 | 12671 | 0 | 1 | PZ | standard | H2O | 1682626473.04 | 0.216 | 0.787 | 68.339 | 0.001 | -0.002 | 1.679 | 44.006 | 29.217 | 2.292 | 65.926 | 0.856 | -1.043 | 0 | 0 | 1 | 38.938 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-24279 | 2023/04/27 14:18:25 | 1-02 | 9 | 2.968 | 19.749 | 12760 | 0 | 1 | PZ | standard | H2O | 1682627010.05 | 0.204 | 0.902 | 64.645 | 0.001 | -0.002 | 1.597 | 43.963 | 29.574 | 2.311 | 67.13 | 0.856 | -0.938 | 0 | 0 | 1 | 38.938 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-24279 | 2023/04/27 14:27:22 | 1-02 | 10 | 2.996 | 19.712 | 12666 | 0 | 1 | PZ | standard | H2O | 1682627546.85 | 0.203 | 0.854 | 68.993 | 0.001 | 0.001 | 1.705 | 43.993 | 28.997 | 2.29 | 65.718 | 0.856 | -1.018 | 0 | 0 | 1 | 38.938 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-24280 | 2023/04/27 14:36:47 | 1-03 | 1 | [5 empty] | UT2 | standard | H2O | [17 empty] | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-24280 | 2023/04/27 14:46:46 | 1-03 | 2 | -13.658 | -96.212 | 13100 | -1 | 1 | UT2 | standard | H2O | 1682628711.63 | 0.211 | 0.816 | 67.692 | 0.002 | -0.003 | 1.673 | 44.551 | 30.533 | 2.175 | 64.251 | 0.856 | -0.944 | 0 | 0 | 1 | 38.938 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-24280 | 2023/04/27 14:55:43 | 1-03 | 3 | -14.294 | -106.125 | 12850 | -1 | 1 | UT2 | standard | H2O | 1682629248.44 | 0.208 | 0.845 | 66.877 | 0.002 | 0.006 | 1.649 | 44.184 | 29.714 | 2.102 | 61.997 | 0.856 | -1.055 | 0 | 0 | 1 | 38.938 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-24280 | 2023/04/27 15:04:41 | 1-03 | 4 | -14.392 | -108.58 | 12775 | 0 | 1 | UT2 | standard | H2O | 1682629786.33 | 0.198 | 0.779 | 65.476 | 0.002 | 0.004 | 1.616 | 43.91 | 29.054 | 2.099 | 63.807 | 0.856 | -0.815 | 0 | 0 | 1 | 38.875 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-24280 | 2023/04/27 15:13:37 | 1-03 | 5 | -14.493 | -109.969 | 12686 | 0 | 1 | UT2 | standard | H2O | 1682630322.77 | 0.22 | 0.818 | 69.482 | 0.002 | 0.005 | 1.727 | 43.773 | 29.251 | 2.078 | 61.923 | 0.856 | -0.934 | 0 | 0 | 1 | 38.875 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-24280 | 2023/04/27 15:22:35 | 1-03 | 6 | -14.449 | -110.298 | 12468 | 0 | 1 | UT2 | standard | H2O | 1682630860.05 | 0.2 | 0.698 | 66.416 | 0.001 | 0.002 | 1.653 | 43.352 | 28.477 | 2.056 | 61.316 | 0.856 | -0.981 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-24280 | 2023/04/27 15:31:31 | 1-03 | 7 | -14.589 | -111.161 | 12812 | 0 | 1 | UT2 | standard | H2O | 1682631396.54 | 0.194 | 0.779 | 64.202 | 0.001 | 0 | 1.58 | 43.986 | 29.538 | 2.103 | 62.473 | 0.856 | -0.977 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-24280 | 2023/04/27 15:40:27 | 1-03 | 8 | -14.563 | -111.23 | 12772 | 0 | 1 | UT2 | standard | H2O | 1682631932.48 | 0.197 | 0.806 | 65.578 | 0.001 | 0 | 1.621 | 43.772 | 29.18 | 2.106 | 63.38 | 0.855 | -0.809 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-24280 | 2023/04/27 15:49:24 | 1-03 | 9 | -14.521 | -111.477 | 12535 | 0 | 1 | UT2 | standard | H2O | 1682632469.26 | 0.204 | 0.784 | 64.191 | 0.001 | 0.003 | 1.587 | 43.288 | 28.966 | 2.043 | 60.996 | 0.856 | -0.927 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-24280 | 2023/04/27 15:58:20 | 1-03 | 10 | -14.544 | -111.474 | 12462 | 0 | 1 | UT2 | standard | H2O | 1682633006.12 | 0.185 | 0.73 | 65.252 | 0.001 | 0.003 | 1.615 | 43.252 | 28.357 | 2.061 | 61.707 | 0.856 | -0.976 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-24280 | 2023/04/27 16:07:16 | 1-03 | 11 | -14.606 | -111.978 | 12569 | 0 | 1 | UT2 | standard | H2O | 1682633541.91 | 0.205 | 0.894 | 67.159 | 0.001 | 0.005 | 1.658 | 43.454 | 28.606 | 2.068 | 61.611 | 0.856 | -0.958 | 0 | 0 | 1 | 38.812 | 1 | 3 | 1 | GJB_Rinse | 0
32 | P-24281 | 2023/04/27 16:17:05 | 1-04 | 1 | -9.247 | -72.426 | 12600 | -1 | 1 | EV | standard | H2O | 1682634130.87 | 0.199 | 0.838 | 68.131 | 0.001 | 0.003 | 1.685 | 43.298 | 29.086 | 2.13 | 62.942 | 0.856 | -0.97 | 0 | 0 | 1 | 38.812 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-24281 | 2023/04/27 16:26:02 | 1-04 | 2 | -9 | -69.15 | 12604 | -1 | 1 | EV | standard | H2O | 1682634667.69 | 0.197 | 0.842 | 69.805 | 0 | 0.001 | 1.725 | 43.394 | 28.913 | 2.133 | 63.373 | 0.856 | -0.97 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-24281 | 2023/04/27 16:34:59 | 1-04 | 3 | -8.869 | -67.662 | 12526 | -1 | 1 | EV | standard | H2O | 1682635204.27 | 0.202 | 0.731 | 63.366 | 0.001 | -0.001 | 1.564 | 43.212 | 28.64 | 2.131 | 63.769 | 0.855 | -0.872 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-24281 | 2023/04/27 16:43:56 | 1-04 | 4 | -8.934 | -67.707 | 12451 | 0 | 1 | EV | standard | H2O | 1682635741.08 | 0.207 | 0.772 | 65.985 | 0.001 | -0.001 | 1.625 | 43.13 | 28.195 | 2.129 | 63.303 | 0.856 | -0.87 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-24281 | 2023/04/27 16:52:52 | 1-04 | 5 | -8.89 | -67.432 | 12445 | 0 | 1 | EV | standard | H2O | 1682636277.06 | 0.199 | 0.807 | 68.384 | 0.001 | 0 | 1.692 | 43.153 | 28.09 | 2.121 | 63.324 | 0.856 | -0.888 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-24281 | 2023/04/27 17:01:48 | 1-04 | 6 | -8.932 | -67.399 | 12453 | 0 | 1 | EV | standard | H2O | 1682636813.82 | 0.21 | 0.794 | 69.242 | 0.001 | -0.001 | 1.712 | 43.081 | 28.45 | 2.128 | 62.459 | 0.856 | -1.01 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-24281 | 2023/04/27 17:10:44 | 1-04 | 7 | -8.904 | -67.133 | 12246 | 0 | 1 | EV | standard | H2O | 1682637349.78 | 0.218 | 0.755 | 70.756 | 0.002 | 0 | 1.753 | 42.588 | 27.854 | 2.088 | 62.142 | 0.855 | -0.878 | 0 | 0 | 1 | 38.875 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-24281 | 2023/04/27 17:19:40 | 1-04 | 8 | -8.946 | -67.447 | 12241 | 0 | 1 | EV | standard | H2O | 1682637885.77 | 0.183 | 0.72 | 63.779 | 0 | -0.001 | 1.586 | 42.59 | 28.093 | 2.08 | 61.33 | 0.856 | -0.956 | 0 | 0 | 1 | 38.812 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-24281 | 2023/04/27 17:28:36 | 1-04 | 9 | -8.909 | -67.086 | 12275 | 0 | 1 | EV | standard | H2O | 1682638421.59 | 0.209 | 0.828 | 67.342 | 0.001 | 0 | 1.674 | 42.658 | 27.908 | 2.09 | 62.357 | 0.856 | -1.051 | 0 | 0 | 1 | 38.812 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-24281 | 2023/04/27 17:37:32 | 1-04 | 10 | -8.915 | -67.075 | 12298 | 0 | 1 | EV | standard | H2O | 1682638957.57 | 0.186 | 0.803 | 67.549 | 0.001 | 0 | 1.678 | 42.695 | 28.102 | 2.1 | 61.604 | 0.856 | -1.06 | 0 | 0 | 1 | 38.812 | 1 | 4 | 1 | GJB_Rinse | 0
42 | P-24282 | 2023/04/27 17:47:22 | 1-05 | 1 | -6.794 | -46.184 | 12511 | -1 | 1 | A00000322632 | 23-097 | H2O | 1682639548.03 | 0.214 | 0.914 | 66.081 | 0.001 | -0.001 | 1.633 | 43.055 | 28.665 | 2.168 | 63.38 | 0.856 | -0.98 | 0 | 0 | 1 | 38.812 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-24282 | 2023/04/27 17:56:18 | 1-05 | 2 | -6.688 | -44.054 | 12302 | -1 | 1 | A00000322632 | 23-097 | H2O | 1682640083.91 | 0.214 | 0.784 | 64.837 | 0.001 | -0.003 | 1.6 | 42.686 | 28.051 | 2.136 | 62.699 | 0.856 | -0.978 | 0 | 0 | 1 | 38.812 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-24282 | 2023/04/27 18:05:15 | 1-05 | 3 | -6.686 | -43.638 | 12521 | -1 | 1 | A00000322632 | 23-097 | H2O | 1682640620.48 | 0.195 | 0.767 | 68.749 | 0.002 | -0.001 | 1.695 | 43.216 | 28.593 | 2.166 | 63.75 | 0.856 | -1.052 | 0 | 0 | 1 | 38.812 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-24282 | 2023/04/27 18:14:11 | 1-05 | 4 | -6.632 | -43.264 | 12395 | 0 | 1 | A00000322632 | 23-097 | H2O | 1682641156.94 | 0.225 | 0.782 | 65.727 | 0.002 | -0.005 | 1.63 | 42.918 | 28.325 | 2.141 | 62.958 | 0.856 | -0.998 | 0 | 0 | 1 | 38.75 | 1 | 5 | 1 | GJB_Rinse | 0
46 | P-24283 | 2023/04/27 18:24:02 | 1-06 | 1 | -6.811 | -44.26 | 12480 | -1 | 1 | A00000322644 | 23-097 | H2O | 1682641747.45 | 0.205 | 0.795 | 67.862 | 0.001 | -0.002 | 1.68 | 43.036 | 28.437 | 2.157 | 64.007 | 0.856 | -0.865 | 0 | 0 | 1 | 38.75 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-24283 | 2023/04/27 18:32:59 | 1-06 | 2 | -6.821 | -44.554 | 12329 | -1 | 1 | A00000322644 | 23-097 | H2O | 1682642285.17 | 0.206 | 0.82 | 68.914 | 0.001 | 0 | 1.705 | 42.705 | 28.142 | 2.147 | 63.299 | 0.856 | -0.807 | 0 | 0 | 1 | 38.688 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-24283 | 2023/04/27 18:41:55 | 1-06 | 3 | -6.766 | -44.307 | 12798 | -1 | 1 | A00000322644 | 23-097 | H2O | 1682642820.49 | 0.197 | 0.745 | 67.41 | 0.001 | -0.001 | 1.663 | 43.771 | 29.651 | 2.201 | 63.698 | 0.857 | -1.133 | 0 | 0 | 1 | 38.562 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-24283 | 2023/04/27 18:50:51 | 1-06 | 4 | -6.83 | -44.231 | 12601 | 0 | 1 | A00000322644 | 23-097 | H2O | 1682643356.96 | 0.203 | 0.837 | 65.521 | 0.001 | 0.001 | 1.615 | 43.256 | 28.888 | 2.188 | 64.112 | 0.856 | -1.038 | 0 | 0 | 1 | 38.562 | 1 | 6 | 1 | GJB_Rinse | 0
50 | P-24284 | 2023/04/27 19:00:41 | 1-07 | 1 | -6.515 | -42.744 | 12620 | -1 | 1 | A00000322660 | 23-097 | H2O | 1682643946.55 | 0.177 | 0.819 | 64.246 | 0.001 | -0.001 | 1.585 | 43.365 | 29.153 | 2.191 | 63.691 | 0.857 | -1.062 | 0 | 0 | 1 | 38.5 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-24284 | 2023/04/27 19:09:37 | 1-07 | 2 | -6.466 | -42.36 | 12682 | -1 | 1 | A00000322660 | 23-097 | H2O | 1682644482.59 | 0.194 | 0.844 | 63.021 | 0.001 | -0.001 | 1.555 | 43.597 | 29.084 | 2.201 | 63.89 | 0.856 | -1.126 | 0 | 0 | 1 | 38.375 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-24284 | 2023/04/27 19:18:33 | 1-07 | 3 | -6.404 | -41.913 | 12377 | -1 | 1 | A00000322660 | 23-097 | H2O | 1682645019 | 0.203 | 0.744 | 65.204 | 0.001 | 0.002 | 1.61 | 42.983 | 28.232 | 2.154 | 63.416 | 0.855 | -1.027 | 0 | 0 | 1 | 38.438 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-24284 | 2023/04/27 19:27:29 | 1-07 | 4 | -6.407 | -41.99 | 12375 | 0 | 1 | A00000322660 | 23-097 | H2O | 1682645554.77 | 0.195 | 0.937 | 69.188 | 0.002 | 0.003 | 1.715 | 42.993 | 28.244 | 2.146 | 62.799 | 0.856 | -0.995 | 0 | 0 | 1 | 38.375 | 1 | 7 | 1 | GJB_Rinse | 0
54 | P-24285 | 2023/04/27 19:37:17 | 1-08 | 1 | -6.435 | -41.317 | 12112 | -1 | 1 | A00000322734 | 23-097 | H2O | 1682646142.41 | 0.198 | 0.916 | 67.248 | 0.001 | 0.002 | 1.671 | 42.451 | 27.618 | 2.096 | 60.673 | 0.856 | -1.095 | 0 | 0 | 1 | 38.438 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-24285 | 2023/04/27 19:46:14 | 1-08 | 2 | -6.494 | -41.635 | 12120 | -1 | 1 | A00000322734 | 23-097 | H2O | 1682646680.14 | 0.201 | 0.707 | 62.641 | 0 | -0.002 | 1.559 | 42.282 | 27.667 | 2.103 | 62.156 | 0.855 | -0.969 | 0 | 0 | 1 | 38.375 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-24285 | 2023/04/27 19:55:11 | 1-08 | 3 | -6.453 | -41.345 | 12283 | -1 | 1 | A00000322734 | 23-097 | H2O | 1682647216.65 | 0.208 | 0.708 | 63.558 | 0.001 | -0.001 | 1.573 | 42.533 | 28.194 | 2.136 | 62.918 | 0.855 | -0.873 | 0 | 0 | 1 | 38.438 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-24285 | 2023/04/27 20:04:07 | 1-08 | 4 | -6.454 | -41.494 | 12315 | 0 | 1 | A00000322734 | 23-097 | H2O | 1682647753.38 | 0.178 | 0.787 | 61.072 | 0.001 | 0 | 1.51 | 42.856 | 28.283 | 2.122 | 61.639 | 0.856 | -1.135 | 0 | 0 | 1 | 38.438 | 1 | 8 | 1 | GJB_Rinse | 0
58 | P-24286 | 2023/04/27 20:13:55 | 1-09 | 1 | -6.676 | -43.023 | 12628 | -1 | 1 | A00000322666 | 23-097 | H2O | 1682648341.02 | 0.213 | 0.729 | 66.923 | 0.002 | 0 | 1.662 | 43.256 | 28.862 | 2.188 | 64.855 | 0.856 | -0.792 | 0 | 0 | 1 | 38.375 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-24286 | 2023/04/27 20:22:51 | 1-09 | 2 | -6.63 | -42.817 | 12308 | -1 | 1 | A00000322666 | 23-097 | H2O | 1682648876.49 | 0.196 | 0.809 | 64.045 | 0.001 | -0.001 | 1.593 | 42.721 | 28.142 | 2.13 | 62.693 | 0.856 | -0.941 | 0 | 0 | 1 | 38.438 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-24286 | 2023/04/27 20:31:47 | 1-09 | 3 | -6.653 | -42.643 | 12296 | -1 | 1 | A00000322666 | 23-097 | H2O | 1682649412.72 | 0.214 | 0.811 | 61.479 | 0.001 | -0.001 | 1.529 | 42.768 | 28.353 | 2.136 | 61.793 | 0.856 | -1.086 | 0 | 0 | 1 | 38.375 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-24286 | 2023/04/27 20:40:43 | 1-09 | 4 | -6.637 | -42.863 | 12304 | 0 | 1 | A00000322666 | 23-097 | H2O | 1682649948.74 | 0.21 | 0.77 | 68.626 | 0.002 | -0.001 | 1.707 | 42.662 | 28.491 | 2.133 | 61.977 | 0.857 | -1.124 | 0 | 0 | 1 | 38.375 | 1 | 9 | 1 | GJB_Rinse | 0
62 | P-24287 | 2023/04/27 20:50:31 | 1-10 | 1 | -6.208 | -39.934 | 12490 | -1 | 1 | A00000322667 | 23-097 | H2O | 1682650536.7 | 0.195 | 0.825 | 64.307 | 0.001 | 0.001 | 1.591 | 43.07 | 28.736 | 2.16 | 63.655 | 0.856 | -1.008 | 0 | 0 | 1 | 38.438 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-24287 | 2023/04/27 20:59:26 | 1-10 | 2 | -6.182 | -39.303 | 12252 | -1 | 1 | A00000322667 | 23-097 | H2O | 1682651072.19 | 0.197 | 0.85 | 63.692 | 0.002 | 0.001 | 1.577 | 42.481 | 28.125 | 2.113 | 62.974 | 0.855 | -0.886 | 0 | 0 | 1 | 38.5 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-24287 | 2023/04/27 21:08:23 | 1-10 | 3 | -6.097 | -38.8 | 12205 | -1 | 1 | A00000322667 | 23-097 | H2O | 1682651609.18 | 0.202 | 0.903 | 66.062 | 0.001 | 0.001 | 1.642 | 42.371 | 27.79 | 2.119 | 63.376 | 0.856 | -0.769 | 0 | 0 | 1 | 38.5 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-24287 | 2023/04/27 21:17:20 | 1-10 | 4 | -6.083 | -38.936 | 12225 | 0 | 1 | A00000322667 | 23-097 | H2O | 1682652146.09 | 0.184 | 0.835 | 62.344 | 0.001 | -0.001 | 1.551 | 42.563 | 28.018 | 2.117 | 61.93 | 0.856 | -1.119 | 0 | 0 | 1 | 38.562 | 1 | 10 | 1 | GJB_Rinse | 0
66 | P-24288 | 2023/04/27 21:27:09 | 1-11 | 1 | -6.805 | -43.551 | 12649 | -1 | 1 | A00000322665 | 23-097 | H2O | 1682652734.63 | 0.192 | 0.804 | 61.651 | 0.001 | -0.004 | 1.523 | 43.425 | 29.007 | 2.181 | 64.158 | 0.856 | -0.993 | 0 | 0 | 1 | 38.562 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-24288 | 2023/04/27 21:36:06 | 1-11 | 2 | -6.753 | -43.592 | 12274 | -1 | 1 | A00000322665 | 23-097 | H2O | 1682653271.51 | 0.187 | 0.839 | 65.126 | 0.001 | -0.001 | 1.609 | 42.537 | 28.569 | 2.108 | 61.539 | 0.856 | -0.97 | 0 | 0 | 1 | 38.562 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-24288 | 2023/04/27 21:45:02 | 1-11 | 3 | -6.734 | -43.597 | 12290 | -1 | 1 | A00000322665 | 23-097 | H2O | 1682653808.43 | 0.189 | 0.825 | 67.78 | 0.001 | -0.002 | 1.677 | 42.731 | 28.19 | 2.117 | 62.125 | 0.856 | -0.97 | 0 | 0 | 1 | 38.562 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-24288 | 2023/04/27 21:53:59 | 1-11 | 4 | -6.761 | -43.677 | 12282 | 0 | 1 | A00000322665 | 23-097 | H2O | 1682654345.06 | 0.194 | 0.768 | 65.015 | 0.001 | 0 | 1.607 | 42.571 | 28.046 | 2.132 | 62.906 | 0.856 | -0.876 | 0 | 0 | 1 | 38.625 | 1 | 11 | 1 | GJB_Rinse | 0
70 | P-24289 | 2023/04/27 22:03:49 | 1-12 | 1 | -6.887 | -44.734 | 12552 | -1 | 1 | A00000322668 | 23-097 | H2O | 1682654934.44 | 0.202 | 0.74 | 59.708 | 0.001 | 0.001 | 1.478 | 43.176 | 28.823 | 2.159 | 63.385 | 0.856 | -0.936 | 0 | 0 | 1 | 38.688 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-24289 | 2023/04/27 22:12:45 | 1-12 | 2 | -6.871 | -44.909 | 12284 | -1 | 1 | A00000322668 | 23-097 | H2O | 1682655470.87 | 0.194 | 0.78 | 63.878 | 0.001 | -0.001 | 1.588 | 42.608 | 28.158 | 2.113 | 62.798 | 0.856 | -0.95 | 0 | 0 | 1 | 38.75 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-24289 | 2023/04/27 22:21:42 | 1-12 | 3 | -6.894 | -44.953 | 12232 | -1 | 1 | A00000322668 | 23-097 | H2O | 1682656007.91 | 0.203 | 0.752 | 67.265 | 0.001 | -0.005 | 1.668 | 42.478 | 27.843 | 2.12 | 63.38 | 0.856 | -0.883 | 0 | 0 | 1 | 38.75 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-24289 | 2023/04/27 22:30:38 | 1-12 | 4 | -6.958 | -45.043 | 12270 | 0 | 1 | A00000322668 | 23-097 | H2O | 1682656544.3 | 0.189 | 0.762 | 66.739 | 0.001 | 0.002 | 1.648 | 42.608 | 28.052 | 2.114 | 62.345 | 0.856 | -1.052 | 0 | 0 | 1 | 38.75 | 1 | 12 | 1 | GJB_Rinse | 0
74 | P-24290 | 2023/04/27 22:40:28 | 1-13 | 1 | -8.821 | -63.839 | 12419 | -1 | 1 | EV | standard | H2O | 1682657133.01 | 0.191 | 0.784 | 65.118 | 0.001 | 0 | 1.609 | 42.906 | 28.66 | 2.123 | 62.141 | 0.856 | -1.057 | 0 | 0 | 1 | 38.812 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-24290 | 2023/04/27 22:49:24 | 1-13 | 2 | -8.979 | -65.983 | 12336 | -1 | 1 | EV | standard | H2O | 1682657670.44 | 0.194 | 0.884 | 65.436 | 0.001 | 0.003 | 1.617 | 42.915 | 27.894 | 2.099 | 62.502 | 0.856 | -1.048 | 0 | 0 | 1 | 38.812 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-24290 | 2023/04/27 22:58:20 | 1-13 | 3 | -8.943 | -65.997 | 12271 | -1 | 1 | EV | standard | H2O | 1682658205.98 | 0.193 | 0.715 | 64.588 | 0 | 0.003 | 1.604 | 42.619 | 28.162 | 2.096 | 61.538 | 0.856 | -1.017 | 0 | 0 | 1 | 38.875 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-24290 | 2023/04/27 23:07:17 | 1-13 | 4 | -8.975 | -66.174 | 12260 | 0 | 1 | EV | standard | H2O | 1682658743.11 | 0.203 | 0.888 | 58.709 | 0.001 | -0.002 | 1.458 | 42.578 | 28.057 | 2.091 | 61.31 | 0.856 | -0.905 | 0 | 0 | 1 | 38.875 | 1 | 13 | 1 | GJB_Rinse | 0
78 | P-24291 | 2023/04/27 23:17:07 | 1-14 | 1 | -4.618 | -38.126 | 12418 | -1 | 1 | A00000322423 | 23-097 | H2O | 1682659332.88 | 0.215 | 0.816 | 64.069 | 0 | 0.002 | 1.581 | 42.876 | 28.455 | 2.162 | 63.15 | 0.856 | -0.995 | 0 | 0 | 1 | 38.812 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-24291 | 2023/04/27 23:26:03 | 1-14 | 2 | -4.327 | -35.175 | 12264 | -1 | 1 | A00000322423 | 23-097 | H2O | 1682659869.21 | 0.218 | 0.839 | 66.841 | 0.001 | 0.002 | 1.665 | 42.444 | 28.524 | 2.13 | 61.622 | 0.856 | -1.053 | 0 | 0 | 1 | 38.875 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-24291 | 2023/04/27 23:35:00 | 1-14 | 3 | -4.278 | -34.314 | 12176 | -1 | 1 | A00000322423 | 23-097 | H2O | 1682660405.82 | 0.207 | 0.856 | 64.586 | 0.001 | -0.004 | 1.597 | 42.361 | 27.762 | 2.125 | 62.589 | 0.856 | -0.914 | 0 | 0 | 1 | 38.938 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-24291 | 2023/04/27 23:43:57 | 1-14 | 4 | -4.23 | -33.784 | 12159 | 0 | 1 | A00000322423 | 23-097 | H2O | 1682660942.18 | 0.202 | 0.877 | 62.622 | 0.001 | 0.002 | 1.559 | 42.426 | 27.555 | 2.115 | 62.771 | 0.856 | -1.015 | 0 | 0 | 1 | 38.938 | 1 | 14 | 1 | GJB_Rinse | 0
82 | P-24292 | 2023/04/27 23:53:44 | 1-15 | 1 | -7.433 | -51.963 | 12481 | -1 | 1 | A00000322427 | 23-097 | H2O | 1682661529.25 | 0.193 | 0.793 | 70.86 | 0.001 | 0 | 1.751 | 42.93 | 29.102 | 2.14 | 62.527 | 0.856 | -1.098 | 0 | 0 | 1 | 39 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-24292 | 2023/04/28 00:02:39 | 1-15 | 2 | -7.619 | -54.002 | 12258 | -1 | 1 | A00000322427 | 23-097 | H2O | 1682662065.24 | 0.209 | 0.779 | 66.987 | 0.001 | 0 | 1.655 | 42.629 | 28.128 | 2.101 | 61.794 | 0.856 | -1.015 | 0 | 0 | 1 | 39 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-24292 | 2023/04/28 00:11:36 | 1-15 | 3 | -7.544 | -54.07 | 12266 | -1 | 1 | A00000322427 | 23-097 | H2O | 1682662601.6 | 0.197 | 0.792 | 61.262 | 0.001 | -0.001 | 1.52 | 42.646 | 28.219 | 2.093 | 61.355 | 0.857 | -1.055 | 0 | 0 | 1 | 38.938 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-24292 | 2023/04/28 00:20:31 | 1-15 | 4 | -7.585 | -54.167 | 12197 | 0 | 1 | A00000322427 | 23-097 | H2O | 1682663137.2 | 0.188 | 0.755 | 65.409 | 0.002 | 0 | 1.617 | 42.494 | 27.957 | 2.089 | 61.812 | 0.856 | -1.001 | 0 | 0 | 1 | 38.938 | 1 | 15 | 1 | GJB_Rinse | 0
86 | P-24293 | 2023/04/28 00:30:19 | 1-16 | 1 | -13.382 | -105.572 | 12333 | -1 | 1 | A00000170752 | 23-097 | H2O | 1682663724.42 | 0.212 | 0.758 | 62.794 | 0.002 | 0 | 1.56 | 42.866 | 28.48 | 2.021 | 59.895 | 0.855 | -0.968 | 0 | 0 | 1 | 39 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-24293 | 2023/04/28 00:39:15 | 1-16 | 2 | -13.694 | -110.597 | 12071 | -1 | 1 | A00000170752 | 23-097 | H2O | 1682664260.72 | 0.195 | 0.851 | 63.047 | 0.001 | 0 | 1.569 | 42.336 | 27.49 | 1.982 | 59.53 | 0.856 | -0.987 | 0 | 0 | 1 | 39 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-24293 | 2023/04/28 00:48:11 | 1-16 | 3 | -13.738 | -111.919 | 12084 | -1 | 1 | A00000170752 | 23-097 | H2O | 1682664797.47 | 0.211 | 0.73 | 63.748 | 0.001 | 0.002 | 1.584 | 42.407 | 27.585 | 1.989 | 58.855 | 0.856 | -1.02 | 0 | 0 | 1 | 38.812 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-24293 | 2023/04/28 00:57:08 | 1-16 | 4 | -13.754 | -112.8 | 12049 | 0 | 1 | A00000170752 | 23-097 | H2O | 1682665334.11 | 0.192 | 0.749 | 64.233 | 0.001 | 0 | 1.589 | 42.322 | 27.011 | 1.978 | 60.639 | 0.856 | -0.899 | 0 | 0 | 1 | 38.812 | 1 | 16 | 1 | GJB_Rinse | 0
90 | P-24294 | 2023/04/28 01:06:57 | 1-17 | 1 | -3.782 | -30.628 | 12428 | -1 | 1 | A00000170760 | 23-097 | H2O | 1682665922.63 | 0.208 | 0.711 | 68.099 | 0.001 | 0.003 | 1.682 | 43.036 | 28.41 | 2.155 | 63.338 | 0.856 | -1.045 | 0 | 0 | 1 | 39 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-24294 | 2023/04/28 01:15:53 | 1-17 | 2 | -3.301 | -22.916 | 12162 | -1 | 1 | A00000170760 | 23-097 | H2O | 1682666458.4 | 0.212 | 0.86 | 61.325 | 0 | 0.001 | 1.516 | 42.475 | 28.043 | 2.122 | 61.255 | 0.856 | -1.14 | 0 | 0 | 1 | 39.125 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-24294 | 2023/04/28 01:24:48 | 1-17 | 3 | -3.213 | -20.909 | 12089 | -1 | 1 | A00000170760 | 23-097 | H2O | 1682666993.84 | 0.204 | 0.82 | 64.049 | 0.001 | -0.004 | 1.587 | 42.36 | 27.848 | 2.116 | 61.49 | 0.855 | -1.158 | 0 | 0 | 1 | 39.312 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-24294 | 2023/04/28 01:33:43 | 1-17 | 4 | -3.149 | -20.053 | 12045 | 0 | 1 | A00000170760 | 23-097 | H2O | 1682667529.27 | 0.184 | 0.852 | 60.946 | 0 | -0.001 | 1.507 | 42.117 | 28.114 | 2.107 | 60.301 | 0.855 | -1.114 | 0 | 0 | 1 | 39.438 | 1 | 17 | 1 | GJB_Rinse | 0
94 | P-24295 | 2023/04/28 01:43:32 | 1-18 | 1 | -5.712 | -38.15 | 12383 | -1 | 1 | A00000322421 | 23-097 | H2O | 1682668117.29 | 0.197 | 0.758 | 62.664 | 0.001 | -0.003 | 1.548 | 42.878 | 28.652 | 2.138 | 61.838 | 0.855 | -1.17 | 0 | 0 | 1 | 39.5 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-24295 | 2023/04/28 01:52:27 | 1-18 | 2 | -5.833 | -40.129 | 12166 | -1 | 1 | A00000322421 | 23-097 | H2O | 1682668652.59 | 0.203 | 0.9 | 67.276 | 0.002 | -0.004 | 1.673 | 42.506 | 27.834 | 2.09 | 61.554 | 0.855 | -1.08 | 0 | 0 | 1 | 39.5 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-24295 | 2023/04/28 02:01:24 | 1-18 | 3 | -5.828 | -40.126 | 12187 | -1 | 1 | A00000322421 | 23-097 | H2O | 1682669188.71 | 0.194 | 0.849 | 66.762 | 0.001 | -0.001 | 1.662 | 42.488 | 27.546 | 2.118 | 62.986 | 0.855 | -0.912 | 0 | 0 | 1 | 39.5 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-24295 | 2023/04/28 02:10:20 | 1-18 | 4 | -5.858 | -40.227 | 12221 | 0 | 1 | A00000322421 | 23-097 | H2O | 1682669725.88 | 0.201 | 0.738 | 63.751 | 0.001 | 0.003 | 1.578 | 42.502 | 28.125 | 2.115 | 62.124 | 0.856 | -0.965 | 0 | 0 | 1 | 39.5 | 1 | 18 | 1 | GJB_Rinse | 0
98 | P-24296 | 2023/04/28 02:20:08 | 1-19 | 1 | -7.643 | -48.355 | 12361 | -1 | 1 | A00000170829 | 23-097 | H2O | 1682670313.78 | 0.198 | 0.726 | 66.158 | 0.001 | 0.001 | 1.64 | 42.801 | 28.668 | 2.118 | 61.606 | 0.855 | -1.01 | 0 | 0 | 1 | 39.562 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-24296 | 2023/04/28 02:29:05 | 1-19 | 2 | -7.762 | -49.481 | 12106 | -1 | 1 | A00000170829 | 23-097 | H2O | 1682670850.26 | 0.207 | 0.748 | 64.858 | 0.001 | -0.002 | 1.615 | 42.274 | 27.806 | 2.074 | 61.152 | 0.855 | -0.995 | 0 | 0 | 1 | 39.562 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-24296 | 2023/04/28 02:38:01 | 1-19 | 3 | -7.724 | -49.511 | 12120 | -1 | 1 | A00000170829 | 23-097 | H2O | 1682671387.02 | 0.224 | 0.888 | 63.158 | 0.001 | 0 | 1.557 | 42.366 | 27.761 | 2.078 | 61.134 | 0.855 | -0.963 | 0 | 0 | 1 | 39.5 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-24296 | 2023/04/28 02:46:56 | 1-19 | 4 | -7.817 | -49.868 | 12123 | 0 | 1 | A00000170829 | 23-097 | H2O | 1682671922.25 | 0.223 | 0.841 | 64.022 | 0.001 | 0.004 | 1.59 | 42.304 | 27.745 | 2.087 | 60.868 | 0.855 | -0.915 | 0 | 0 | 1 | 39.562 | 1 | 19 | 1 | GJB_Rinse | 0
102 | P-24297 | 2023/04/28 02:56:45 | 1-20 | 1 | -14.314 | -105.474 | 12269 | -1 | 1 | A00000170754 | 23-097 | H2O | 1682672509.95 | 0.209 | 0.817 | 67.724 | 0.001 | 0.002 | 1.677 | 42.629 | 28.228 | 2.026 | 60.321 | 0.855 | -1.071 | 0 | 0 | 1 | 39.562 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-24297 | 2023/04/28 03:05:39 | 1-20 | 2 | -14.624 | -111.152 | 11974 | -1 | 1 | A00000170754 | 23-097 | H2O | 1682673045.71 | 0.223 | 0.836 | 65.624 | 0.002 | 0.004 | 1.646 | 42.059 | 27.144 | 1.978 | 59.203 | 0.855 | -1.042 | 0 | 0 | 1 | 39.562 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-24297 | 2023/04/28 03:14:36 | 1-20 | 3 | -14.737 | -112.662 | 11967 | -1 | 1 | A00000170754 | 23-097 | H2O | 1682673581.81 | 0.191 | 0.675 | 65.777 | 0.001 | 0.003 | 1.639 | 42.051 | 27.204 | 1.98 | 59.519 | 0.855 | -0.883 | 0 | 0 | 1 | 39.5 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-24297 | 2023/04/28 03:23:31 | 1-20 | 4 | -14.741 | -113.445 | 12052 | 0 | 1 | A00000170754 | 23-097 | H2O | 1682674116.99 | 0.197 | 0.72 | 64.097 | 0.001 | 0.003 | 1.589 | 42.284 | 27.371 | 1.966 | 59.451 | 0.855 | -1.018 | 0 | 0 | 1 | 39.562 | 1 | 20 | 1 | GJB_Rinse | 0
106 | P-24298 | 2023/04/28 03:33:20 | 1-21 | 1 | -5.007 | -39.298 | 12404 | -1 | 1 | A00000170758 | 23-097 | H2O | 1682674705.67 | 0.176 | 0.719 | 61.525 | 0 | 0.002 | 1.522 | 42.78 | 28.533 | 2.15 | 63.027 | 0.856 | -1.004 | 0 | 0 | 1 | 39.438 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-24298 | 2023/04/28 03:42:16 | 1-21 | 2 | -4.505 | -32.211 | 12153 | -1 | 1 | A00000170758 | 23-097 | H2O | 1682675241.93 | 0.188 | 0.807 | 62.33 | 0.001 | 0.001 | 1.552 | 42.388 | 27.99 | 2.107 | 61.821 | 0.855 | -1.081 | 0 | 0 | 1 | 39.375 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-24298 | 2023/04/28 03:51:13 | 1-21 | 3 | -4.443 | -30.494 | 12156 | -1 | 1 | A00000170758 | 23-097 | H2O | 1682675778.14 | 0.205 | 0.841 | 61.454 | 0.001 | -0.007 | 1.52 | 42.347 | 27.802 | 2.131 | 62.28 | 0.855 | -1.001 | 0 | 0 | 1 | 39.375 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-24298 | 2023/04/28 04:00:09 | 1-21 | 4 | -4.43 | -29.799 | 12174 | 0 | 1 | A00000170758 | 23-097 | H2O | 1682676314.92 | 0.21 | 0.8 | 64.246 | 0.001 | -0.003 | 1.598 | 42.277 | 27.782 | 2.122 | 63.049 | 0.855 | -0.84 | 0 | 0 | 1 | 39.375 | 1 | 21 | 1 | GJB_Rinse | 0
110 | P-24299 | 2023/04/28 04:09:54 | 1-22 | 1 | -8.717 | -62.306 | 12437 | -1 | 1 | EV | standard | H2O | 1682676900.07 | 0.199 | 0.88 | 66.033 | 0.001 | -0.001 | 1.623 | 42.957 | 28.304 | 2.121 | 62.917 | 0.855 | -0.963 | 0 | 0 | 1 | 39.438 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-24299 | 2023/04/28 04:18:51 | 1-22 | 2 | -8.917 | -65.497 | 12150 | -1 | 1 | EV | standard | H2O | 1682677435.82 | 0.23 | 0.734 | 64.572 | 0.002 | 0 | 1.605 | 42.36 | 27.766 | 2.068 | 60.776 | 0.856 | -1.074 | 0 | 0 | 1 | 39.438 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-24299 | 2023/04/28 04:27:47 | 1-22 | 3 | -8.947 | -66.14 | 12169 | -1 | 1 | EV | standard | H2O | 1682677972.45 | 0.182 | 0.803 | 65.895 | 0.001 | 0.001 | 1.631 | 42.402 | 27.853 | 2.062 | 61.46 | 0.855 | -1.047 | 0 | 0 | 1 | 39.438 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-24299 | 2023/04/28 04:36:43 | 1-22 | 4 | -8.964 | -66.517 | 12152 | 0 | 1 | EV | standard | H2O | 1682678508.17 | 0.212 | 0.731 | 63.748 | 0.001 | -0.003 | 1.577 | 42.442 | 27.881 | 2.063 | 60.451 | 0.855 | -1.183 | 0 | 0 | 1 | 39.438 | 1 | 22 | 1 | GJB_Rinse | 0
114 | P-24300 | 2023/04/28 04:46:29 | 1-23 | 1 | -5.704 | -37.664 | 12355 | -1 | 1 | A00000322425 | 23-097 | H2O | 1682679095.11 | 0.207 | 0.756 | 65.325 | 0.001 | 0 | 1.618 | 42.817 | 28.129 | 2.142 | 63.661 | 0.855 | -0.974 | 0 | 0 | 1 | 39.562 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-24300 | 2023/04/28 04:55:25 | 1-23 | 2 | -5.56 | -35.224 | 12113 | -1 | 1 | A00000322425 | 23-097 | H2O | 1682679630.95 | 0.193 | 0.788 | 63.281 | 0 | 0 | 1.567 | 42.387 | 27.91 | 2.092 | 60.699 | 0.855 | -1.168 | 0 | 0 | 1 | 39.5 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-24300 | 2023/04/28 05:04:21 | 1-23 | 3 | -5.505 | -34.058 | 12072 | -1 | 1 | A00000322425 | 23-097 | H2O | 1682680166.6 | 0.228 | 0.83 | 65.426 | 0.001 | -0.002 | 1.619 | 42.266 | 27.738 | 2.094 | 60.69 | 0.855 | -0.965 | 0 | 0 | 1 | 39.5 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-24300 | 2023/04/28 05:13:17 | 1-23 | 4 | -5.503 | -33.739 | 12072 | 0 | 1 | A00000322425 | 23-097 | H2O | 1682680702.71 | 0.201 | 1.013 | 61.063 | 0.001 | -0.002 | 1.508 | 42.281 | 27.619 | 2.098 | 60.989 | 0.855 | -1.047 | 0 | 0 | 1 | 39.438 | 1 | 23 | 1 | GJB_Rinse | 0
118 | P-24301 | 2023/04/28 05:23:04 | 1-24 | 1 | -4.852 | -22.734 | 12418 | -1 | 1 | A00000170750 | 23-097 | H2O | 1682681289.68 | 0.216 | 0.81 | 63.529 | 0.001 | -0.003 | 1.57 | 42.876 | 28.81 | 2.17 | 62.829 | 0.855 | -1.056 | 0 | 0 | 1 | 39.562 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-24301 | 2023/04/28 05:32:00 | 1-24 | 2 | -4.784 | -21.601 | 12250 | -1 | 1 | A00000170750 | 23-097 | H2O | 1682681825.48 | 0.192 | 0.812 | 61.813 | 0.001 | 0 | 1.521 | 42.525 | 28.203 | 2.151 | 63.021 | 0.855 | -0.975 | 0 | 0 | 1 | 39.562 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-24301 | 2023/04/28 05:40:56 | 1-24 | 3 | -4.765 | -21.115 | 12174 | -1 | 1 | A00000170750 | 23-097 | H2O | 1682682361.73 | 0.21 | 0.754 | 68.488 | 0.001 | -0.002 | 1.693 | 42.589 | 27.981 | 2.126 | 61.26 | 0.855 | -1.282 | 0 | 0 | 1 | 39.562 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-24301 | 2023/04/28 05:49:51 | 1-24 | 4 | -4.797 | -20.976 | 12201 | 0 | 1 | A00000170750 | 23-097 | H2O | 1682682896.58 | 0.21 | 0.901 | 65.228 | 0 | -0.003 | 1.616 | 42.462 | 27.89 | 2.14 | 62.65 | 0.855 | -0.82 | 0 | 0 | 1 | 39.625 | 1 | 24 | 1 | GJB_Rinse | 0
122 | P-24302 | 2023/04/28 05:59:38 | 1-25 | 1 | -7.961 | -54.032 | 12522 | -1 | 1 | A00000170756 | 23-097 | H2O | 1682683483.9 | 0.172 | 0.778 | 64.447 | 0.001 | -0.001 | 1.583 | 43.335 | 29.252 | 2.137 | 60.963 | 0.855 | -1.384 | 0 | 0 | 1 | 39.688 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-24302 | 2023/04/28 06:08:34 | 1-25 | 2 | -8.049 | -56.718 | 12204 | -1 | 1 | A00000170756 | 23-097 | H2O | 1682684019.5 | 0.187 | 0.756 | 68.832 | 0.001 | 0.004 | 1.716 | 42.51 | 28.282 | 2.081 | 60.521 | 0.856 | -1.276 | 0 | 0 | 1 | 39.812 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-24302 | 2023/04/28 06:17:29 | 1-25 | 3 | -8.121 | -57.895 | 12260 | -1 | 1 | A00000170756 | 23-097 | H2O | 1682684554.42 | 0.184 | 0.758 | 66.519 | 0.001 | 0.001 | 1.646 | 42.69 | 28.164 | 2.09 | 60.926 | 0.855 | -1.14 | 0 | 0 | 1 | 39.812 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-24302 | 2023/04/28 06:26:25 | 1-25 | 4 | -8.063 | -57.889 | 12169 | 0 | 1 | A00000170756 | 23-097 | H2O | 1682685090.98 | 0.221 | 0.827 | 65.308 | 0.002 | 0.001 | 1.61 | 42.502 | 28.176 | 2.072 | 60.154 | 0.855 | -1.256 | 0 | 0 | 1 | 39.812 | 1 | 25 | 1 | GJB_Rinse | 0
126 | P-24303 | 2023/04/28 06:36:12 | 1-26 | 1 | -6.459 | -44.371 | 12442 | -1 | 1 | A00000322705 | 23-097 | H2O | 1682685678.26 | 0.218 | 0.761 | 62.526 | 0.001 | -0.002 | 1.548 | 43.075 | 28.74 | 2.128 | 62.135 | 0.855 | -1.186 | 0 | 0 | 1 | 39.812 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-24303 | 2023/04/28 06:45:08 | 1-26 | 2 | -6.415 | -43.079 | 12174 | -1 | 1 | A00000322705 | 23-097 | H2O | 1682686213.05 | 0.198 | 0.73 | 62.176 | 0 | 0.002 | 1.54 | 42.376 | 27.837 | 2.095 | 62.689 | 0.855 | -1.043 | 0 | 0 | 1 | 39.75 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-24303 | 2023/04/28 06:54:03 | 1-26 | 3 | -6.401 | -42.86 | 12242 | -1 | 1 | A00000322705 | 23-097 | H2O | 1682686748.41 | 0.195 | 0.812 | 64.428 | 0 | -0.001 | 1.595 | 42.579 | 27.933 | 2.115 | 62.237 | 0.855 | -1.026 | 0 | 0 | 1 | 39.75 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-24303 | 2023/04/28 07:02:59 | 1-26 | 4 | -6.382 | -42.789 | 12238 | 0 | 1 | A00000322705 | 23-097 | H2O | 1682687284.87 | 0.197 | 0.861 | 65.438 | 0.001 | 0 | 1.619 | 42.514 | 28.053 | 2.123 | 62.203 | 0.855 | -1.027 | 0 | 0 | 1 | 39.688 | 1 | 26 | 1 | GJB_Rinse | 0
130 | P-24304 | 2023/04/28 07:12:47 | 1-27 | 1 | -6.42 | -42.75 | 12404 | -1 | 1 | A00000322627 | 23-097 | H2O | 1682687872.78 | 0.222 | 0.801 | 67.106 | 0.002 | -0.003 | 1.655 | 42.711 | 28.825 | 2.137 | 62.645 | 0.855 | -0.879 | 0 | 0 | 1 | 39.688 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-24304 | 2023/04/28 07:21:43 | 1-27 | 2 | -6.429 | -42.814 | 12138 | -1 | 1 | A00000322627 | 23-097 | H2O | 1682688409.18 | 0.203 | 0.653 | 63.008 | 0.001 | 0.001 | 1.556 | 42.216 | 28.129 | 2.086 | 61.495 | 0.855 | -0.916 | 0 | 0 | 1 | 39.625 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-24304 | 2023/04/28 07:30:39 | 1-27 | 3 | -6.415 | -42.866 | 12108 | -1 | 1 | A00000322627 | 23-097 | H2O | 1682688944.45 | 0.215 | 0.837 | 64.218 | 0.002 | 0.002 | 1.598 | 42.339 | 27.742 | 2.085 | 61.053 | 0.855 | -0.978 | 0 | 0 | 1 | 39.562 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-24304 | 2023/04/28 07:39:35 | 1-27 | 4 | -6.436 | -42.933 | 12142 | 0 | 1 | A00000322627 | 23-097 | H2O | 1682689480.78 | 0.207 | 0.714 | 66.566 | 0.001 | 0.002 | 1.641 | 42.316 | 27.521 | 2.102 | 63.093 | 0.856 | -0.877 | 0 | 0 | 1 | 39.438 | 1 | 27 | 1 | GJB_Rinse | 0
134 | P-24305 | 2023/04/28 07:49:23 | 1-28 | 1 | -6.434 | -42.811 | 12212 | -1 | 1 | A00000322639 | 23-097 | H2O | 1682690069.09 | 0.191 | 0.807 | 62.451 | 0.002 | -0.001 | 1.563 | 42.512 | 27.875 | 2.113 | 61.997 | 0.855 | -1.041 | 0 | 0 | 1 | 39.5 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-24305 | 2023/04/28 07:58:18 | 1-28 | 2 | -6.472 | -42.913 | 11931 | -1 | 1 | A00000322639 | 23-097 | H2O | 1682690603.96 | 0.221 | 0.918 | 62.796 | 0.002 | 0 | 1.57 | 41.867 | 27.109 | 2.065 | 61.367 | 0.855 | -0.886 | 0 | 0 | 1 | 39.562 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-24305 | 2023/04/28 08:07:14 | 1-28 | 3 | -6.461 | -43.018 | 11933 | -1 | 1 | A00000322639 | 23-097 | H2O | 1682691139.98 | 0.203 | 0.802 | 65.646 | 0.001 | -0.004 | 1.627 | 41.942 | 27.206 | 2.058 | 60.511 | 0.855 | -0.978 | 0 | 0 | 1 | 39.562 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-24305 | 2023/04/28 08:16:11 | 1-28 | 4 | -6.382 | -42.693 | 12057 | 0 | 1 | A00000322639 | 23-097 | H2O | 1682691676.71 | 0.208 | 0.845 | 63.441 | 0.001 | 0.001 | 1.575 | 42.228 | 27.479 | 2.078 | 61.393 | 0.855 | -1.099 | 0 | 0 | 1 | 39.562 | 1 | 28 | 1 | GJB_Rinse | 0
138 | P-24306 | 2023/04/28 08:25:55 | 1-29 | 1 | -6.345 | -42.695 | 12333 | -1 | 1 | A00000322654 | 23-097 | H2O | 1682692261.42 | 0.207 | 0.824 | 62.911 | 0.001 | -0.004 | 1.555 | 42.87 | 28.449 | 2.128 | 61.724 | 0.855 | -1.164 | 0 | 0 | 1 | 39.688 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-24306 | 2023/04/28 08:34:51 | 1-29 | 2 | -6.331 | -42.581 | 12085 | -1 | 1 | A00000322654 | 23-097 | H2O | 1682692797.17 | 0.194 | 0.814 | 66.59 | 0.001 | 0.003 | 1.644 | 42.324 | 27.632 | 2.098 | 61.744 | 0.855 | -0.995 | 0 | 0 | 1 | 39.75 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-24306 | 2023/04/28 08:43:46 | 1-29 | 3 | -6.334 | -42.769 | 12127 | -1 | 1 | A00000322654 | 23-097 | H2O | 1682693332.28 | 0.196 | 0.747 | 62.64 | 0 | 0 | 1.56 | 42.499 | 27.962 | 2.094 | 60.745 | 0.855 | -1.22 | 0 | 0 | 1 | 39.75 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-24306 | 2023/04/28 08:52:42 | 1-29 | 4 | -6.338 | -42.714 | 12117 | 0 | 1 | A00000322654 | 23-097 | H2O | 1682693868.32 | 0.216 | 0.762 | 63.248 | 0.001 | -0.001 | 1.564 | 42.392 | 27.577 | 2.104 | 62.122 | 0.855 | -0.96 | 0 | 0 | 1 | 39.75 | 1 | 29 | 1 | GJB_Rinse | 0
142 | P-24307 | 2023/04/28 09:02:28 | 1-30 | 1 | -6.287 | -42.555 | 12351 | -1 | 1 | A00000322729 | 23-097 | H2O | 1682694453.59 | 0.194 | 0.834 | 67.969 | 0 | 0 | 1.683 | 42.748 | 28.303 | 2.139 | 63.215 | 0.855 | -0.887 | 0 | 0 | 1 | 39.75 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-24307 | 2023/04/28 09:11:24 | 1-30 | 2 | -6.343 | -42.739 | 12010 | -1 | 1 | A00000322729 | 23-097 | H2O | 1682694989.02 | 0.188 | 0.759 | 63.881 | 0.001 | 0.001 | 1.589 | 42.163 | 27.315 | 2.089 | 61.3 | 0.855 | -0.934 | 0 | 0 | 1 | 39.75 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-24307 | 2023/04/28 09:20:20 | 1-30 | 3 | -6.263 | -41.961 | 12131 | -1 | 1 | A00000322729 | 23-097 | H2O | 1682695525.39 | 0.214 | 0.79 | 67.04 | 0.001 | -0.001 | 1.662 | 42.385 | 27.42 | 2.111 | 63.07 | 0.855 | -0.81 | 0 | 0 | 1 | 39.688 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-24307 | 2023/04/28 09:29:15 | 1-30 | 4 | -6.248 | -42.218 | 12067 | 0 | 1 | A00000322729 | 23-097 | H2O | 1682696060.56 | 0.194 | 0.831 | 65.924 | 0.002 | -0.001 | 1.645 | 42.26 | 27.417 | 2.102 | 61.97 | 0.855 | -0.907 | 0 | 0 | 1 | 39.688 | 1 | 30 | 1 | GJB_Rinse | 0
146 | P-24308 | 2023/04/28 09:38:59 | 1-31 | 1 | -8.742 | -63.273 | 12467 | -1 | 1 | EV | standard | H2O | 1682696644.46 | 0.192 | 0.788 | 65.638 | 0.002 | 0 | 1.625 | 42.98 | 28.847 | 2.114 | 61.739 | 0.855 | -1.03 | 0 | 0 | 1 | 39.688 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-24308 | 2023/04/28 09:47:55 | 1-31 | 2 | -8.905 | -65.25 | 12189 | -1 | 1 | EV | standard | H2O | 1682697180.62 | 0.207 | 0.863 | 63.19 | 0.001 | 0.002 | 1.56 | 42.323 | 27.917 | 2.067 | 62.172 | 0.855 | -0.889 | 0 | 0 | 1 | 39.75 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-24308 | 2023/04/28 09:56:50 | 1-31 | 3 | -8.919 | -65.947 | 12210 | -1 | 1 | EV | standard | H2O | 1682697715.98 | 0.195 | 0.812 | 60.468 | 0 | 0 | 1.496 | 42.538 | 28.246 | 2.066 | 59.996 | 0.855 | -1.096 | 0 | 0 | 1 | 39.812 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-24308 | 2023/04/28 10:05:46 | 1-31 | 4 | -8.934 | -65.953 | 12235 | 0 | 1 | EV | standard | H2O | 1682698251.65 | 0.18 | 0.81 | 65.491 | 0.002 | -0.001 | 1.623 | 42.669 | 27.881 | 2.071 | 60.983 | 0.855 | -1.051 | 0 | 0 | 1 | 39.938 | 1 | 31 | 1 | GJB_Rinse | 0
150 | P-24309 | 2023/04/28 10:15:32 | 1-32 | 1 | -6.475 | -45.205 | 12167 | -1 | 1 | A00000322757 | 23-097 | H2O | 1682698837.48 | 0.213 | 0.789 | 67.819 | 0.001 | 0 | 1.687 | 42.539 | 28.197 | 2.088 | 60.314 | 0.855 | -1.294 | 0 | 0 | 1 | 40 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-24309 | 2023/04/28 10:24:28 | 1-32 | 2 | -6.455 | -43.726 | 12318 | -1 | 1 | A00000322757 | 23-097 | H2O | 1682699373.78 | 0.206 | 0.868 | 63.104 | 0.001 | -0.001 | 1.563 | 42.779 | 28.424 | 2.127 | 61.866 | 0.855 | -1.052 | 0 | 0 | 1 | 40.125 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-24309 | 2023/04/28 10:33:24 | 1-32 | 3 | -6.363 | -42.739 | 12352 | -1 | 1 | A00000322757 | 23-097 | H2O | 1682699910.08 | 0.186 | 0.818 | 65.541 | 0.002 | 0.001 | 1.62 | 42.942 | 28.625 | 2.119 | 61.203 | 0.855 | -1.179 | 0 | 0 | 1 | 40.188 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-24309 | 2023/04/28 10:42:21 | 1-32 | 4 | -6.348 | -42.5 | 12310 | 0 | 1 | A00000322757 | 23-097 | H2O | 1682700446.24 | 0.196 | 0.744 | 65.935 | 0.001 | 0.001 | 1.643 | 42.593 | 28.336 | 2.124 | 63.241 | 0.855 | -0.929 | 0 | 0 | 1 | 40.312 | 1 | 32 | 1 | GJB_Rinse | 0
154 | P-24310 | 2023/04/28 10:52:07 | 1-33 | 1 | -6.446 | -43.024 | 12387 | -1 | 1 | A00000322747 | 23-097 | H2O | 1682701032.29 | 0.205 | 0.835 | 65.866 | 0.001 | -0.003 | 1.631 | 43.009 | 28.515 | 2.133 | 61.841 | 0.855 | -1.146 | 0 | 0 | 1 | 40.312 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-24310 | 2023/04/28 11:01:03 | 1-33 | 2 | -6.426 | -43.019 | 12166 | -1 | 1 | A00000322747 | 23-097 | H2O | 1682701568.08 | 0.202 | 0.743 | 65.14 | 0.001 | -0.001 | 1.612 | 42.485 | 28.249 | 2.09 | 60.135 | 0.855 | -1.11 | 0 | 0 | 1 | 40.375 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-24310 | 2023/04/28 11:09:59 | 1-33 | 3 | -6.43 | -42.815 | 12138 | -1 | 1 | A00000322747 | 23-097 | H2O | 1682702104.21 | 0.188 | 0.692 | 62.047 | 0.001 | -0.001 | 1.545 | 42.454 | 27.885 | 2.09 | 60.763 | 0.855 | -1.116 | 0 | 0 | 1 | 40.312 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-24310 | 2023/04/28 11:18:55 | 1-33 | 4 | -6.402 | -42.953 | 12142 | 0 | 1 | A00000322747 | 23-097 | H2O | 1682702640.57 | 0.215 | 0.757 | 65.007 | 0.001 | 0 | 1.609 | 42.388 | 27.705 | 2.094 | 62.2 | 0.855 | -0.964 | 0 | 0 | 1 | 40.125 | 1 | 33 | 1 | GJB_Rinse | 0
158 | P-24311 | 2023/04/28 11:28:42 | 1-34 | 1 | -6.118 | -40.757 | 12089 | -1 | 1 | A00000322748 | 23-097 | H2O | 1682703227.58 | 0.201 | 0.766 | 60.746 | 0.001 | 0 | 1.509 | 42.161 | 27.583 | 2.094 | 62.373 | 0.855 | -0.867 | 0 | 0 | 1 | 39.75 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-24311 | 2023/04/28 11:37:37 | 1-34 | 2 | -6.044 | -40.156 | 12186 | -1 | 1 | A00000322748 | 23-097 | H2O | 1682703762.86 | 0.213 | 0.817 | 69.534 | 0.001 | -0.001 | 1.722 | 42.474 | 28.047 | 2.106 | 61.864 | 0.856 | -1.147 | 0 | 0 | 1 | 39.5 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-24311 | 2023/04/28 11:46:33 | 1-34 | 3 | -6.021 | -40.121 | 12173 | -1 | 1 | A00000322748 | 23-097 | H2O | 1682704298.28 | 0.203 | 0.877 | 66.274 | 0.001 | -0.003 | 1.65 | 42.35 | 27.394 | 2.123 | 64.049 | 0.856 | -0.702 | 0 | 0 | 1 | 39.188 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-24311 | 2023/04/28 11:55:29 | 1-34 | 4 | -6.012 | -39.964 | 12024 | 0 | 1 | A00000322748 | 23-097 | H2O | 1682704834.49 | 0.202 | 0.672 | 63.228 | 0.001 | 0 | 1.576 | 42.199 | 27.086 | 2.082 | 62.66 | 0.856 | -0.943 | 0 | 0 | 1 | 38.938 | 1 | 34 | 1 | GJB_Rinse | 0
162 | P-24312 | 2023/04/28 12:05:15 | 1-35 | 1 | -6.33 | -42.008 | 12458 | -1 | 1 | A00000322749 | 23-097 | H2O | 1682705421.08 | 0.186 | 0.791 | 67.731 | 0.001 | -0.001 | 1.687 | 43.13 | 28.762 | 2.139 | 62.662 | 0.855 | -1.141 | 0 | 0 | 1 | 38.812 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-24312 | 2023/04/28 12:14:11 | 1-35 | 2 | -6.457 | -42.65 | 12142 | -1 | 1 | A00000322749 | 23-097 | H2O | 1682705957.01 | 0.197 | 0.833 | 68.685 | 0.002 | 0.001 | 1.707 | 42.258 | 27.86 | 2.092 | 62.254 | 0.855 | -0.853 | 0 | 0 | 1 | 38.75 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-24312 | 2023/04/28 12:23:07 | 1-35 | 3 | -6.384 | -42.109 | 12164 | -1 | 1 | A00000322749 | 23-097 | H2O | 1682706492.74 | 0.214 | 0.731 | 63.966 | 0.001 | 0.002 | 1.588 | 42.522 | 27.8 | 2.101 | 61.537 | 0.856 | -1.16 | 0 | 0 | 1 | 38.812 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-24312 | 2023/04/28 12:32:03 | 1-35 | 4 | -6.393 | -42.352 | 12110 | 0 | 1 | A00000322749 | 23-097 | H2O | 1682707028.89 | 0.195 | 0.799 | 63.9 | 0.001 | -0.001 | 1.591 | 42.425 | 27.662 | 2.082 | 61.043 | 0.855 | -1.087 | 0 | 0 | 1 | 38.812 | 1 | 35 | 1 | GJB_Rinse | 0
166 | P-24313 | 2023/04/28 12:41:46 | 1-36 | 1 | -6.495 | -43.437 | 12501 | -1 | 1 | A00000322746 | 23-097 | H2O | 1682707612.23 | 0.183 | 0.808 | 63.94 | 0.001 | 0 | 1.577 | 43.233 | 28.652 | 2.149 | 62.806 | 0.856 | -1.155 | 0 | 0 | 1 | 38.75 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-24313 | 2023/04/28 12:50:41 | 1-36 | 2 | -6.532 | -43.643 | 12136 | -1 | 1 | A00000322746 | 23-097 | H2O | 1682708146.57 | 0.205 | 0.895 | 59.335 | 0.001 | 0 | 1.487 | 42.394 | 27.763 | 2.087 | 61.523 | 0.856 | -1.075 | 0 | 0 | 1 | 38.75 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-24313 | 2023/04/28 12:59:36 | 1-36 | 3 | -6.559 | -43.464 | 12160 | -1 | 1 | A00000322746 | 23-097 | H2O | 1682708682.42 | 0.222 | 0.888 | 64.461 | 0.002 | 0.001 | 1.595 | 42.432 | 27.549 | 2.101 | 62.623 | 0.855 | -0.85 | 0 | 0 | 1 | 38.75 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-24313 | 2023/04/28 13:08:32 | 1-36 | 4 | -6.567 | -43.894 | 12182 | 0 | 1 | A00000322746 | 23-097 | H2O | 1682709218.21 | 0.182 | 0.868 | 67.469 | 0.001 | 0.004 | 1.663 | 42.468 | 28.07 | 2.11 | 61.718 | 0.856 | -1.047 | 0 | 0 | 1 | 38.75 | 1 | 36 | 1 | GJB_Rinse | 0
170 | P-24314 | 2023/04/28 13:18:17 | 1-37 | 1 | -6.363 | -41.29 | 12378 | -1 | 1 | A00000322687 | 23-097 | H2O | 1682709802.35 | 0.208 | 0.845 | 64.856 | 0.001 | 0.001 | 1.615 | 42.689 | 28.594 | 2.141 | 62.404 | 0.856 | -0.97 | 0 | 0 | 1 | 38.812 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-24314 | 2023/04/28 13:27:13 | 1-37 | 2 | -6.344 | -40.871 | 12113 | -1 | 1 | A00000322687 | 23-097 | H2O | 1682710338.9 | 0.206 | 0.737 | 61.899 | 0.002 | 0.001 | 1.53 | 42.397 | 28.009 | 2.094 | 60.034 | 0.855 | -1.276 | 0 | 0 | 1 | 38.75 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-24314 | 2023/04/28 13:36:10 | 1-37 | 3 | -6.372 | -41.044 | 12170 | -1 | 1 | A00000322687 | 23-097 | H2O | 1682710875.46 | 0.214 | 0.806 | 67.563 | 0.001 | 0.003 | 1.679 | 42.6 | 27.937 | 2.09 | 60.942 | 0.856 | -1.176 | 0 | 0 | 1 | 38.75 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-24314 | 2023/04/28 13:45:05 | 1-37 | 4 | -6.391 | -40.925 | 12044 | 0 | 1 | A00000322687 | 23-097 | H2O | 1682711410.98 | 0.198 | 0.932 | 63.326 | 0.002 | 0.001 | 1.574 | 42.281 | 27.463 | 2.081 | 61.043 | 0.856 | -1.07 | 0 | 0 | 1 | 38.688 | 1 | 37 | 1 | GJB_Rinse | 0
174 | P-24315 | 2023/04/28 13:54:50 | 1-38 | 1 | -6.53 | -41.836 | 12409 | -1 | 1 | A00000322622 | 23-097 | H2O | 1682711995.17 | 0.213 | 0.864 | 65.853 | 0.001 | -0.001 | 1.626 | 42.891 | 28.598 | 2.142 | 62.971 | 0.856 | -1.041 | 0 | 0 | 1 | 38.75 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-24315 | 2023/04/28 14:03:45 | 1-38 | 2 | -6.54 | -41.882 | 12231 | -1 | 1 | A00000322622 | 23-097 | H2O | 1682712530.58 | 0.207 | 0.876 | 63.588 | 0 | -0.006 | 1.579 | 42.621 | 27.966 | 2.107 | 61.86 | 0.856 | -1.024 | 0 | 0 | 1 | 38.625 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-24315 | 2023/04/28 14:12:41 | 1-38 | 3 | -6.562 | -41.98 | 12056 | -1 | 1 | A00000322622 | 23-097 | H2O | 1682713066.8 | 0.194 | 0.819 | 63.112 | 0.001 | 0.002 | 1.569 | 42.268 | 27.476 | 2.072 | 60.841 | 0.856 | -1.048 | 0 | 0 | 1 | 38.625 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-24315 | 2023/04/28 14:21:37 | 1-38 | 4 | -6.618 | -42.393 | 12098 | 0 | 1 | A00000322622 | 23-097 | H2O | 1682713602.44 | 0.212 | 0.905 | 62.222 | 0.001 | -0.003 | 1.534 | 42.25 | 27.623 | 2.083 | 61.775 | 0.856 | -0.947 | 0 | 0 | 1 | 38.625 | 1 | 38 | 1 | GJB_Rinse | 0
178 | P-24316 | 2023/04/28 14:31:21 | 1-39 | 1 | 10.42 | -43.608 | 11988 | -1 | 1 | WOOD4-D-5 | 23-074 | H2O | 1682714186.83 | 0.238 | 1.354 | 61.686 | -0.002 | 0.004 | 1.534 | 392.433 | -0.847 | 4.855 | 64.23 | 0.857 | -9.414 | 0 | 0 | 1 | 38.688 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-24316 | 2023/04/28 14:40:17 | 1-39 | 2 | 10.522 | -46.133 | 11943 | -1 | 1 | WOOD4-D-5 | 23-074 | H2O | 1682714723.21 | 0.239 | 1.496 | 66.643 | -0.001 | -0.002 | 1.658 | 394.005 | -1.287 | 4.87 | 63.784 | 0.857 | -9.584 | 0 | 0 | 1 | 38.688 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-24316 | 2023/04/28 14:49:14 | 1-39 | 3 | 10.614 | -46.672 | 11893 | -1 | 1 | WOOD4-D-5 | 23-074 | H2O | 1682715259.46 | 0.27 | 1.569 | 63.803 | -0.002 | 0.007 | 1.593 | 393.79 | -1.021 | 4.856 | 63.255 | 0.856 | -9.761 | 0 | 0 | 1 | 38.688 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-24316 | 2023/04/28 14:58:10 | 1-39 | 4 | 10.575 | -46.702 | 11909 | 0 | 1 | WOOD4-D-5 | 23-074 | H2O | 1682715795.71 | 0.23 | 1.534 | 62.021 | -0.001 | -0.006 | 1.54 | 394.438 | -1.21 | 4.874 | 63.151 | 0.857 | -9.639 | 0 | 0 | 1 | 38.625 | 1 | 39 | 1 | GJB_Rinse | 0
182 | P-24317 | 2023/04/28 15:07:55 | 1-40 | 1 | -8.538 | -66.286 | 12441 | -1 | 1 | EV | standard | H2O | 1682716380.77 | 0.205 | 0.833 | 65.42 | 0.002 | 0.002 | 1.616 | 47.786 | 28.17 | 2.13 | 62.592 | 0.856 | -1.112 | 0 | 0 | 1 | 38.625 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-24317 | 2023/04/28 15:16:51 | 1-40 | 2 | -8.781 | -66.186 | 12142 | -1 | 1 | EV | standard | H2O | 1682716916.24 | 0.19 | 0.812 | 59.499 | 0.001 | 0.002 | 1.479 | 44.297 | 27.469 | 2.072 | 60.711 | 0.856 | -1.001 | 0 | 0 | 1 | 38.625 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-24317 | 2023/04/28 15:25:47 | 1-40 | 3 | -8.786 | -66.374 | 12154 | -1 | 1 | EV | standard | H2O | 1682717452.29 | 0.194 | 0.861 | 59.761 | 0.001 | -0.001 | 1.485 | 43.76 | 27.754 | 2.055 | 60.432 | 0.856 | -1.156 | 0 | 0 | 1 | 38.688 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-24317 | 2023/04/28 15:34:41 | 1-40 | 4 | -8.83 | -66.311 | 12104 | 0 | 1 | EV | standard | H2O | 1682717986.82 | 0.216 | 0.928 | 66.567 | 0.002 | 0 | 1.656 | 43.189 | 27.444 | 2.066 | 60.934 | 0.856 | -0.886 | 0 | 0 | 1 | 38.688 | 1 | 40 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 41.9063333333
varray: [43.879, 44.686, 44.58, 44.506, 44.619, 44.432, 43.847, 44.006, 43.963, 43.993, 0.0, 44.551, 44.184, 43.91, 43.773, 43.352, 43.986, 43.772, 43.288, 43.252, 43.454]
except: True
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_shift
baseline_shift_STD | 9.61211166359
varray: [43.879, 44.686, 44.58, 44.506, 44.619, 44.432, 43.847, 44.006, 43.963, 43.993, 0.0, 44.551, 44.184, 43.91, 43.773, 43.352, 43.986, 43.772, 43.288, 43.252, 43.454]
except: True
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.984142857143
varray: [-1.129, -1.22, -1.22, -1.207, -1.221, -1.116, -1.179, -1.043, -0.938, -1.018, 0.0, -0.944, -1.055, -0.815, -0.934, -0.981, -0.977, -0.809, -0.927, -0.976, -0.958]
except: True
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.258954298229
varray: [-1.129, -1.22, -1.22, -1.207, -1.221, -1.116, -1.179, -1.043, -0.938, -1.018, 0.0, -0.944, -1.055, -0.815, -0.934, -0.981, -0.977, -0.809, -0.927, -0.976, -0.958]
except: True
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: ch4_ppm
residuals_MEAN | 2.08023809524
varray: [2.261, 2.27, 2.264, 2.269, 2.285, 2.289, 2.263, 2.292, 2.311, 2.29, 0.0, 2.175, 2.102, 2.099, 2.078, 2.056, 2.103, 2.106, 2.043, 2.061, 2.068]
except: True
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: residuals
residuals_STD | 0.486824701999
varray: [2.261, 2.27, 2.264, 2.269, 2.285, 2.289, 2.263, 2.292, 2.311, 2.29, 0.0, 2.175, 2.102, 2.099, 2.078, 2.056, 2.103, 2.106, 2.043, 2.061, 2.068]
except: True
inst: std_standards
operation: std_standards(residuals)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: residuals
slope_shift_MEAN | 27.9004761905
varray: [29.248, 29.653, 29.044, 29.698, 29.894, 29.616, 29.293, 29.217, 29.574, 28.997, 0.0, 30.533, 29.714, 29.054, 29.251, 28.477, 29.538, 29.18, 28.966, 28.357, 28.606]
except: True
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: slope_shift
slope_shift_STD | 6.41218051539
varray: [29.248, 29.653, 29.044, 29.698, 29.894, 29.616, 29.293, 29.217, 29.574, 28.997, 0.0, 30.533, 29.714, 29.054, 29.251, 28.477, 29.538, 29.18, 28.966, 28.357, 28.606]
except: True
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: slope_shift
baseline_curvature_MEAN | 60.8590952381
varray: [64.794, 65.08, 65.073, 64.762, 65.17, 66.259, 64.668, 65.926, 67.13, 65.718, 0.0, 64.251, 61.997, 63.807, 61.923, 61.316, 62.473, 63.38, 60.996, 61.707, 61.611]
except: True
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_curvature
baseline_curvature_STD | 14.0607064826
varray: [64.794, 65.08, 65.073, 64.762, 65.17, 66.259, 64.668, 65.926, 67.13, 65.718, 0.0, 64.251, 61.997, 63.807, 61.923, 61.316, 62.473, 63.38, 60.996, 61.707, 61.611]
except: True
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.08023809524
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 41.9063333333
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.984142857143
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 27.9004761905
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.730237052998
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.81047514824
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.973649403998
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 3.05388749924
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 1.460474106
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 3.54071220123
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.208023809524
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.776862894688
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.76100575183
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 4.8060558318
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 46.7123891651
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 51.5184449969
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 28.8363349908
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 70.7426683241
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 173.018009945
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 214.924343278
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 32.0609025769
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 59.9613787674
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 42.1821194479
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 103.041214686
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -8.65152418615
actual_array: [ 1.93 -15.84]
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822, 2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996, 0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606, -9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915, -6.794, -6.688, -6.686, -6.632, -6.811, -6.821, -6.766, -6.83, -6.515, -6.466, -6.404, -6.407, -6.435, -6.494, -6.453, -6.454, -6.676, -6.63, -6.653, -6.637, -6.208, -6.182, -6.097, -6.083, -6.805, -6.753, -6.734, -6.761, -6.887, -6.871, -6.894, -6.958, -8.821, -8.979, -8.943, -8.975, -4.618, -4.327, -4.278, -4.23, -7.433, -7.619, -7.544, -7.585, -13.382, -13.694, -13.738, -13.754, -3.782, -3.301, -3.213, -3.149, -5.712, -5.833, -5.828, -5.858, -7.643, -7.762, -7.724, -7.817, -14.314, -14.624, -14.737, -14.741, -5.007, -4.505, -4.443, -4.43, -8.717, -8.917, -8.947, -8.964, -5.704, -5.56, -5.505, -5.503, -4.852, -4.784, -4.765, -4.797, -7.961, -8.049, -8.121, -8.063, -6.459, -6.415, -6.401, -6.382, -6.42, -6.429, -6.415, -6.436, -6.434, -6.472, -6.461, -6.382, -6.345, -6.331, -6.334, -6.338, -6.287, -6.343, -6.263, -6.248, -8.742, -8.905, -8.919, -8.934, -6.475, -6.455, -6.363, -6.348, -6.446, -6.426, -6.43, -6.402, -6.118, -6.044, -6.021, -6.012, -6.33, -6.457, -6.384, -6.393, -6.495, -6.532, -6.559, -6.567, -6.363, -6.344, -6.372, -6.391, -6.53, -6.54, -6.562, -6.618, 10.42, 10.522, 10.614, 10.575, -8.538, -8.781, -8.786, -8.83]
fit_rsq: 1.0
fit_intercept: 1.0849564946
except: True
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.66305945946
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0, 5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0, 16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0, 22245.0, 22781.0, 23318.0, 23854.0, 24445.0, 24982.0, 25518.0, 26054.0, 26644.0, 27180.0, 27716.0, 28252.0, 28840.0, 29377.0, 29914.0, 30450.0, 31038.0, 31574.0, 32110.0, 32646.0, 33234.0, 33769.0, 34306.0, 34843.0, 35432.0, 35969.0, 36505.0, 37042.0, 37632.0, 38168.0, 38705.0, 39241.0, 39831.0, 40367.0, 40903.0, 41440.0, 42030.0, 42566.0, 43103.0, 43640.0, 44227.0, 44762.0, 45299.0, 45834.0, 46422.0, 46958.0, 47494.0, 48031.0, 48620.0, 49156.0, 49691.0, 50226.0, 50815.0, 51350.0, 51887.0, 52423.0, 53011.0, 53548.0, 54084.0, 54619.0, 55208.0, 55742.0, 56279.0, 56814.0, 57403.0, 57939.0, 58476.0, 59012.0, 59597.0, 60134.0, 60670.0, 61206.0, 61792.0, 62328.0, 62864.0, 63400.0, 63987.0, 64523.0, 65059.0, 65594.0, 66181.0, 66717.0, 67252.0, 67788.0, 68375.0, 68911.0, 69446.0, 69982.0, 70570.0, 71106.0, 71642.0, 72178.0, 72766.0, 73301.0, 73837.0, 74374.0, 74958.0, 75494.0, 76029.0, 76565.0, 77151.0, 77687.0, 78223.0, 78758.0, 79342.0, 79878.0, 80413.0, 80949.0, 81535.0, 82071.0, 82607.0, 83144.0, 83730.0, 84266.0, 84802.0, 85338.0, 85925.0, 86460.0, 86996.0, 87532.0, 88118.0, 88654.0, 89190.0, 89726.0, 90309.0, 90844.0, 91379.0, 91915.0, 92500.0, 93036.0, 93573.0, 94108.0, 94693.0, 95228.0, 95764.0, 96300.0, 96884.0, 97420.0, 97957.0, 98493.0, 99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -60.9479103344
actual_array: [ 18.1 -119. ]
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487, 7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712, 0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978, -72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075, -46.184, -44.054, -43.638, -43.264, -44.26, -44.554, -44.307, -44.231, -42.744, -42.36, -41.913, -41.99, -41.317, -41.635, -41.345, -41.494, -43.023, -42.817, -42.643, -42.863, -39.934, -39.303, -38.8, -38.936, -43.551, -43.592, -43.597, -43.677, -44.734, -44.909, -44.953, -45.043, -63.839, -65.983, -65.997, -66.174, -38.126, -35.175, -34.314, -33.784, -51.963, -54.002, -54.07, -54.167, -105.572, -110.597, -111.919, -112.8, -30.628, -22.916, -20.909, -20.053, -38.15, -40.129, -40.126, -40.227, -48.355, -49.481, -49.511, -49.868, -105.474, -111.152, -112.662, -113.445, -39.298, -32.211, -30.494, -29.799, -62.306, -65.497, -66.14, -66.517, -37.664, -35.224, -34.058, -33.739, -22.734, -21.601, -21.115, -20.976, -54.032, -56.718, -57.895, -57.889, -44.371, -43.079, -42.86, -42.789, -42.75, -42.814, -42.866, -42.933, -42.811, -42.913, -43.018, -42.693, -42.695, -42.581, -42.769, -42.714, -42.555, -42.739, -41.961, -42.218, -63.273, -65.25, -65.947, -65.953, -45.205, -43.726, -42.739, -42.5, -43.024, -43.019, -42.815, -42.953, -40.757, -40.156, -40.121, -39.964, -42.008, -42.65, -42.109, -42.352, -43.437, -43.643, -43.464, -43.894, -41.29, -40.871, -41.044, -40.925, -41.836, -41.882, -41.98, -42.393, -43.608, -46.133, -46.672, -46.702, -66.286, -66.186, -66.374, -66.311]
fit_rsq: 1.0
fit_intercept: 1.93644088588
except: True
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -49.7368108108
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0, 5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0, 10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0, 16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0, 22245.0, 22781.0, 23318.0, 23854.0, 24445.0, 24982.0, 25518.0, 26054.0, 26644.0, 27180.0, 27716.0, 28252.0, 28840.0, 29377.0, 29914.0, 30450.0, 31038.0, 31574.0, 32110.0, 32646.0, 33234.0, 33769.0, 34306.0, 34843.0, 35432.0, 35969.0, 36505.0, 37042.0, 37632.0, 38168.0, 38705.0, 39241.0, 39831.0, 40367.0, 40903.0, 41440.0, 42030.0, 42566.0, 43103.0, 43640.0, 44227.0, 44762.0, 45299.0, 45834.0, 46422.0, 46958.0, 47494.0, 48031.0, 48620.0, 49156.0, 49691.0, 50226.0, 50815.0, 51350.0, 51887.0, 52423.0, 53011.0, 53548.0, 54084.0, 54619.0, 55208.0, 55742.0, 56279.0, 56814.0, 57403.0, 57939.0, 58476.0, 59012.0, 59597.0, 60134.0, 60670.0, 61206.0, 61792.0, 62328.0, 62864.0, 63400.0, 63987.0, 64523.0, 65059.0, 65594.0, 66181.0, 66717.0, 67252.0, 67788.0, 68375.0, 68911.0, 69446.0, 69982.0, 70570.0, 71106.0, 71642.0, 72178.0, 72766.0, 73301.0, 73837.0, 74374.0, 74958.0, 75494.0, 76029.0, 76565.0, 77151.0, 77687.0, 78223.0, 78758.0, 79342.0, 79878.0, 80413.0, 80949.0, 81535.0, 82071.0, 82607.0, 83144.0, 83730.0, 84266.0, 84802.0, 85338.0, 85925.0, 86460.0, 86996.0, 87532.0, 88118.0, 88654.0, 89190.0, 89726.0, 90309.0, 90844.0, 91379.0, 91915.0, 92500.0, 93036.0, 93573.0, 94108.0, 94693.0, 95228.0, 95764.0, 96300.0, 96884.0, 97420.0, 97957.0, 98493.0, 99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2095
varray: [2.697, 2.19, 2.154, 2.16, 2.149, 2.141, 2.151, 2.128, 2.159, 2.166]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: residuals
SAMP_residuals_STD | 0.172055062246
varray: [2.697, 2.19, 2.154, 2.16, 2.149, 2.141, 2.151, 2.128, 2.159, 2.166]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1229
varray: [-1.404, -1.085, -1.096, -1.003, -1.083, -1.111, -1.223, -1.254, -1.119, -0.851]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.148488271142
varray: [-1.404, -1.085, -1.096, -1.003, -1.083, -1.111, -1.223, -1.254, -1.119, -0.851]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 51.8671
varray: [78.99, 52.747, 51.443, 51.412, 50.174, 47.591, 47.434, 46.755, 46.355, 45.77]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 9.83931084364
varray: [78.99, 52.747, 51.443, 51.412, 50.174, 47.591, 47.434, 46.755, 46.355, 45.77]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.1971
varray: [38.375, 29.663, 29.482, 29.018, 28.907, 29.089, 29.635, 29.118, 29.335, 29.349]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.88482818475
varray: [38.375, 29.663, 29.482, 29.018, 28.907, 29.089, 29.635, 29.118, 29.335, 29.349]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: slope_shift
SAMP_d18O_STD | 0.193000748415
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.38582217136
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.6101
varray: [77.371, 63.328, 62.559, 64.449, 63.304, 62.779, 62.447, 61.824, 63.249, 64.791]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.57168593871
varray: [77.371, 63.328, 62.559, 64.449, 63.304, 62.779, 62.447, 61.824, 63.249, 64.791]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2095
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1229
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 51.8671
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.1971
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.129261904762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.129261904762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.213803365
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 3.7510047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9232
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
dD_mean | -68.8389
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
d18O_stdev | 0.193000748415
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
dD_stdev | 2.38582217136
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9232
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.8389
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.193000748415
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.38582217136
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.1752232423
actual_array: [ 1.93 -15.84]
varray: [-9.435, -9.045, -8.902, -8.865, -8.872, -8.853, -8.821, -8.802, -8.815, -8.822]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.9232
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -83.4785694448
actual_array: [ 18.1 -119. ]
varray: [-75.152, -70.259, -68.826, -68.381, -68.147, -67.671, -67.553, -67.485, -67.428, -67.487]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -68.8389
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [1.0, 537.0, 1074.0, 1609.0, 2146.0, 2682.0, 3218.0, 3754.0, 4290.0, 4827.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.6968333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 9.96076666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2794
varray: [2.261, 2.27, 2.264, 2.269, 2.285, 2.289, 2.263, 2.292, 2.311, 2.29]
inst: mean_all
operation: mean_all(residuals)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: residuals
SAMP_residuals_STD | 0.0164465464122
varray: [2.261, 2.27, 2.264, 2.269, 2.285, 2.289, 2.263, 2.292, 2.311, 2.29]
inst: std_all
operation: std_all(residuals)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1291
varray: [-1.129, -1.22, -1.22, -1.207, -1.221, -1.116, -1.179, -1.043, -0.938, -1.018]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0999871658431
varray: [-1.129, -1.22, -1.22, -1.207, -1.221, -1.116, -1.179, -1.043, -0.938, -1.018]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 44.2511
varray: [43.879, 44.686, 44.58, 44.506, 44.619, 44.432, 43.847, 44.006, 43.963, 43.993]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.340230134533
varray: [43.879, 44.686, 44.58, 44.506, 44.619, 44.432, 43.847, 44.006, 43.963, 43.993]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.4234
varray: [29.248, 29.653, 29.044, 29.698, 29.894, 29.616, 29.293, 29.217, 29.574, 28.997]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.302855374358
varray: [29.248, 29.653, 29.044, 29.698, 29.894, 29.616, 29.293, 29.217, 29.574, 28.997]
inst: std_all
operation: std_all(slope_shift)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: slope_shift
SAMP_d18O_STD | 0.259387397963
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.72903432391
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.458
varray: [64.794, 65.08, 65.073, 64.762, 65.17, 66.259, 64.668, 65.926, 67.13, 65.718]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.792133125736
varray: [64.794, 65.08, 65.073, 64.762, 65.17, 66.259, 64.668, 65.926, 67.13, 65.718]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2794
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1291
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 44.2511
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.4234
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.199161904762
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.199161904762
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 9.57399565068
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 4.5989047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.8134
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.2821
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.8134
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
dD_mean | 17.2821
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
d18O_stdev | 0.259387397963
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
dD_stdev | 3.72903432391
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.8134
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.2821
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.259387397963
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.72903432391
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.13, 2.667, 2.79, 2.854, 2.91, 2.953, 2.943, 2.923, 2.968, 2.996]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.8134
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [7.531, 15.006, 16.88, 18.014, 18.437, 18.769, 19.42, 19.303, 19.749, 19.712]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 17.2821
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [5416.0, 5952.0, 6488.0, 7025.0, 7561.0, 8099.0, 8634.0, 9171.0, 9708.0, 10245.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.6269333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.34476666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.89918181818
varray: [0.0, 2.175, 2.102, 2.099, 2.078, 2.056, 2.103, 2.106, 2.043, 2.061, 2.068]
except: True
inst: mean_all
operation: mean_all(residuals)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: residuals
SAMP_residuals_STD | 0.630894732611
varray: [0.0, 2.175, 2.102, 2.099, 2.078, 2.056, 2.103, 2.106, 2.043, 2.061, 2.068]
except: True
inst: std_all
operation: std_all(residuals)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.852363636364
varray: [0.0, -0.944, -1.055, -0.815, -0.934, -0.981, -0.977, -0.809, -0.927, -0.976, -0.958]
except: True
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.291551118237
varray: [0.0, -0.944, -1.055, -0.815, -0.934, -0.981, -0.977, -0.809, -0.927, -0.976, -0.958]
except: True
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.7747272727
varray: [0.0, 44.551, 44.184, 43.91, 43.773, 43.352, 43.986, 43.772, 43.288, 43.252, 43.454]
except: True
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 13.1979222159
varray: [0.0, 44.551, 44.184, 43.91, 43.773, 43.352, 43.986, 43.772, 43.288, 43.252, 43.454]
except: True
inst: std_all
operation: std_all(baseline_shift)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.516
varray: [0.0, 30.533, 29.714, 29.054, 29.251, 28.477, 29.538, 29.18, 28.966, 28.357, 28.606]
except: True
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: slope_shift
SAMP_slope_shift_STD | 8.81599795826
varray: [0.0, 30.533, 29.714, 29.054, 29.251, 28.477, 29.538, 29.18, 28.966, 28.357, 28.606]
except: True
inst: std_all
operation: std_all(slope_shift)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: slope_shift
SAMP_d18O_STD | 4.35323339182
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 33.1306864989
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 56.6782727273
varray: [0.0, 64.251, 61.997, 63.807, 61.923, 61.316, 62.473, 63.38, 60.996, 61.707, 61.611]
except: True
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 18.8272429372
varray: [0.0, 64.251, 61.997, 63.807, 61.923, 61.316, 62.473, 63.38, 60.996, 61.707, 61.611]
except: True
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.89918181818
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.852363636364
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.7747272727
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.516
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.181056277056
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.181056277056
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.70363240971
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.18082251082
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.1008181818
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
dD_mean | -98.9549090909
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
d18O_stdev | 4.35323339182
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
dD_stdev | 33.1306864989
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.1008181818
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
measured_dD_mean | -98.9549090909
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 4.35323339182
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
except: True
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
measured_dD_stdev | 33.1306864989
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
except: True
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [0.0, -13.658, -14.294, -14.392, -14.493, -14.449, -14.589, -14.563, -14.521, -14.544, -14.606]
fit_rsq: 1.0
fit_intercept: 1.0849564946
except: True
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.1008181818
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [0.0, -96.212, -106.125, -108.58, -109.969, -110.298, -111.161, -111.23, -111.477, -111.474, -111.978]
fit_rsq: 1.0
fit_intercept: 1.93644088588
except: True
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -98.9549090909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [10810.0, 11409.0, 11946.0, 12484.0, 13020.0, 13558.0, 14094.0, 14630.0, 15167.0, 15703.0, 16239.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0071515152
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.13160606061
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.113
varray: [2.13, 2.133, 2.131, 2.129, 2.121, 2.128, 2.088, 2.08, 2.09, 2.1]
inst: mean_all
operation: mean_all(residuals)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: residuals
SAMP_residuals_STD | 0.021002645336
varray: [2.13, 2.133, 2.131, 2.129, 2.121, 2.128, 2.088, 2.08, 2.09, 2.1]
inst: std_all
operation: std_all(residuals)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9525
varray: [-0.97, -0.97, -0.872, -0.87, -0.888, -1.01, -0.878, -0.956, -1.051, -1.06]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0731714121468
varray: [-0.97, -0.97, -0.872, -0.87, -0.888, -1.01, -0.878, -0.956, -1.051, -1.06]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.9799
varray: [43.298, 43.394, 43.212, 43.13, 43.153, 43.081, 42.588, 42.59, 42.658, 42.695]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.312631643383
varray: [43.298, 43.394, 43.212, 43.13, 43.153, 43.081, 42.588, 42.59, 42.658, 42.695]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3331
varray: [29.086, 28.913, 28.64, 28.195, 28.09, 28.45, 27.854, 28.093, 27.908, 28.102]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.423013908361
varray: [29.086, 28.913, 28.64, 28.195, 28.09, 28.45, 27.854, 28.093, 27.908, 28.102]
inst: std_all
operation: std_all(slope_shift)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: slope_shift
SAMP_d18O_STD | 0.108666871165
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.65051116998
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.6603
varray: [62.942, 63.373, 63.769, 63.303, 63.324, 62.459, 62.142, 61.33, 62.357, 61.604]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.81354424178
varray: [62.942, 63.373, 63.769, 63.303, 63.324, 62.459, 62.142, 61.33, 62.357, 61.604]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.113
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.9799
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3331
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0327619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0327619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.57491129678
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8012047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9546
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
dD_mean | -68.0517
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
d18O_stdev | 0.108666871165
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
dD_stdev | 1.65051116998
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9546
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.0517
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.108666871165
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.65051116998
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.2102848453
actual_array: [ 1.93 -15.84]
varray: [-9.247, -9.0, -8.869, -8.934, -8.89, -8.932, -8.904, -8.946, -8.909, -8.915]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.9546
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.5500775571
actual_array: [ 18.1 -119. ]
varray: [-72.426, -69.15, -67.662, -67.707, -67.432, -67.399, -67.133, -67.447, -67.086, -67.075]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -68.0517
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [16828.0, 17365.0, 17902.0, 18439.0, 18975.0, 19511.0, 20047.0, 20583.0, 21119.0, 21655.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7933333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.07356666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15275
varray: [2.168, 2.136, 2.166, 2.141]
inst: mean_all
operation: mean_all(residuals)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: residuals
SAMP_residuals_STD | 0.0166007027964
varray: [2.168, 2.136, 2.166, 2.141]
inst: std_all
operation: std_all(residuals)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.002
varray: [-0.98, -0.978, -1.052, -0.998]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0345253530033
varray: [-0.98, -0.978, -1.052, -0.998]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.96875
varray: [43.055, 42.686, 43.216, 42.918]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.22442129281
varray: [43.055, 42.686, 43.216, 42.918]
inst: std_all
operation: std_all(baseline_shift)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4085
varray: [28.665, 28.051, 28.593, 28.325]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.279649184992
varray: [28.665, 28.051, 28.593, 28.325]
inst: std_all
operation: std_all(slope_shift)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0678232998313
varray: [-6.794, -6.688, -6.686, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.30647260464
varray: [-46.184, -44.054, -43.638, -43.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.19675
varray: [63.38, 62.699, 63.75, 62.958]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.463473390103
varray: [63.38, 62.699, 63.75, 62.958]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.002
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.96875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0725119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0725119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.48575025753
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.3376547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.794, -6.688, -6.686, -6.632]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-46.184, -44.054, -43.638, -43.264]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.7
varray: [-6.794, -6.688, -6.686, -6.632]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
dD_mean | -44.285
varray: [-46.184, -44.054, -43.638, -43.264]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0678232998313
varray: [-6.794, -6.688, -6.686, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
dD_stdev | 1.30647260464
varray: [-46.184, -44.054, -43.638, -43.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.7
varray: [-6.794, -6.688, -6.686, -6.632]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.285
varray: [-46.184, -44.054, -43.638, -43.264]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0678232998313
varray: [-6.794, -6.688, -6.686, -6.632]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.30647260464
varray: [-46.184, -44.054, -43.638, -43.264]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.69277242077
actual_array: [ 1.93 -15.84]
varray: [-6.794, -6.688, -6.686, -6.632]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.7
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.5175722863
actual_array: [ 18.1 -119. ]
varray: [-46.184, -44.054, -43.638, -43.264]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.285
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [22245.0, 22781.0, 23318.0, 23854.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7535833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.06241666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.17325
varray: [2.157, 2.147, 2.201, 2.188]
inst: mean_all
operation: mean_all(residuals)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: residuals
SAMP_residuals_STD | 0.0254345565455
varray: [2.157, 2.147, 2.201, 2.188]
inst: std_all
operation: std_all(residuals)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.96075
varray: [-0.865, -0.807, -1.133, -1.038]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.151046074648
varray: [-0.865, -0.807, -1.133, -1.038]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.192
varray: [43.036, 42.705, 43.771, 43.256]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.44752728036
varray: [43.036, 42.705, 43.771, 43.256]
inst: std_all
operation: std_all(baseline_shift)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.7795
varray: [28.437, 28.142, 29.651, 28.888]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.657012683794
varray: [28.437, 28.142, 29.651, 28.888]
inst: std_all
operation: std_all(slope_shift)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0284136117615
varray: [-6.811, -6.821, -6.766, -6.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.147365758121
varray: [-44.26, -44.554, -44.307, -44.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.779
varray: [64.007, 63.299, 63.698, 64.112]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.36507259552
varray: [64.007, 63.299, 63.698, 64.112]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.17325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.96075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.192
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.7795
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0930119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0930119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.47121437564
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.9199047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.811, -6.821, -6.766, -6.83]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-44.26, -44.554, -44.307, -44.231]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.807
varray: [-6.811, -6.821, -6.766, -6.83]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
dD_mean | -44.338
varray: [-44.26, -44.554, -44.307, -44.231]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0284136117615
varray: [-6.811, -6.821, -6.766, -6.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
dD_stdev | 0.147365758121
varray: [-44.26, -44.554, -44.307, -44.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.807
varray: [-6.811, -6.821, -6.766, -6.83]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.338
varray: [-44.26, -44.554, -44.307, -44.231]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0284136117615
varray: [-6.811, -6.821, -6.766, -6.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.147365758121
varray: [-44.26, -44.554, -44.307, -44.231]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.81224985776
actual_array: [ 1.93 -15.84]
varray: [-6.811, -6.821, -6.766, -6.83]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.807
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.5800850785
actual_array: [ 18.1 -119. ]
varray: [-44.26, -44.554, -44.307, -44.231]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.338
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [24445.0, 24982.0, 25518.0, 26054.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7330833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.28566666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.173
varray: [2.191, 2.201, 2.154, 2.146]
inst: mean_all
operation: mean_all(residuals)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: residuals
SAMP_residuals_STD | 0.0270678160675
varray: [2.191, 2.201, 2.154, 2.146]
inst: std_all
operation: std_all(residuals)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0525
varray: [-1.062, -1.126, -1.027, -0.995]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0561218911537
varray: [-1.062, -1.126, -1.027, -0.995]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.2345
varray: [43.365, 43.597, 42.983, 42.993]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.300006111049
varray: [43.365, 43.597, 42.983, 42.993]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.67825
varray: [29.153, 29.084, 28.232, 28.244]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.509160338204
varray: [29.153, 29.084, 28.232, 28.244]
inst: std_all
operation: std_all(slope_shift)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0530094331228
varray: [-6.515, -6.466, -6.404, -6.407]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.381790845883
varray: [-42.744, -42.36, -41.913, -41.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.449
varray: [63.691, 63.89, 63.416, 62.799]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.474915431068
varray: [63.691, 63.89, 63.416, 62.799]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.173
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.2345
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.67825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0927619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0927619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.45919652054
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.5899047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.515, -6.466, -6.404, -6.407]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.744, -42.36, -41.913, -41.99]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.448
varray: [-6.515, -6.466, -6.404, -6.407]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
dD_mean | -42.25175
varray: [-42.744, -42.36, -41.913, -41.99]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0530094331228
varray: [-6.515, -6.466, -6.404, -6.407]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
dD_stdev | 0.381790845883
varray: [-42.744, -42.36, -41.913, -41.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.448
varray: [-6.515, -6.466, -6.404, -6.407]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.25175
varray: [-42.744, -42.36, -41.913, -41.99]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0530094331228
varray: [-6.515, -6.466, -6.404, -6.407]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.381790845883
varray: [-42.744, -42.36, -41.913, -41.99]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.4113863075
actual_array: [ 1.93 -15.84]
varray: [-6.515, -6.466, -6.404, -6.407]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.1193810632
actual_array: [ 18.1 -119. ]
varray: [-42.744, -42.36, -41.913, -41.99]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.25175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [26644.0, 27180.0, 27716.0, 28252.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7333333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.32816666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11425
varray: [2.096, 2.103, 2.136, 2.122]
inst: mean_all
operation: mean_all(residuals)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: residuals
SAMP_residuals_STD | 0.0181911150474
varray: [2.096, 2.103, 2.136, 2.122]
inst: std_all
operation: std_all(residuals)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.018
varray: [-1.095, -0.969, -0.873, -1.135]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.119783137378
varray: [-1.095, -0.969, -0.873, -1.135]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.5305
varray: [42.451, 42.282, 42.533, 42.856]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.240851960064
varray: [42.451, 42.282, 42.533, 42.856]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.9405
varray: [27.618, 27.667, 28.194, 28.283]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.346591498261
varray: [27.618, 27.667, 28.194, 28.283]
inst: std_all
operation: std_all(slope_shift)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0249131825881
varray: [-6.435, -6.494, -6.453, -6.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.147031459219
varray: [-41.317, -41.635, -41.345, -41.494]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.8465
varray: [60.673, 62.156, 62.918, 61.639]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.942348307864
varray: [60.673, 62.156, 62.918, 61.639]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11425
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.018
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.5305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.9405
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0340119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0340119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.63500057228
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.987404761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.435, -6.494, -6.453, -6.454]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-41.317, -41.635, -41.345, -41.494]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.459
varray: [-6.435, -6.494, -6.453, -6.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
dD_mean | -41.44775
varray: [-41.317, -41.635, -41.345, -41.494]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0249131825881
varray: [-6.435, -6.494, -6.453, -6.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
dD_stdev | 0.147031459219
varray: [-41.317, -41.635, -41.345, -41.494]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.459
varray: [-6.435, -6.494, -6.453, -6.454]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.44775
varray: [-41.317, -41.635, -41.345, -41.494]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0249131825881
varray: [-6.435, -6.494, -6.453, -6.454]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.147031459219
varray: [-41.317, -41.635, -41.345, -41.494]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.42366903467
actual_array: [ 1.93 -15.84]
varray: [-6.435, -6.494, -6.453, -6.454]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.1710737998
actual_array: [ 18.1 -119. ]
varray: [-41.317, -41.635, -41.345, -41.494]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.44775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [28840.0, 29377.0, 29914.0, 30450.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7920833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.624166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14675
varray: [2.188, 2.13, 2.136, 2.133]
inst: mean_all
operation: mean_all(residuals)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: residuals
SAMP_residuals_STD | 0.0276088753846
varray: [2.188, 2.13, 2.136, 2.133]
inst: std_all
operation: std_all(residuals)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98575
varray: [-0.792, -0.941, -1.086, -1.124]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.151332690894
varray: [-0.792, -0.941, -1.086, -1.124]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.85175
varray: [43.256, 42.721, 42.768, 42.662]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.272966878333
varray: [43.256, 42.721, 42.768, 42.662]
inst: std_all
operation: std_all(baseline_shift)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.462
varray: [28.862, 28.142, 28.353, 28.491]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.302832186753
varray: [28.862, 28.142, 28.353, 28.491]
inst: std_all
operation: std_all(slope_shift)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0204124145232
varray: [-6.676, -6.63, -6.653, -6.637]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.156319118046
varray: [-43.023, -42.817, -42.643, -42.863]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.8295
varray: [64.855, 62.693, 61.793, 61.977]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.40503511225
varray: [64.855, 62.693, 61.793, 61.977]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.85175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.462
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0665119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0665119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.19732173515
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.9704047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.676, -6.63, -6.653, -6.637]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-43.023, -42.817, -42.643, -42.863]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.649
varray: [-6.676, -6.63, -6.653, -6.637]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
dD_mean | -42.8365
varray: [-43.023, -42.817, -42.643, -42.863]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0204124145232
varray: [-6.676, -6.63, -6.653, -6.637]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
dD_stdev | 0.156319118046
varray: [-43.023, -42.817, -42.643, -42.863]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.649
varray: [-6.676, -6.63, -6.653, -6.637]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.8365
varray: [-43.023, -42.817, -42.643, -42.863]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0204124145232
varray: [-6.676, -6.63, -6.653, -6.637]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.156319118046
varray: [-43.023, -42.817, -42.643, -42.863]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.63582523118
actual_array: [ 1.93 -15.84]
varray: [-6.676, -6.63, -6.653, -6.637]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.649
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.8090858795
actual_array: [ 18.1 -119. ]
varray: [-43.023, -42.817, -42.643, -42.863]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.8365
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [31038.0, 31574.0, 32110.0, 32646.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7595833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.945416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12725
varray: [2.16, 2.113, 2.119, 2.117]
inst: mean_all
operation: mean_all(residuals)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: residuals
SAMP_residuals_STD | 0.0219753649951
varray: [2.16, 2.113, 2.119, 2.117]
inst: std_all
operation: std_all(residuals)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9455
varray: [-1.008, -0.886, -0.769, -1.119]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.151328560865
varray: [-1.008, -0.886, -0.769, -1.119]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.62125
varray: [43.07, 42.481, 42.371, 42.563]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.309335174204
varray: [43.07, 42.481, 42.371, 42.563]
inst: std_all
operation: std_all(baseline_shift)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.16725
varray: [28.736, 28.125, 27.79, 28.018]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.404085283903
varray: [28.736, 28.125, 27.79, 28.018]
inst: std_all
operation: std_all(slope_shift)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: slope_shift
SAMP_d18O_STD | 0.061808845106
varray: [-6.208, -6.182, -6.097, -6.083]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.507141909791
varray: [-39.934, -39.303, -38.8, -38.936]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.98375
varray: [63.655, 62.974, 63.376, 61.93]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.756068945269
varray: [63.655, 62.974, 63.376, 61.93]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9455
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.62125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.16725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0470119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0470119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.25992903743
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.1246547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.208, -6.182, -6.097, -6.083]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-39.934, -39.303, -38.8, -38.936]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.1425
varray: [-6.208, -6.182, -6.097, -6.083]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
dD_mean | -39.24325
varray: [-39.934, -39.303, -38.8, -38.936]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
d18O_stdev | 0.061808845106
varray: [-6.208, -6.182, -6.097, -6.083]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
dD_stdev | 0.507141909791
varray: [-39.934, -39.303, -38.8, -38.936]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.1425
varray: [-6.208, -6.182, -6.097, -6.083]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.24325
varray: [-39.934, -39.303, -38.8, -38.936]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.061808845106
varray: [-6.208, -6.182, -6.097, -6.083]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.507141909791
varray: [-39.934, -39.303, -38.8, -38.936]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.07026147573
actual_array: [ 1.93 -15.84]
varray: [-6.208, -6.182, -6.097, -6.083]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.1425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.5708954885
actual_array: [ 18.1 -119. ]
varray: [-39.934, -39.303, -38.8, -38.936]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.24325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [33234.0, 33769.0, 34306.0, 34843.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7790833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.714916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1345
varray: [2.181, 2.108, 2.117, 2.132]
inst: mean_all
operation: mean_all(residuals)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: residuals
SAMP_residuals_STD | 0.0325422801905
varray: [2.181, 2.108, 2.117, 2.132]
inst: std_all
operation: std_all(residuals)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95225
varray: [-0.993, -0.97, -0.97, -0.876]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0519767576262
varray: [-0.993, -0.97, -0.97, -0.876]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.816
varray: [43.425, 42.537, 42.731, 42.571]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.414717574581
varray: [43.425, 42.537, 42.731, 42.571]
inst: std_all
operation: std_all(baseline_shift)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.453
varray: [29.007, 28.569, 28.19, 28.046]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.430189880557
varray: [29.007, 28.569, 28.19, 28.046]
inst: std_all
operation: std_all(slope_shift)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0300485717908
varray: [-6.805, -6.753, -6.734, -6.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.052696457819
varray: [-43.551, -43.592, -43.597, -43.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.682
varray: [64.158, 61.539, 62.125, 62.906]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.13217342606
varray: [64.158, 61.539, 62.125, 62.906]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1345
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.816
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.453
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0542619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0542619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.6084468353
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8229047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.805, -6.753, -6.734, -6.761]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-43.551, -43.592, -43.597, -43.677]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.76325
varray: [-6.805, -6.753, -6.734, -6.761]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
dD_mean | -43.60425
varray: [-43.551, -43.592, -43.597, -43.677]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0300485717908
varray: [-6.805, -6.753, -6.734, -6.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
dD_stdev | 0.052696457819
varray: [-43.551, -43.592, -43.597, -43.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.76325
varray: [-6.805, -6.753, -6.734, -6.761]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.60425
varray: [-43.551, -43.592, -43.597, -43.677]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0300485717908
varray: [-6.805, -6.753, -6.734, -6.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.052696457819
varray: [-43.551, -43.592, -43.597, -43.677]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.76339810198
actual_array: [ 1.93 -15.84]
varray: [-6.805, -6.753, -6.734, -6.761]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.76325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.714636752
actual_array: [ 18.1 -119. ]
varray: [-43.551, -43.592, -43.597, -43.677]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.60425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [35432.0, 35969.0, 36505.0, 37042.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7718333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.909666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1265
varray: [2.159, 2.113, 2.12, 2.114]
inst: mean_all
operation: mean_all(residuals)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: residuals
SAMP_residuals_STD | 0.0218860686282
varray: [2.159, 2.113, 2.12, 2.114]
inst: std_all
operation: std_all(residuals)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95525
varray: [-0.936, -0.95, -0.883, -1.052]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0706605736367
varray: [-0.936, -0.95, -0.883, -1.052]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.7175
varray: [43.176, 42.608, 42.478, 42.608]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.311749365142
varray: [43.176, 42.608, 42.478, 42.608]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.219
varray: [28.823, 28.158, 27.843, 28.052]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.423399732326
varray: [28.823, 28.158, 27.843, 28.052]
inst: std_all
operation: std_all(slope_shift)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0382317494586
varray: [-6.887, -6.871, -6.894, -6.958]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.129762282656
varray: [-44.734, -44.909, -44.953, -45.043]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.977
varray: [63.385, 62.798, 63.38, 62.345]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.503434206228
varray: [63.385, 62.798, 63.38, 62.345]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1265
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.7175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.219
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0462619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0462619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.22387547213
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.1179047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.887, -6.871, -6.894, -6.958]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-44.734, -44.909, -44.953, -45.043]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.9025
varray: [-6.887, -6.871, -6.894, -6.958]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
dD_mean | -44.90975
varray: [-44.734, -44.909, -44.953, -45.043]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0382317494586
varray: [-6.887, -6.871, -6.894, -6.958]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
dD_stdev | 0.129762282656
varray: [-44.734, -44.909, -44.953, -45.043]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.9025
varray: [-6.887, -6.871, -6.894, -6.958]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
measured_dD_mean | -44.90975
varray: [-44.734, -44.909, -44.953, -45.043]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0382317494586
varray: [-6.887, -6.871, -6.894, -6.958]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.129762282656
varray: [-44.734, -44.909, -44.953, -45.043]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.9188862618
actual_array: [ 1.93 -15.84]
varray: [-6.887, -6.871, -6.894, -6.958]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.9025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.2544565684
actual_array: [ 18.1 -119. ]
varray: [-44.734, -44.909, -44.953, -45.043]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.90975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [37632.0, 38168.0, 38705.0, 39241.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7798333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.811166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10225
varray: [2.123, 2.099, 2.096, 2.091]
inst: mean_all
operation: mean_all(residuals)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: residuals
SAMP_residuals_STD | 0.0142214626533
varray: [2.123, 2.099, 2.096, 2.091]
inst: std_all
operation: std_all(residuals)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00675
varray: [-1.057, -1.048, -1.017, -0.905]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0699636810543
varray: [-1.057, -1.048, -1.017, -0.905]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.7545
varray: [42.906, 42.915, 42.619, 42.578]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.18094658512
varray: [42.906, 42.915, 42.619, 42.578]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.19325
varray: [28.66, 27.894, 28.162, 28.057]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.33012459769
varray: [28.66, 27.894, 28.162, 28.057]
inst: std_all
operation: std_all(slope_shift)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0741057802514
varray: [-8.821, -8.979, -8.943, -8.975]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.10957690886
varray: [-63.839, -65.983, -65.997, -66.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.87275
varray: [62.141, 62.502, 61.538, 61.31]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.546704902118
varray: [62.141, 62.502, 61.538, 61.31]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.7545
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.19325
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0220119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0220119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.05814352753
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.0136547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-8.821, -8.979, -8.943, -8.975]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-63.839, -65.983, -65.997, -66.174]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9295
varray: [-8.821, -8.979, -8.943, -8.975]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
dD_mean | -65.49825
varray: [-63.839, -65.983, -65.997, -66.174]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0741057802514
varray: [-8.821, -8.979, -8.943, -8.975]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
dD_stdev | 1.10957690886
varray: [-63.839, -65.983, -65.997, -66.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9295
varray: [-8.821, -8.979, -8.943, -8.975]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.49825
varray: [-63.839, -65.983, -65.997, -66.174]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0741057802514
varray: [-8.821, -8.979, -8.943, -8.975]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.10957690886
varray: [-63.839, -65.983, -65.997, -66.174]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.1822578952
actual_array: [ 1.93 -15.84]
varray: [-8.821, -8.979, -8.943, -8.975]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.9295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.5383173807
actual_array: [ 18.1 -119. ]
varray: [-63.839, -65.983, -65.997, -66.174]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.49825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [39831.0, 40367.0, 40903.0, 41440.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8040833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.848166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.133
varray: [2.162, 2.13, 2.125, 2.115]
inst: mean_all
operation: mean_all(residuals)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: residuals
SAMP_residuals_STD | 0.0203141986469
varray: [2.162, 2.13, 2.125, 2.115]
inst: std_all
operation: std_all(residuals)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.99425
varray: [-0.995, -1.053, -0.914, -1.015]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0586593271924
varray: [-0.995, -1.053, -0.914, -1.015]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.52675
varray: [42.876, 42.444, 42.361, 42.426]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.235546704498
varray: [42.876, 42.444, 42.361, 42.426]
inst: std_all
operation: std_all(baseline_shift)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.074
varray: [28.455, 28.524, 27.762, 27.555]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.487977458496
varray: [28.455, 28.524, 27.762, 27.555]
inst: std_all
operation: std_all(slope_shift)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: slope_shift
SAMP_d18O_STD | 0.174389172064
varray: [-4.618, -4.327, -4.278, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.93756313876
varray: [-38.126, -35.175, -34.314, -33.784]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.533
varray: [63.15, 61.622, 62.589, 62.771]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.650740603723
varray: [63.15, 61.622, 62.589, 62.771]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.133
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.99425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.52675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.074
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0527619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0527619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.5363397047
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.6739047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-4.618, -4.327, -4.278, -4.23]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-38.126, -35.175, -34.314, -33.784]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.36325
varray: [-4.618, -4.327, -4.278, -4.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
dD_mean | -35.34975
varray: [-38.126, -35.175, -34.314, -33.784]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
d18O_stdev | 0.174389172064
varray: [-4.618, -4.327, -4.278, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
dD_stdev | 1.93756313876
varray: [-38.126, -35.175, -34.314, -33.784]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.36325
varray: [-4.618, -4.327, -4.278, -4.23]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.34975
varray: [-38.126, -35.175, -34.314, -33.784]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.174389172064
varray: [-4.618, -4.327, -4.278, -4.23]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.93756313876
varray: [-38.126, -35.175, -34.314, -33.784]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.0835303565
actual_array: [ 1.93 -15.84]
varray: [-4.618, -4.327, -4.278, -4.23]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.36325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.9785642321
actual_array: [ 18.1 -119. ]
varray: [-38.126, -35.175, -34.314, -33.784]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -35.34975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [42030.0, 42566.0, 43103.0, 43640.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7733333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.620416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10575
varray: [2.14, 2.101, 2.093, 2.089]
inst: mean_all
operation: mean_all(residuals)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: residuals
SAMP_residuals_STD | 0.023371991785
varray: [2.14, 2.101, 2.093, 2.089]
inst: std_all
operation: std_all(residuals)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04225
varray: [-1.098, -1.015, -1.055, -1.001]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0436453510315
varray: [-1.098, -1.015, -1.055, -1.001]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.67475
varray: [42.93, 42.629, 42.646, 42.494]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.183250966346
varray: [42.93, 42.629, 42.646, 42.494]
inst: std_all
operation: std_all(baseline_shift)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3515
varray: [29.102, 28.128, 28.219, 27.957]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.511986002413
varray: [29.102, 28.128, 28.219, 27.957]
inst: std_all
operation: std_all(slope_shift)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0808718121474
varray: [-7.433, -7.619, -7.544, -7.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.06049689611
varray: [-51.963, -54.002, -54.07, -54.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.872
varray: [62.527, 61.794, 61.355, 61.812]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.485110983865
varray: [62.527, 61.794, 61.355, 61.812]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.67475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3515
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0255119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0255119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.22639349891
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.0129047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-7.433, -7.619, -7.544, -7.585]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-51.963, -54.002, -54.07, -54.167]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.54525
varray: [-7.433, -7.619, -7.544, -7.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
dD_mean | -53.5505
varray: [-51.963, -54.002, -54.07, -54.167]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0808718121474
varray: [-7.433, -7.619, -7.544, -7.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
dD_stdev | 1.06049689611
varray: [-51.963, -54.002, -54.07, -54.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.54525
varray: [-7.433, -7.619, -7.544, -7.585]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
measured_dD_mean | -53.5505
varray: [-51.963, -54.002, -54.07, -54.167]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0808718121474
varray: [-7.433, -7.619, -7.544, -7.585]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.06049689611
varray: [-51.963, -54.002, -54.07, -54.167]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.63658834239
actual_array: [ 1.93 -15.84]
varray: [-7.433, -7.619, -7.544, -7.585]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.54525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -65.4461058053
actual_array: [ 18.1 -119. ]
varray: [-51.963, -54.002, -54.07, -54.167]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -53.5505
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [44227.0, 44762.0, 45299.0, 45834.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8005833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.768416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.9925
varray: [2.021, 1.982, 1.989, 1.978]
inst: mean_all
operation: mean_all(residuals)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: residuals
SAMP_residuals_STD | 0.0195362910161
varray: [2.021, 1.982, 1.989, 1.978]
inst: std_all
operation: std_all(residuals)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9685
varray: [-0.968, -0.987, -1.02, -0.899]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0510718448201
varray: [-0.968, -0.987, -1.02, -0.899]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.48275
varray: [42.866, 42.336, 42.407, 42.322]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.258195500864
varray: [42.866, 42.336, 42.407, 42.322]
inst: std_all
operation: std_all(baseline_shift)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.6415
varray: [28.48, 27.49, 27.585, 27.011]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.612850443964
varray: [28.48, 27.49, 27.585, 27.011]
inst: std_all
operation: std_all(slope_shift)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: slope_shift
SAMP_d18O_STD | 0.175179907524
varray: [-13.382, -13.694, -13.738, -13.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.22950037415
varray: [-105.572, -110.597, -111.919, -112.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.72975
varray: [59.895, 59.53, 58.855, 60.639]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.743669001192
varray: [59.895, 59.53, 58.855, 60.639]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.9925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9685
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.48275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.6415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0877380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0877380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.21769486094
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1293452381
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-13.382, -13.694, -13.738, -13.754]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-105.572, -110.597, -111.919, -112.8]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -13.642
varray: [-13.382, -13.694, -13.738, -13.754]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
dD_mean | -110.222
varray: [-105.572, -110.597, -111.919, -112.8]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
d18O_stdev | 0.175179907524
varray: [-13.382, -13.694, -13.738, -13.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
dD_stdev | 3.22950037415
varray: [-105.572, -110.597, -111.919, -112.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
measured_d18O_mean | -13.642
varray: [-13.382, -13.694, -13.738, -13.754]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
measured_dD_mean | -110.222
varray: [-105.572, -110.597, -111.919, -112.8]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.175179907524
varray: [-13.382, -13.694, -13.738, -13.754]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.22950037415
varray: [-105.572, -110.597, -111.919, -112.8]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -16.4442898746
actual_array: [ 1.93 -15.84]
varray: [-13.382, -13.694, -13.738, -13.754]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -132.289383267
actual_array: [ 18.1 -119. ]
varray: [-105.572, -110.597, -111.919, -112.8]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -110.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [46422.0, 46958.0, 47494.0, 48031.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9138333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.576416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.125
varray: [2.155, 2.122, 2.116, 2.107]
inst: mean_all
operation: mean_all(residuals)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: residuals
SAMP_residuals_STD | 0.0209284495365
varray: [2.155, 2.122, 2.116, 2.107]
inst: std_all
operation: std_all(residuals)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.11425
varray: [-1.045, -1.14, -1.158, -1.114]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0495740187329
varray: [-1.045, -1.14, -1.158, -1.114]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.497
varray: [43.036, 42.475, 42.36, 42.117]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.38909039224
varray: [43.036, 42.475, 42.36, 42.117]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.10375
varray: [28.41, 28.043, 27.848, 28.114]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.233089932573
varray: [28.41, 28.043, 27.848, 28.114]
inst: std_all
operation: std_all(slope_shift)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: slope_shift
SAMP_d18O_STD | 0.287337635312
varray: [-3.782, -3.301, -3.213, -3.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.81942323658
varray: [-30.628, -22.916, -20.909, -20.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.596
varray: [63.338, 61.255, 61.49, 60.301]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.27005327972
varray: [63.338, 61.255, 61.49, 60.301]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.11425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.497
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.10375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0447619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0447619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.15176834154
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.736904761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-3.782, -3.301, -3.213, -3.149]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-30.628, -22.916, -20.909, -20.053]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.36125
varray: [-3.782, -3.301, -3.213, -3.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
dD_mean | -23.6265
varray: [-30.628, -22.916, -20.909, -20.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
d18O_stdev | 0.287337635312
varray: [-3.782, -3.301, -3.213, -3.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
dD_stdev | 4.81942323658
varray: [-30.628, -22.916, -20.909, -20.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.36125
varray: [-3.782, -3.301, -3.213, -3.149]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
measured_dD_mean | -23.6265
varray: [-30.628, -22.916, -20.909, -20.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.287337635312
varray: [-3.782, -3.301, -3.213, -3.149]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.81942323658
varray: [-30.628, -22.916, -20.909, -20.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.96468557276
actual_array: [ 1.93 -15.84]
varray: [-3.782, -3.301, -3.213, -3.149]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -3.36125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -30.1511474088
actual_array: [ 18.1 -119. ]
varray: [-30.628, -22.916, -20.909, -20.053]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -23.6265
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [48620.0, 49156.0, 49691.0, 50226.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7813333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.590666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11525
varray: [2.138, 2.09, 2.118, 2.115]
inst: mean_all
operation: mean_all(residuals)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: residuals
SAMP_residuals_STD | 0.01968713624
varray: [2.138, 2.09, 2.118, 2.115]
inst: std_all
operation: std_all(residuals)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03175
varray: [-1.17, -1.08, -0.912, -0.965]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.115811268882
varray: [-1.17, -1.08, -0.912, -0.965]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.5935
varray: [42.878, 42.506, 42.488, 42.502]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.189823602326
varray: [42.878, 42.506, 42.488, 42.502]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.03925
varray: [28.652, 27.834, 27.546, 28.125]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.471960008899
varray: [28.652, 27.834, 27.546, 28.125]
inst: std_all
operation: std_all(slope_shift)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0651683716742
varray: [-5.712, -5.833, -5.828, -5.858]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.00642767583
varray: [-38.15, -40.129, -40.126, -40.227]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.1255
varray: [61.838, 61.554, 62.986, 62.124]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.619066770766
varray: [61.838, 61.554, 62.986, 62.124]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.5935
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.03925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0350119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0350119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.68307199267
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.2664047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-5.712, -5.833, -5.828, -5.858]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-38.15, -40.129, -40.126, -40.227]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.80775
varray: [-5.712, -5.833, -5.828, -5.858]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
dD_mean | -39.658
varray: [-38.15, -40.129, -40.126, -40.227]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0651683716742
varray: [-5.712, -5.833, -5.828, -5.858]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
dD_stdev | 1.00642767583
varray: [-38.15, -40.129, -40.126, -40.227]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.80775
varray: [-5.712, -5.833, -5.828, -5.858]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
measured_dD_mean | -39.658
varray: [-38.15, -40.129, -40.126, -40.227]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0651683716742
varray: [-5.712, -5.833, -5.828, -5.858]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.00642767583
varray: [-38.15, -40.129, -40.126, -40.227]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.69647575581
actual_array: [ 1.93 -15.84]
varray: [-5.712, -5.833, -5.828, -5.858]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.80775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.060087575
actual_array: [ 18.1 -119. ]
varray: [-38.15, -40.129, -40.126, -40.227]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.658
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [50815.0, 51350.0, 51887.0, 52423.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7910833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.687166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08925
varray: [2.118, 2.074, 2.078, 2.087]
inst: mean_all
operation: mean_all(residuals)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: residuals
SAMP_residuals_STD | 0.0199227675454
varray: [2.118, 2.074, 2.078, 2.087]
inst: std_all
operation: std_all(residuals)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97075
varray: [-1.01, -0.995, -0.963, -0.915]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0420188449786
varray: [-1.01, -0.995, -0.963, -0.915]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.43625
varray: [42.801, 42.274, 42.366, 42.304]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.246165763934
varray: [42.801, 42.274, 42.366, 42.304]
inst: std_all
operation: std_all(baseline_shift)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.995
varray: [28.668, 27.806, 27.761, 27.745]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.449409241857
varray: [28.668, 27.806, 27.761, 27.745]
inst: std_all
operation: std_all(slope_shift)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0730958275143
varray: [-7.643, -7.762, -7.724, -7.817]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.656474104592
varray: [-48.355, -49.481, -49.511, -49.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.19
varray: [61.606, 61.152, 61.134, 60.868]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.306224318651
varray: [61.606, 61.152, 61.134, 60.868]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.43625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.995
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0090119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0090119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.433215062378
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.330904761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-7.643, -7.762, -7.724, -7.817]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-48.355, -49.481, -49.511, -49.868]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.7365
varray: [-7.643, -7.762, -7.724, -7.817]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
dD_mean | -49.30375
varray: [-48.355, -49.481, -49.511, -49.868]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0730958275143
varray: [-7.643, -7.762, -7.724, -7.817]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
dD_stdev | 0.656474104592
varray: [-48.355, -49.481, -49.511, -49.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.7365
varray: [-7.643, -7.762, -7.724, -7.817]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.30375
varray: [-48.355, -49.481, -49.511, -49.868]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0730958275143
varray: [-7.643, -7.762, -7.724, -7.817]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.656474104592
varray: [-48.355, -49.481, -49.511, -49.868]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.85014030335
actual_array: [ 1.93 -15.84]
varray: [-7.643, -7.762, -7.724, -7.817]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.7365
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.4371208912
actual_array: [ 18.1 -119. ]
varray: [-48.355, -49.481, -49.511, -49.868]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -49.30375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [53011.0, 53548.0, 54084.0, 54619.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8170833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.529916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.9875
varray: [2.026, 1.978, 1.98, 1.966]
inst: mean_all
operation: mean_all(residuals)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: residuals
SAMP_residuals_STD | 0.0264007575649
varray: [2.026, 1.978, 1.98, 1.966]
inst: std_all
operation: std_all(residuals)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0035
varray: [-1.071, -1.042, -0.883, -1.018]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0832045671823
varray: [-1.071, -1.042, -0.883, -1.018]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.25575
varray: [42.629, 42.059, 42.051, 42.284]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.271260483668
varray: [42.629, 42.059, 42.051, 42.284]
inst: std_all
operation: std_all(baseline_shift)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.48675
varray: [28.228, 27.144, 27.204, 27.371]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.50341326628
varray: [28.228, 27.144, 27.204, 27.371]
inst: std_all
operation: std_all(slope_shift)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: slope_shift
SAMP_d18O_STD | 0.200796746322
varray: [-14.314, -14.624, -14.737, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.6008668748
varray: [-105.474, -111.152, -112.662, -113.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.6235
varray: [60.321, 59.203, 59.519, 59.451]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.484425088808
varray: [60.321, 59.203, 59.519, 59.451]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.9875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0035
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.25575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.48675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0927380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0927380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.45805196292
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2355952381
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-14.314, -14.624, -14.737, -14.741]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-105.474, -111.152, -112.662, -113.445]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.604
varray: [-14.314, -14.624, -14.737, -14.741]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
dD_mean | -110.68325
varray: [-105.474, -111.152, -112.662, -113.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
d18O_stdev | 0.200796746322
varray: [-14.314, -14.624, -14.737, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
dD_stdev | 3.6008668748
varray: [-105.474, -111.152, -112.662, -113.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.604
varray: [-14.314, -14.624, -14.737, -14.741]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
measured_dD_mean | -110.68325
varray: [-105.474, -111.152, -112.662, -113.445]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.200796746322
varray: [-14.314, -14.624, -14.737, -14.741]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.6008668748
varray: [-105.474, -111.152, -112.662, -113.445]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -17.5184701959
actual_array: [ 1.93 -15.84]
varray: [-14.314, -14.624, -14.737, -14.741]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.604
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -132.833421483
actual_array: [ 18.1 -119. ]
varray: [-105.474, -111.152, -112.662, -113.445]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -110.68325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [55208.0, 55742.0, 56279.0, 56814.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9188333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.349416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1275
varray: [2.15, 2.107, 2.131, 2.122]
inst: mean_all
operation: mean_all(residuals)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: residuals
SAMP_residuals_STD | 0.0179722007556
varray: [2.15, 2.107, 2.131, 2.122]
inst: std_all
operation: std_all(residuals)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9815
varray: [-1.004, -1.081, -1.001, -0.84]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101339363856
varray: [-1.004, -1.081, -1.001, -0.84]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.448
varray: [42.78, 42.388, 42.347, 42.277]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.226028021862
varray: [42.78, 42.388, 42.347, 42.277]
inst: std_all
operation: std_all(baseline_shift)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.02675
varray: [28.533, 27.99, 27.802, 27.782]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.35026406705
varray: [28.533, 27.99, 27.802, 27.782]
inst: std_all
operation: std_all(slope_shift)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: slope_shift
SAMP_d18O_STD | 0.275781767587
varray: [-5.007, -4.505, -4.443, -4.43]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 4.35139674894
varray: [-39.298, -32.211, -30.494, -29.799]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.54425
varray: [63.027, 61.821, 62.28, 63.049]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.600205173253
varray: [63.027, 61.821, 62.28, 63.049]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9815
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.448
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.02675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0472619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0472619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.27194689253
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.6851547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-5.007, -4.505, -4.443, -4.43]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-39.298, -32.211, -30.494, -29.799]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.59625
varray: [-5.007, -4.505, -4.443, -4.43]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
dD_mean | -32.9505
varray: [-39.298, -32.211, -30.494, -29.799]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
d18O_stdev | 0.275781767587
varray: [-5.007, -4.505, -4.443, -4.43]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
dD_stdev | 4.35139674894
varray: [-39.298, -32.211, -30.494, -29.799]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.59625
varray: [-5.007, -4.505, -4.443, -4.43]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
measured_dD_mean | -32.9505
varray: [-39.298, -32.211, -30.494, -29.799]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.275781767587
varray: [-5.007, -4.505, -4.443, -4.43]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
measured_dD_stdev | 4.35139674894
varray: [-39.298, -32.211, -30.494, -29.799]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.34370085013
actual_array: [ 1.93 -15.84]
varray: [-5.007, -4.505, -4.443, -4.43]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.59625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -41.1486808966
actual_array: [ 18.1 -119. ]
varray: [-39.298, -32.211, -30.494, -29.799]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -32.9505
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [57403.0, 57939.0, 58476.0, 59012.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7788333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.541666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0785
varray: [2.121, 2.068, 2.062, 2.063]
inst: mean_all
operation: mean_all(residuals)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: residuals
SAMP_residuals_STD | 0.0284546422692
varray: [2.121, 2.068, 2.062, 2.063]
inst: std_all
operation: std_all(residuals)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06675
varray: [-0.963, -1.074, -1.047, -1.183]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0907758227724
varray: [-0.963, -1.074, -1.047, -1.183]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.54025
varray: [42.957, 42.36, 42.402, 42.442]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.279843259701
varray: [42.957, 42.36, 42.402, 42.442]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.951
varray: [28.304, 27.766, 27.853, 27.881]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.24037332076
varray: [28.304, 27.766, 27.853, 27.881]
inst: std_all
operation: std_all(slope_shift)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: slope_shift
SAMP_d18O_STD | 0.114494177436
varray: [-8.717, -8.917, -8.947, -8.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.9194299501
varray: [-62.306, -65.497, -66.14, -66.517]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.401
varray: [62.917, 60.776, 61.46, 60.451]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09466311408
varray: [62.917, 60.776, 61.46, 60.451]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0785
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06675
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.54025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.951
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0017380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0017380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0835527068788
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.541904761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-8.717, -8.917, -8.947, -8.964]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-62.306, -65.497, -66.14, -66.517]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.88625
varray: [-8.717, -8.917, -8.947, -8.964]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
dD_mean | -65.115
varray: [-62.306, -65.497, -66.14, -66.517]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
d18O_stdev | 0.114494177436
varray: [-8.717, -8.917, -8.947, -8.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
dD_stdev | 1.9194299501
varray: [-62.306, -65.497, -66.14, -66.517]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.88625
varray: [-8.717, -8.917, -8.947, -8.964]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.115
varray: [-62.306, -65.497, -66.14, -66.517]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.114494177436
varray: [-8.717, -8.917, -8.947, -8.964]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.9194299501
varray: [-62.306, -65.497, -66.14, -66.517]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.1339644452
actual_array: [ 1.93 -15.84]
varray: [-8.717, -8.917, -8.947, -8.964]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.88625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.0862791236
actual_array: [ 18.1 -119. ]
varray: [-62.306, -65.497, -66.14, -66.517]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.115
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [59597.0, 60134.0, 60670.0, 61206.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8278333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.633916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1065
varray: [2.142, 2.092, 2.094, 2.098]
inst: mean_all
operation: mean_all(residuals)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: residuals
SAMP_residuals_STD | 0.0237977589981
varray: [2.142, 2.092, 2.094, 2.098]
inst: std_all
operation: std_all(residuals)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0385
varray: [-0.974, -1.168, -0.965, -1.047]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0938171981391
varray: [-0.974, -1.168, -0.965, -1.047]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.43775
varray: [42.817, 42.387, 42.266, 42.281]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.258505157911
varray: [42.817, 42.387, 42.266, 42.281]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.849
varray: [28.129, 27.91, 27.738, 27.619]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.221616786368
varray: [28.129, 27.91, 27.738, 27.619]
inst: std_all
operation: std_all(slope_shift)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0944351629426
varray: [-5.704, -5.56, -5.505, -5.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.78019200743
varray: [-37.664, -35.224, -34.058, -33.739]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.50975
varray: [63.661, 60.699, 60.69, 60.989]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.44087505241
varray: [63.661, 60.699, 60.69, 60.989]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1065
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0385
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.43775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.849
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0262619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0262619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.26244706421
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.650654761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-5.704, -5.56, -5.505, -5.503]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-37.664, -35.224, -34.058, -33.739]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.568
varray: [-5.704, -5.56, -5.505, -5.503]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
dD_mean | -35.17125
varray: [-37.664, -35.224, -34.058, -33.739]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0944351629426
varray: [-5.704, -5.56, -5.505, -5.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
dD_stdev | 1.78019200743
varray: [-37.664, -35.224, -34.058, -33.739]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.568
varray: [-5.704, -5.56, -5.505, -5.503]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
measured_dD_mean | -35.17125
varray: [-37.664, -35.224, -34.058, -33.739]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0944351629426
varray: [-5.704, -5.56, -5.505, -5.503]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.78019200743
varray: [-37.664, -35.224, -34.058, -33.739]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.42876813416
actual_array: [ 1.93 -15.84]
varray: [-5.704, -5.56, -5.505, -5.503]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.568
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.7680258658
actual_array: [ 18.1 -119. ]
varray: [-37.664, -35.224, -34.058, -33.739]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -35.17125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [61792.0, 62328.0, 62864.0, 63400.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7998333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.531416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14675
varray: [2.17, 2.151, 2.126, 2.14]
inst: mean_all
operation: mean_all(residuals)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: residuals
SAMP_residuals_STD | 0.0185719322276
varray: [2.17, 2.151, 2.126, 2.14]
inst: std_all
operation: std_all(residuals)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03325
varray: [-1.056, -0.975, -1.282, -0.82]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.192581368085
varray: [-1.056, -0.975, -1.282, -0.82]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.613
varray: [42.876, 42.525, 42.589, 42.462]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.182838726751
varray: [42.876, 42.525, 42.589, 42.462]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.221
varray: [28.81, 28.203, 27.981, 27.89]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.414087752375
varray: [28.81, 28.203, 27.981, 27.89]
inst: std_all
operation: std_all(slope_shift)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0373853803868
varray: [-4.852, -4.784, -4.765, -4.797]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.79799519631
varray: [-22.734, -21.601, -21.115, -20.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.44
varray: [62.829, 63.021, 61.26, 62.65]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.801120465348
varray: [62.829, 63.021, 61.26, 62.65]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.613
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.221
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0665119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0665119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.19732173515
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.5809047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-4.852, -4.784, -4.765, -4.797]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-22.734, -21.601, -21.115, -20.976]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.7995
varray: [-4.852, -4.784, -4.765, -4.797]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
dD_mean | -21.6065
varray: [-22.734, -21.601, -21.115, -20.976]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0373853803868
varray: [-4.852, -4.784, -4.765, -4.797]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
dD_stdev | 0.79799519631
varray: [-22.734, -21.601, -21.115, -20.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.7995
varray: [-4.852, -4.784, -4.765, -4.797]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
measured_dD_mean | -21.6065
varray: [-22.734, -21.601, -21.115, -20.976]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0373853803868
varray: [-4.852, -4.784, -4.765, -4.797]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.79799519631
varray: [-22.734, -21.601, -21.115, -20.976]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.57065214982
actual_array: [ 1.93 -15.84]
varray: [-4.852, -4.784, -4.765, -4.797]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.7995
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -27.7685843837
actual_array: [ 18.1 -119. ]
varray: [-22.734, -21.601, -21.115, -20.976]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -21.6065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [63987.0, 64523.0, 65059.0, 65594.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7595833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.706666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.095
varray: [2.137, 2.081, 2.09, 2.072]
inst: mean_all
operation: mean_all(residuals)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: residuals
SAMP_residuals_STD | 0.0289482296523
varray: [2.137, 2.081, 2.09, 2.072]
inst: std_all
operation: std_all(residuals)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.264
varray: [-1.384, -1.276, -1.14, -1.256]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0999733297768
varray: [-1.384, -1.276, -1.14, -1.256]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.75925
varray: [43.335, 42.51, 42.69, 42.502]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.39352541214
varray: [43.335, 42.51, 42.69, 42.502]
inst: std_all
operation: std_all(baseline_shift)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.4685
varray: [29.252, 28.282, 28.164, 28.176]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.525017777477
varray: [29.252, 28.282, 28.164, 28.176]
inst: std_all
operation: std_all(slope_shift)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0661387430583
varray: [-7.961, -8.049, -8.121, -8.063]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.82049489608
varray: [-54.032, -56.718, -57.895, -57.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.641
varray: [60.963, 60.521, 60.926, 60.154]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.381435009755
varray: [60.963, 60.521, 60.926, 60.154]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.095
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.264
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.75925
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.4685
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0147619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0147619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.709625729655
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.218095238095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-7.961, -8.049, -8.121, -8.063]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-54.032, -56.718, -57.895, -57.889]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.0485
varray: [-7.961, -8.049, -8.121, -8.063]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
dD_mean | -56.6335
varray: [-54.032, -56.718, -57.895, -57.889]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0661387430583
varray: [-7.961, -8.049, -8.121, -8.063]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
dD_stdev | 1.82049489608
varray: [-54.032, -56.718, -57.895, -57.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.0485
varray: [-7.961, -8.049, -8.121, -8.063]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.6335
varray: [-54.032, -56.718, -57.895, -57.889]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0661387430583
varray: [-7.961, -8.049, -8.121, -8.063]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.82049489608
varray: [-54.032, -56.718, -57.895, -57.889]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1985231103
actual_array: [ 1.93 -15.84]
varray: [-7.961, -8.049, -8.121, -8.063]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.0485
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.0824631351
actual_array: [ 18.1 -119. ]
varray: [-54.032, -56.718, -57.895, -57.889]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -56.6335
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [66181.0, 66717.0, 67252.0, 67788.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8113333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.852916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11525
varray: [2.128, 2.095, 2.115, 2.123]
inst: mean_all
operation: mean_all(residuals)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: residuals
SAMP_residuals_STD | 0.0145229703114
varray: [2.128, 2.095, 2.115, 2.123]
inst: std_all
operation: std_all(residuals)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0705
varray: [-1.186, -1.043, -1.026, -1.027]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0773929368009
varray: [-1.186, -1.043, -1.026, -1.027]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.636
varray: [43.075, 42.376, 42.579, 42.514]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.304660466749
varray: [43.075, 42.376, 42.579, 42.514]
inst: std_all
operation: std_all(baseline_shift)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.14075
varray: [28.74, 27.837, 27.933, 28.053]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.409155532775
varray: [28.74, 27.837, 27.933, 28.053]
inst: std_all
operation: std_all(slope_shift)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0327554066784
varray: [-6.459, -6.415, -6.401, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.741182107627
varray: [-44.371, -43.079, -42.86, -42.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.316
varray: [62.135, 62.689, 62.237, 62.203]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.252256483234
varray: [62.135, 62.689, 62.237, 62.203]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.636
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.14075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0350119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0350119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.68307199267
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.4569047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.459, -6.415, -6.401, -6.382]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-44.371, -43.079, -42.86, -42.789]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.41425
varray: [-6.459, -6.415, -6.401, -6.382]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
dD_mean | -43.27475
varray: [-44.371, -43.079, -42.86, -42.789]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0327554066784
varray: [-6.459, -6.415, -6.401, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
dD_stdev | 0.741182107627
varray: [-44.371, -43.079, -42.86, -42.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.41425
varray: [-6.459, -6.415, -6.401, -6.382]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.27475
varray: [-44.371, -43.079, -42.86, -42.789]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0327554066784
varray: [-6.459, -6.415, -6.401, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.741182107627
varray: [-44.371, -43.079, -42.86, -42.789]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.37370066733
actual_array: [ 1.93 -15.84]
varray: [-6.459, -6.415, -6.401, -6.382]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.41425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.3259959021
actual_array: [ 18.1 -119. ]
varray: [-44.371, -43.079, -42.86, -42.789]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.27475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [68375.0, 68911.0, 69446.0, 69982.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7910833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.729666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1025
varray: [2.137, 2.086, 2.085, 2.102]
inst: mean_all
operation: mean_all(residuals)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: residuals
SAMP_residuals_STD | 0.0242830530755
varray: [2.137, 2.086, 2.085, 2.102]
inst: std_all
operation: std_all(residuals)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9125
varray: [-0.879, -0.916, -0.978, -0.877]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0472052256994
varray: [-0.879, -0.916, -0.978, -0.877]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.3955
varray: [42.711, 42.216, 42.339, 42.316]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.217004608246
varray: [42.711, 42.216, 42.339, 42.316]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.05425
varray: [28.825, 28.129, 27.742, 27.521]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.571984484055
varray: [28.825, 28.129, 27.742, 27.521]
inst: std_all
operation: std_all(slope_shift)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00934523051258
varray: [-6.42, -6.429, -6.415, -6.436]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0776718524735
varray: [-42.75, -42.814, -42.866, -42.933]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.0715
varray: [62.645, 61.495, 61.053, 63.093]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.956043757715
varray: [62.645, 61.495, 61.053, 63.093]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.3955
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.05425
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0222619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0222619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.07016138263
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.2124047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.42, -6.429, -6.415, -6.436]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.75, -42.814, -42.866, -42.933]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.425
varray: [-6.42, -6.429, -6.415, -6.436]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
dD_mean | -42.84075
varray: [-42.75, -42.814, -42.866, -42.933]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00934523051258
varray: [-6.42, -6.429, -6.415, -6.436]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
dD_stdev | 0.0776718524735
varray: [-42.75, -42.814, -42.866, -42.933]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.425
varray: [-6.42, -6.429, -6.415, -6.436]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.84075
varray: [-42.75, -42.814, -42.866, -42.933]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00934523051258
varray: [-6.42, -6.429, -6.415, -6.436]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0776718524735
varray: [-42.75, -42.814, -42.866, -42.933]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.3857042416
actual_array: [ 1.93 -15.84]
varray: [-6.42, -6.429, -6.415, -6.436]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.8140986977
actual_array: [ 18.1 -119. ]
varray: [-42.75, -42.814, -42.866, -42.933]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.84075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [70570.0, 71106.0, 71642.0, 72178.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8038333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.489166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0785
varray: [2.113, 2.065, 2.058, 2.078]
inst: mean_all
operation: mean_all(residuals)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: residuals
SAMP_residuals_STD | 0.0244472220644
varray: [2.113, 2.065, 2.058, 2.078]
inst: std_all
operation: std_all(residuals)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.001
varray: [-1.041, -0.886, -0.978, -1.099]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0912103795263
varray: [-1.041, -0.886, -0.978, -1.099]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.13725
varray: [42.512, 41.867, 41.942, 42.228]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.294296194335
varray: [42.512, 41.867, 41.942, 42.228]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.41725
varray: [27.875, 27.109, 27.206, 27.479]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.3430222296
varray: [27.875, 27.109, 27.206, 27.479]
inst: std_all
operation: std_all(slope_shift)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0401445305532
varray: [-6.434, -6.472, -6.461, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.139112364655
varray: [-42.811, -42.913, -43.018, -42.693]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.317
varray: [61.997, 61.367, 60.511, 61.393]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.611095191712
varray: [61.997, 61.367, 60.511, 61.393]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0785
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.001
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.13725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.41725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0017380952381
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0017380952381
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0835527068788
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.457904761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.434, -6.472, -6.461, -6.382]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.811, -42.913, -43.018, -42.693]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.43725
varray: [-6.434, -6.472, -6.461, -6.382]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
dD_mean | -42.85875
varray: [-42.811, -42.913, -43.018, -42.693]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0401445305532
varray: [-6.434, -6.472, -6.461, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
dD_stdev | 0.139112364655
varray: [-42.811, -42.913, -43.018, -42.693]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.43725
varray: [-6.434, -6.472, -6.461, -6.382]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.85875
varray: [-42.811, -42.913, -43.018, -42.693]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0401445305532
varray: [-6.434, -6.472, -6.461, -6.382]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.139112364655
varray: [-42.811, -42.913, -43.018, -42.693]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.39938273322
actual_array: [ 1.93 -15.84]
varray: [-6.434, -6.472, -6.461, -6.382]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.43725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.8353294573
actual_array: [ 18.1 -119. ]
varray: [-42.811, -42.913, -43.018, -42.693]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.85875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [72766.0, 73301.0, 73837.0, 74374.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8278333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.230916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.106
varray: [2.128, 2.098, 2.094, 2.104]
inst: mean_all
operation: mean_all(residuals)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: residuals
SAMP_residuals_STD | 0.0152315462117
varray: [2.128, 2.098, 2.094, 2.104]
inst: std_all
operation: std_all(residuals)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.08475
varray: [-1.164, -0.995, -1.22, -0.96]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.126742192396
varray: [-1.164, -0.995, -1.22, -0.96]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.52125
varray: [42.87, 42.324, 42.499, 42.392]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.243402786892
varray: [42.87, 42.324, 42.499, 42.392]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.905
varray: [28.449, 27.632, 27.962, 27.577]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.400540468201
varray: [28.449, 27.632, 27.962, 27.577]
inst: std_all
operation: std_all(slope_shift)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00605530070819
varray: [-6.345, -6.331, -6.334, -6.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0789994725721
varray: [-42.695, -42.581, -42.769, -42.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.58375
varray: [61.724, 61.744, 60.745, 62.122]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.588377642335
varray: [61.724, 61.744, 60.745, 62.122]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.106
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.08475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.52125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.905
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0257619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0257619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.23841135401
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.724654761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.345, -6.331, -6.334, -6.338]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.695, -42.581, -42.769, -42.714]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.337
varray: [-6.345, -6.331, -6.334, -6.338]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
dD_mean | -42.68975
varray: [-42.695, -42.581, -42.769, -42.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00605530070819
varray: [-6.345, -6.331, -6.334, -6.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
dD_stdev | 0.0789994725721
varray: [-42.695, -42.581, -42.769, -42.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.337
varray: [-6.345, -6.331, -6.334, -6.338]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.68975
varray: [-42.695, -42.581, -42.769, -42.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00605530070819
varray: [-6.345, -6.331, -6.334, -6.338]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0789994725721
varray: [-42.695, -42.581, -42.769, -42.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.28744242427
actual_array: [ 1.93 -15.84]
varray: [-6.345, -6.331, -6.334, -6.338]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.337
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.6359962142
actual_array: [ 18.1 -119. ]
varray: [-42.695, -42.581, -42.769, -42.714]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.68975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [74958.0, 75494.0, 76029.0, 76565.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8003333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.614916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11025
varray: [2.139, 2.089, 2.111, 2.102]
inst: mean_all
operation: mean_all(residuals)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: residuals
SAMP_residuals_STD | 0.0211876536376
varray: [2.139, 2.089, 2.111, 2.102]
inst: std_all
operation: std_all(residuals)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8845
varray: [-0.887, -0.934, -0.81, -0.907]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0532697537695
varray: [-0.887, -0.934, -0.81, -0.907]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.389
varray: [42.748, 42.163, 42.385, 42.26]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.25600390622
varray: [42.748, 42.163, 42.385, 42.26]
inst: std_all
operation: std_all(baseline_shift)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.61375
varray: [28.303, 27.315, 27.42, 27.417]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.462084678387
varray: [28.303, 27.315, 27.42, 27.417]
inst: std_all
operation: std_all(slope_shift)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0417163037672
varray: [-6.287, -6.343, -6.263, -6.248]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.34677502313
varray: [-42.555, -42.739, -41.961, -42.218]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.38875
varray: [63.215, 61.3, 63.07, 61.97]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.914242628263
varray: [63.215, 61.3, 63.07, 61.97]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8845
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.389
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.61375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0300119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0300119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.44271489069
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.5296547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.287, -6.343, -6.263, -6.248]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.555, -42.739, -41.961, -42.218]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.28525
varray: [-6.287, -6.343, -6.263, -6.248]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
dD_mean | -42.36825
varray: [-42.555, -42.739, -41.961, -42.218]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0417163037672
varray: [-6.287, -6.343, -6.263, -6.248]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
dD_stdev | 0.34677502313
varray: [-42.555, -42.739, -41.961, -42.218]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.28525
varray: [-6.287, -6.343, -6.263, -6.248]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.36825
varray: [-42.555, -42.739, -41.961, -42.218]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0417163037672
varray: [-6.287, -6.343, -6.263, -6.248]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.34677502313
varray: [-42.555, -42.739, -41.961, -42.218]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.22965777601
actual_array: [ 1.93 -15.84]
varray: [-6.287, -6.343, -6.263, -6.248]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.28525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.2567912575
actual_array: [ 18.1 -119. ]
varray: [-42.555, -42.739, -41.961, -42.218]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.36825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [77151.0, 77687.0, 78223.0, 78758.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7960833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.482666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0795
varray: [2.114, 2.067, 2.066, 2.071]
inst: mean_all
operation: mean_all(residuals)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: residuals
SAMP_residuals_STD | 0.0231012265187
varray: [2.114, 2.067, 2.066, 2.071]
inst: std_all
operation: std_all(residuals)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0165
varray: [-1.03, -0.889, -1.096, -1.051]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.089347635671
varray: [-1.03, -0.889, -1.096, -1.051]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.6275
varray: [42.98, 42.323, 42.538, 42.669]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.27489937553
varray: [42.98, 42.323, 42.538, 42.669]
inst: std_all
operation: std_all(baseline_shift)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.22275
varray: [28.847, 27.917, 28.246, 27.881]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.447401665173
varray: [28.847, 27.917, 28.246, 27.881]
inst: std_all
operation: std_all(slope_shift)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0894538987412
varray: [-8.742, -8.905, -8.919, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.26561115013
varray: [-63.273, -65.25, -65.947, -65.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.2225
varray: [61.739, 62.172, 59.996, 60.983]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.953936580701
varray: [61.739, 62.172, 59.996, 60.983]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0795
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.6275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.22275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.000738095238095
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.000738095238095
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.0354812864828
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.363404761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-8.742, -8.905, -8.919, -8.934]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-63.273, -65.25, -65.947, -65.953]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.875
varray: [-8.742, -8.905, -8.919, -8.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
dD_mean | -65.10575
varray: [-63.273, -65.25, -65.947, -65.953]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0894538987412
varray: [-8.742, -8.905, -8.919, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
dD_stdev | 1.26561115013
varray: [-63.273, -65.25, -65.947, -65.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.875
varray: [-8.742, -8.905, -8.919, -8.934]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.10575
varray: [-63.273, -65.25, -65.947, -65.953]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0894538987412
varray: [-8.742, -8.905, -8.919, -8.934]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.26561115013
varray: [-63.273, -65.25, -65.947, -65.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.1214025651
actual_array: [ 1.93 -15.84]
varray: [-8.742, -8.905, -8.919, -8.934]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.0753688721
actual_array: [ 18.1 -119. ]
varray: [-63.273, -65.25, -65.947, -65.953]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.10575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [79342.0, 79878.0, 80413.0, 80949.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8268333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.721166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1145
varray: [2.088, 2.127, 2.119, 2.124]
inst: mean_all
operation: mean_all(residuals)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: residuals
SAMP_residuals_STD | 0.0179722007556
varray: [2.088, 2.127, 2.119, 2.124]
inst: std_all
operation: std_all(residuals)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1135
varray: [-1.294, -1.052, -1.179, -0.929]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.157789944758
varray: [-1.294, -1.052, -1.179, -0.929]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.71325
varray: [42.539, 42.779, 42.942, 42.593]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.183913702589
varray: [42.539, 42.779, 42.942, 42.593]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3955
varray: [28.197, 28.424, 28.625, 28.336]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.17928097129
varray: [28.197, 28.424, 28.625, 28.336]
inst: std_all
operation: std_all(slope_shift)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0640383999384
varray: [-6.475, -6.455, -6.363, -6.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.22882016585
varray: [-45.205, -43.726, -42.739, -42.5]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.656
varray: [60.314, 61.866, 61.203, 63.241]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23322044528
varray: [60.314, 61.866, 61.203, 63.241]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1145
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1135
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.71325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3955
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0342619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0342619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.64701842738
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.796904761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.475, -6.455, -6.363, -6.348]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-45.205, -43.726, -42.739, -42.5]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.41025
varray: [-6.475, -6.455, -6.363, -6.348]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
dD_mean | -43.5425
varray: [-45.205, -43.726, -42.739, -42.5]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0640383999384
varray: [-6.475, -6.455, -6.363, -6.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
dD_stdev | 1.22882016585
varray: [-45.205, -43.726, -42.739, -42.5]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.41025
varray: [-6.475, -6.455, -6.363, -6.348]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.5425
varray: [-45.205, -43.726, -42.739, -42.5]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0640383999384
varray: [-6.475, -6.455, -6.363, -6.348]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.22882016585
varray: [-45.205, -43.726, -42.739, -42.5]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.36923422109
actual_array: [ 1.93 -15.84]
varray: [-6.475, -6.455, -6.363, -6.348]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.41025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.6418034516
actual_array: [ 18.1 -119. ]
varray: [-45.205, -43.726, -42.739, -42.5]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.5425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [81535.0, 82071.0, 82607.0, 83144.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7918333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.806916666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10175
varray: [2.133, 2.09, 2.09, 2.094]
inst: mean_all
operation: mean_all(residuals)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: residuals
SAMP_residuals_STD | 0.0209184926162
varray: [2.133, 2.09, 2.09, 2.094]
inst: std_all
operation: std_all(residuals)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.084
varray: [-1.146, -1.11, -1.116, -0.964]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0815352684426
varray: [-1.146, -1.11, -1.116, -0.964]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.584
varray: [43.009, 42.485, 42.454, 42.388]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.286206219359
varray: [43.009, 42.485, 42.454, 42.388]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.0885
varray: [28.515, 28.249, 27.885, 27.705]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.363385470265
varray: [28.515, 28.249, 27.885, 27.705]
inst: std_all
operation: std_all(slope_shift)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0181842422626
varray: [-6.446, -6.426, -6.43, -6.402]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0973665753737
varray: [-43.024, -43.019, -42.815, -42.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.23475
varray: [61.841, 60.135, 60.763, 62.2]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.954156652058
varray: [61.841, 60.135, 60.763, 62.2]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.084
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.584
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.0885
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0215119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0215119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.03410781733
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.375654761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.446, -6.426, -6.43, -6.402]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-43.024, -43.019, -42.815, -42.953]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.426
varray: [-6.446, -6.426, -6.43, -6.402]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
dD_mean | -42.95275
varray: [-43.024, -43.019, -42.815, -42.953]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0181842422626
varray: [-6.446, -6.426, -6.43, -6.402]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
dD_stdev | 0.0973665753737
varray: [-43.024, -43.019, -42.815, -42.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.426
varray: [-6.446, -6.426, -6.43, -6.402]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.95275
varray: [-43.024, -43.019, -42.815, -42.953]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0181842422626
varray: [-6.446, -6.426, -6.43, -6.402]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0973665753737
varray: [-43.024, -43.019, -42.815, -42.953]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.38682085317
actual_array: [ 1.93 -15.84]
varray: [-6.446, -6.426, -6.43, -6.402]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.9462012021
actual_array: [ 18.1 -119. ]
varray: [-43.024, -43.019, -42.815, -42.953]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.95275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [83730.0, 84266.0, 84802.0, 85338.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8045833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.677666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10125
varray: [2.094, 2.106, 2.123, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: residuals
SAMP_residuals_STD | 0.0175
varray: [2.094, 2.106, 2.123, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.91475
varray: [-0.867, -1.147, -0.702, -0.943]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.18464447821
varray: [-0.867, -1.147, -0.702, -0.943]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.296
varray: [42.161, 42.474, 42.35, 42.199]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.144030089449
varray: [42.161, 42.474, 42.35, 42.199]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.5275
varray: [27.583, 28.047, 27.394, 27.086]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.402370061842
varray: [27.583, 28.047, 27.394, 27.086]
inst: std_all
operation: std_all(slope_shift)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0480927922528
varray: [-6.118, -6.044, -6.021, -6.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.348482902498
varray: [-40.757, -40.156, -40.121, -39.964]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.7365
varray: [62.373, 61.864, 64.049, 62.66]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.934861308787
varray: [62.373, 61.864, 64.049, 62.66]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.91475
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.296
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.5275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0210119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0210119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.01007210713
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.8774047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.118, -6.044, -6.021, -6.012]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-40.757, -40.156, -40.121, -39.964]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.04875
varray: [-6.118, -6.044, -6.021, -6.012]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
dD_mean | -40.2495
varray: [-40.757, -40.156, -40.121, -39.964]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0480927922528
varray: [-6.118, -6.044, -6.021, -6.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
dD_stdev | 0.348482902498
varray: [-40.757, -40.156, -40.121, -39.964]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.04875
varray: [-6.118, -6.044, -6.021, -6.012]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.2495
varray: [-40.757, -40.156, -40.121, -39.964]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0480927922528
varray: [-6.118, -6.044, -6.021, -6.012]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.348482902498
varray: [-40.757, -40.156, -40.121, -39.964]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.96557914192
actual_array: [ 1.93 -15.84]
varray: [-6.118, -6.044, -6.021, -6.012]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.04875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -49.7577539261
actual_array: [ 18.1 -119. ]
varray: [-40.757, -40.156, -40.121, -39.964]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.2495
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [85925.0, 86460.0, 86996.0, 87532.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8050833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.389666666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1035
varray: [2.139, 2.092, 2.101, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: residuals
SAMP_residuals_STD | 0.024906491791
varray: [2.139, 2.092, 2.101, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06025
varray: [-1.141, -0.853, -1.16, -1.087]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.141584780256
varray: [-1.141, -0.853, -1.16, -1.087]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.58375
varray: [43.13, 42.258, 42.522, 42.425]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.380138900053
varray: [43.13, 42.258, 42.522, 42.425]
inst: std_all
operation: std_all(baseline_shift)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.021
varray: [28.762, 27.86, 27.8, 27.662]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.50090717703
varray: [28.762, 27.86, 27.8, 27.662]
inst: std_all
operation: std_all(slope_shift)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0520576603393
varray: [-6.33, -6.457, -6.384, -6.393]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.285953813287
varray: [-42.008, -42.65, -42.109, -42.352]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.874
varray: [62.662, 62.254, 61.537, 61.043]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.723296619652
varray: [62.662, 62.254, 61.537, 61.043]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1035
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.58375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.021
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0232619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0232619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.11823280302
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.0149047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.33, -6.457, -6.384, -6.393]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-42.008, -42.65, -42.109, -42.352]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.391
varray: [-6.33, -6.457, -6.384, -6.393]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
dD_mean | -42.27975
varray: [-42.008, -42.65, -42.109, -42.352]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0520576603393
varray: [-6.33, -6.457, -6.384, -6.393]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
dD_stdev | 0.285953813287
varray: [-42.008, -42.65, -42.109, -42.352]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.391
varray: [-6.33, -6.457, -6.384, -6.393]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.27975
varray: [-42.008, -42.65, -42.109, -42.352]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0520576603393
varray: [-6.33, -6.457, -6.384, -6.393]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.285953813287
varray: [-42.008, -42.65, -42.109, -42.352]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.34773944854
actual_array: [ 1.93 -15.84]
varray: [-6.33, -6.457, -6.384, -6.393]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.391
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.1524066893
actual_array: [ 18.1 -119. ]
varray: [-42.008, -42.65, -42.109, -42.352]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.27975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [88118.0, 88654.0, 89190.0, 89726.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8028333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.677416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11175
varray: [2.149, 2.087, 2.101, 2.11]
inst: mean_all
operation: mean_all(residuals)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: residuals
SAMP_residuals_STD | 0.0265753645318
varray: [2.149, 2.087, 2.101, 2.11]
inst: std_all
operation: std_all(residuals)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03175
varray: [-1.155, -1.075, -0.85, -1.047]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129520590384
varray: [-1.155, -1.075, -0.85, -1.047]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.63175
varray: [43.233, 42.394, 42.432, 42.468]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.401970459114
varray: [43.233, 42.394, 42.432, 42.468]
inst: std_all
operation: std_all(baseline_shift)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.0085
varray: [28.652, 27.763, 27.549, 28.07]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.479334608251
varray: [28.652, 27.763, 27.549, 28.07]
inst: std_all
operation: std_all(slope_shift)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0324897419709
varray: [-6.495, -6.532, -6.559, -6.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.210546115297
varray: [-43.437, -43.643, -43.464, -43.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.1675
varray: [62.806, 61.523, 62.623, 61.718]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.640986999348
varray: [62.806, 61.523, 62.623, 61.718]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.63175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.0085
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0315119047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0315119047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.51482202129
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.3084047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.495, -6.532, -6.559, -6.567]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-43.437, -43.643, -43.464, -43.894]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.53825
varray: [-6.495, -6.532, -6.559, -6.567]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
dD_mean | -43.6095
varray: [-43.437, -43.643, -43.464, -43.894]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0324897419709
varray: [-6.495, -6.532, -6.559, -6.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
dD_stdev | 0.210546115297
varray: [-43.437, -43.643, -43.464, -43.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.53825
varray: [-6.495, -6.532, -6.559, -6.567]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.6095
varray: [-43.437, -43.643, -43.464, -43.894]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0324897419709
varray: [-6.495, -6.532, -6.559, -6.567]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.210546115297
varray: [-43.437, -43.643, -43.464, -43.894]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.51216050084
actual_array: [ 1.93 -15.84]
varray: [-6.495, -6.532, -6.559, -6.567]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.53825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.7208290569
actual_array: [ 18.1 -119. ]
varray: [-43.437, -43.643, -43.464, -43.894]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.6095
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [90309.0, 90844.0, 91379.0, 91915.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7945833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.725416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1015
varray: [2.141, 2.094, 2.09, 2.081]
inst: mean_all
operation: mean_all(residuals)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: residuals
SAMP_residuals_STD | 0.0268886593195
varray: [2.141, 2.094, 2.09, 2.081]
inst: std_all
operation: std_all(residuals)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.123
varray: [-0.97, -1.276, -1.176, -1.07]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132206908544
varray: [-0.97, -1.276, -1.176, -1.07]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.49175
varray: [42.689, 42.397, 42.6, 42.281]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.186206650436
varray: [42.689, 42.397, 42.6, 42.281]
inst: std_all
operation: std_all(baseline_shift)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.00075
varray: [28.594, 28.009, 27.937, 27.463]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.463771764988
varray: [28.594, 28.009, 27.937, 27.463]
inst: std_all
operation: std_all(slope_shift)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0195362910161
varray: [-6.363, -6.344, -6.372, -6.391]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.186258780554
varray: [-41.29, -40.871, -41.044, -40.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.10575
varray: [62.404, 60.034, 60.942, 61.043]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.977215900744
varray: [62.404, 60.034, 60.942, 61.043]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1015
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.123
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.49175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.00075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0212619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0212619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.02208996223
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.246654761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.363, -6.344, -6.372, -6.391]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-41.29, -40.871, -41.044, -40.925]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.3675
varray: [-6.363, -6.344, -6.372, -6.391]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
dD_mean | -41.0325
varray: [-41.29, -40.871, -41.044, -40.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0195362910161
varray: [-6.363, -6.344, -6.372, -6.391]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
dD_stdev | 0.186258780554
varray: [-41.29, -40.871, -41.044, -40.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.3675
varray: [-6.363, -6.344, -6.372, -6.391]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
measured_dD_mean | -41.0325
varray: [-41.29, -40.871, -41.044, -40.925]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0195362910161
varray: [-6.363, -6.344, -6.372, -6.391]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.186258780554
varray: [-41.29, -40.871, -41.044, -40.925]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.32149907687
actual_array: [ 1.93 -15.84]
varray: [-6.363, -6.344, -6.372, -6.391]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.3675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -50.68129197
actual_array: [ 18.1 -119. ]
varray: [-41.29, -40.871, -41.044, -40.925]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.0325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [92500.0, 93036.0, 93573.0, 94108.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8048333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.585416666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.101
varray: [2.142, 2.107, 2.072, 2.083]
inst: mean_all
operation: mean_all(residuals)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: residuals
SAMP_residuals_STD | 0.0309946231896
varray: [2.142, 2.107, 2.072, 2.083]
inst: std_all
operation: std_all(residuals)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.015
varray: [-1.041, -1.024, -1.048, -0.947]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0464399253516
varray: [-1.041, -1.024, -1.048, -0.947]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.5075
varray: [42.891, 42.621, 42.268, 42.25]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.307474118152
varray: [42.891, 42.621, 42.268, 42.25]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.91575
varray: [28.598, 27.966, 27.476, 27.623]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.499023296049
varray: [28.598, 27.966, 27.476, 27.623]
inst: std_all
operation: std_all(slope_shift)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0393403948463
varray: [-6.53, -6.54, -6.562, -6.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.254033298342
varray: [-41.836, -41.882, -41.98, -42.393]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.86175
varray: [62.971, 61.86, 60.841, 61.775]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.871759666804
varray: [62.971, 61.86, 60.841, 61.775]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.101
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.015
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.5075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.91575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0207619047619
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0207619047619
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.998054252032
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.0026547619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-6.53, -6.54, -6.562, -6.618]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-41.836, -41.882, -41.98, -42.393]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.5625
varray: [-6.53, -6.54, -6.562, -6.618]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
dD_mean | -42.02275
varray: [-41.836, -41.882, -41.98, -42.393]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0393403948463
varray: [-6.53, -6.54, -6.562, -6.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
dD_stdev | 0.254033298342
varray: [-41.836, -41.882, -41.98, -42.393]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.5625
varray: [-6.53, -6.54, -6.562, -6.618]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
measured_dD_mean | -42.02275
varray: [-41.836, -41.882, -41.98, -42.393]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0393403948463
varray: [-6.53, -6.54, -6.562, -6.618]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.254033298342
varray: [-41.836, -41.882, -41.98, -42.393]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.53923833119
actual_array: [ 1.93 -15.84]
varray: [-6.53, -6.54, -6.562, -6.618]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.5625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -51.8492786212
actual_array: [ 18.1 -119. ]
varray: [-41.836, -41.882, -41.98, -42.393]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.02275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [94693.0, 95228.0, 95764.0, 96300.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8053333333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.601166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.86375
varray: [4.855, 4.87, 4.856, 4.874]
inst: mean_all
operation: mean_all(residuals)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: residuals
SAMP_residuals_STD | 0.00967384790729
varray: [4.855, 4.87, 4.856, 4.874]
inst: std_all
operation: std_all(residuals)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -9.5995
varray: [-9.414, -9.584, -9.761, -9.639]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.144098346046
varray: [-9.414, -9.584, -9.761, -9.639]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 393.6665
varray: [392.433, 394.005, 393.79, 394.438]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.865364855615
varray: [392.433, 394.005, 393.79, 394.438]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | -1.09125
varray: [-0.847, -1.287, -1.021, -1.21]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.197494936644
varray: [-0.847, -1.287, -1.021, -1.21]
inst: std_all
operation: std_all(slope_shift)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0840927067785
varray: [10.42, 10.522, 10.614, 10.575]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.4705933893
varray: [-43.608, -46.133, -46.672, -46.702]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.605
varray: [64.23, 63.784, 63.255, 63.151]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.500427150875
varray: [64.23, 63.784, 63.255, 63.151]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.86375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 9.5995
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 393.6665
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 1.09125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 2.78351190476
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 2.78351190476
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 133.807370951
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.7459047619
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [10.42, 10.522, 10.614, 10.575]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-43.608, -46.133, -46.672, -46.702]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 10.53275
varray: [10.42, 10.522, 10.614, 10.575]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
dD_mean | -45.77875
varray: [-43.608, -46.133, -46.672, -46.702]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0840927067785
varray: [10.42, 10.522, 10.614, 10.575]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
dD_stdev | 1.4705933893
varray: [-43.608, -46.133, -46.672, -46.702]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
measured_d18O_mean | 10.53275
varray: [10.42, 10.522, 10.614, 10.575]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
measured_dD_mean | -45.77875
varray: [-43.608, -46.133, -46.672, -46.702]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0840927067785
varray: [10.42, 10.522, 10.614, 10.575]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.4705933893
varray: [-43.608, -46.133, -46.672, -46.702]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 10.5495154504
actual_array: [ 1.93 -15.84]
varray: [10.42, 10.522, 10.614, 10.575]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 10.53275
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -56.2794304638
actual_array: [ 18.1 -119. ]
varray: [-43.608, -46.133, -46.672, -46.702]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -45.77875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [96884.0, 97420.0, 97957.0, 98493.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 37.0425833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 351.760166667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.08075
varray: [2.13, 2.072, 2.055, 2.066]
inst: mean_all
operation: mean_all(residuals)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: residuals
SAMP_residuals_STD | 0.0335795076398
varray: [2.13, 2.072, 2.055, 2.066]
inst: std_all
operation: std_all(residuals)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03875
varray: [-1.112, -1.001, -1.156, -0.886]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.120928008887
varray: [-1.112, -1.001, -1.156, -0.886]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 44.758
varray: [47.786, 44.297, 43.76, 43.189]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.06874116312
varray: [47.786, 44.297, 43.76, 43.189]
inst: std_all
operation: std_all(baseline_shift)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.70925
varray: [28.17, 27.469, 27.754, 27.444]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.337821821873
varray: [28.17, 27.469, 27.754, 27.444]
inst: std_all
operation: std_all(slope_shift)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: slope_shift
SAMP_d18O_STD | 0.132343933245
varray: [-8.538, -8.781, -8.786, -8.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.078159559023
varray: [-66.286, -66.186, -66.374, -66.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.16725
varray: [62.592, 60.711, 60.432, 60.934]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.97178096812
varray: [62.592, 60.711, 60.432, 60.934]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.08075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 44.758
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.70925
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.000511904761905
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.000511904761905
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.0246079890123
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.308154761905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -13.1008181818
varray: [-8.538, -8.781, -8.786, -8.83]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -98.9549090909
varray: [-66.286, -66.186, -66.374, -66.311]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.73375
varray: [-8.538, -8.781, -8.786, -8.83]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
dD_mean | -66.28925
varray: [-66.286, -66.186, -66.374, -66.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
d18O_stdev | 0.132343933245
varray: [-8.538, -8.781, -8.786, -8.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
dD_stdev | 0.078159559023
varray: [-66.286, -66.186, -66.374, -66.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.73375
varray: [-8.538, -8.781, -8.786, -8.83]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.28925
varray: [-66.286, -66.186, -66.374, -66.311]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.132343933245
varray: [-8.538, -8.781, -8.786, -8.83]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.078159559023
varray: [-66.286, -66.186, -66.374, -66.311]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.9636811822
actual_array: [ 1.93 -15.84]
varray: [-8.538, -8.781, -8.786, -8.83]
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.73375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.4712913177
actual_array: [ 18.1 -119. ]
varray: [-66.286, -66.186, -66.374, -66.311]
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.28925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: [99078.0, 99614.0, 100150.0, 100684.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.8255833333
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.85166666667
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | -75.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.7467050391
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -9.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.92478659
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -75.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.045
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.259
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.3112265304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -9.045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.1535584309
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -70.259
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1515510773
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.4633540671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -68.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.865
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1102364495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.865
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.9384825096
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -68.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.872
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.147
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1180527304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.872
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.6624826344
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -68.147
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.671
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0968371108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1010469909
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.671
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.553
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0611055408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9618675667
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.553
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.485
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0398899212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8816624747
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.815
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.428
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0544058715
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.815
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8144317359
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.428
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.822
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0622221524
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.822
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.884021448
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.13
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.531
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.16690765947
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.13
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 6.59870776566
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 7.531
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.667
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.006
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.76652806753
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.667
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.4153704449
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 15.006
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.79
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.88
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.90387128948
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.79
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.6257284196
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 16.88
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97533442936
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.9632662762
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 18.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.91
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.03786467676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.91
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.4621891275
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 18.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.769
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.08587897386
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.853778694
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 18.769
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.42
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07471285826
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6216245006
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 19.42
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.923
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.303
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05238062704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.923
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.483624563
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 19.303
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.968
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.749
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10262814727
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.968
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 21.0096756072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 19.749
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.996
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.712
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8134
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.2821
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.712
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.13389327097
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 2.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.9660346012
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 19.712
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.0
operation: col(d(18_16)Mean)
inst: col
except: True
dD_actual | 0.0
operation: col(d(D_H)Mean)
inst: col
except: True
DTL_DELTA_18_16O_CALIBRATE | -1.21147496464
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -2.28400616577
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: 0.0
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.658
operation: col(d(18_16)Mean)
inst: col
dD_actual | -96.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.4621556595
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.658
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.764775356
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -96.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.294
operation: col(d(18_16)Mean)
inst: col
dD_actual | -106.125
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.1723206121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.294
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.457026478
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -106.125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.392
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.281748545
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.392
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -130.352666194
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -108.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.493
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.969
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.3945263127
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.493
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -131.990973146
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -109.969
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.298
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.345395404
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -132.379024252
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -110.298
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.589
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.161
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.5017210225
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.589
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -133.396921228
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.161
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.563
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.23
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4726891219
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.563
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -133.478305806
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.23
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.521
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.477
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4257914364
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.521
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -133.769639008
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.477
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.4514735023
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -133.766100548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 3 - 11
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.606
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.978
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.520703419
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.606
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -134.360561818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.978
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.426
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.5367820657
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -9.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.7095059929
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -72.426
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2609790102
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -9.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -83.8455077404
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -69.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.869
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.662
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1147028958
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.869
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0904316111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.662
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.934
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.707
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1872826472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.934
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1435085102
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.707
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1381517385
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8191496825
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1850494241
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.7802266232
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.904
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1537843004
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.904
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.4664831753
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.946
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.447
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2006819859
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.946
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8368419822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.909
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.086
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1593673582
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.909
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.411047303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.086
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.915
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.075
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1660670275
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.915
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3980729499
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -67.075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.794
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.184
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.79773390747
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.794
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.7574174271
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -46.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.688
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.054
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.67937308205
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.688
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.245110871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.054
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.686
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.638
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.67713985893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.686
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.7544444263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.638
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.632
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.264
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.61684283465
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.632
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.3133164207
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.264
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.811
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.26
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.816716304
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.811
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.4880851201
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.26
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.554
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.82788241961
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.8348541941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.554
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.766
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.307
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.76646878377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.766
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.5435209925
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.307
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.83
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83793192365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.83
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.4538800074
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.515
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.744
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.48619928206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.515
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6999833647
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.744
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.466
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.36
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.43148531559
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.466
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.2470604926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.36
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.404
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.36225539883
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.404
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.7198299618
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.407
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.99
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.36560523351
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.407
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.8106504336
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.99
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.435
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.39687035721
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.435
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.0168559208
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.635
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.46275043929
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.3919326742
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.635
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.453
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.345
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.4169693653
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.453
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.0498815469
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.345
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.454
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.494
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.41808597686
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.454
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.2256250571
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.494
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.676
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.023
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.66597374332
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.676
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.029060139
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.023
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.817
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.61460961153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7860858899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.817
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.653
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.64029167743
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.653
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5808552135
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.637
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.62242589246
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.637
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.8403422756
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.208
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.934
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.14339953295
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.208
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.3856258893
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.934
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.182
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.303
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.11436763237
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.182
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.6413698156
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.303
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.097
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.8
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.01945564972
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.097
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.0480880327
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -38.8
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.083
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.936
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.00382308787
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.083
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.2084982165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -38.936
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.805
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.551
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.81001663464
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.6518290881
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.551
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.753
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.592
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.75195283349
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.753
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.7001880406
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.592
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.734
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.597
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.73073721384
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.734
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.7060854738
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.597
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.677
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.76088572597
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.8004444055
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.677
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.887
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.90157878261
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.887
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.0471617904
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.871
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.909
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.88371299764
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.871
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.2535719534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.894
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.90939506353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.894
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.3054693658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.958
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.043
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.98085820341
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.958
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.411623164
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -45.043
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.821
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.839
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0611055408
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.821
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.5812541632
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -63.839
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.979
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2375301674
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.979
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.1100735324
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.997
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1973321512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.1265863454
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.997
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.975
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.174
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2330637212
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.3353554818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.174
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.36798715157
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.2531140332
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -38.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.175
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.04305318743
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.7724489407
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -35.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.278
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.98833922096
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.278
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.7569109385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -34.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.23
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.784
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.93474186605
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.23
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.1317830161
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -33.784
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.433
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.963
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.51124869471
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.433
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.5736707547
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -51.963
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.619
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.002
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.71893844498
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.619
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.978644026
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -54.002
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.07
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.63519257793
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0588491179
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -54.07
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.167
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.68097365192
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1732593226
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -54.167
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.572
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1539708688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -126.804770363
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -105.572
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.694
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.597
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.5023536757
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.694
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -132.731690759
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -110.597
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.738
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.919
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.5514845844
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.738
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -134.290972106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.919
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.754
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.8
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.5693503694
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -13.754
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -135.330099841
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -112.8
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.782
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.628
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.43449988689
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -3.782
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.4093231611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -30.628
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.301
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.916
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.89740972624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -3.301
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.3131221467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -22.916
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.213
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.909
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7991479089
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -3.213
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -26.9458924481
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -20.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.72768476902
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -3.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -25.9362518791
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -20.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.712
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.15
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.58956019888
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.712
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.2814217127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -38.15
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.833
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.129
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.72467019772
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.833
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6156257853
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.129
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.828
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.126
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.71908713992
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.828
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6120873253
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.126
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.858
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.227
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.75258548673
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.858
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.7312154766
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.227
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.355
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.74573712244
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.3180829357
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -48.355
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.762
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.481
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.87861389815
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.762
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.6461848992
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -49.481
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.83618265885
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.6815694986
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -49.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.817
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.868
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.94002753398
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.817
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.1026462312
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -49.868
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.314
operation: col(d(18_16)Mean)
inst: col
dD_actual | -105.474
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.1946528433
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.314
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -126.689180672
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -105.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.624
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.152
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.5408024271
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.624
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -133.386305848
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -111.152
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.737
operation: col(d(18_16)Mean)
inst: col
dD_actual | -112.662
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.6669795335
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.737
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -135.167330684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -112.662
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.741
operation: col(d(18_16)Mean)
inst: col
dD_actual | -113.445
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -17.6714459797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -14.741
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -136.090868727
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -113.445
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.007
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.298
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.80234904865
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.6354723824
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.298
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.505
operation: col(d(18_16)Mean)
inst: col
dD_actual | -32.211
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.24181004522
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.505
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -40.2764505218
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -32.211
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.443
operation: col(d(18_16)Mean)
inst: col
dD_actual | -30.494
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.17258012846
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.443
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -38.2512719505
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -30.494
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.43
operation: col(d(18_16)Mean)
inst: col
dD_actual | -29.799
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.15806417817
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.43
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -37.4315287315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -29.799
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.717
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.306
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9449779385
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.717
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7731011348
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -62.306
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.917
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.497
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1683002507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.917
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.5368430224
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.497
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.947
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.14
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.2017985975
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.947
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.2952529358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.517
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.220780994
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.7399194014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.517
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.704
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.664
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.5806273064
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.704
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.7081912028
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -37.664
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.56
operation: col(d(18_16)Mean)
inst: col
dD_actual | -35.224
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.41983524167
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.56
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.8302437864
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -35.224
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.505
operation: col(d(18_16)Mean)
inst: col
dD_actual | -34.058
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.35842160584
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.505
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.4549623571
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -34.058
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.503
operation: col(d(18_16)Mean)
inst: col
dD_actual | -33.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.35618838272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -5.503
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.078706117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -33.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.852
operation: col(d(18_16)Mean)
inst: col
dD_actual | -22.734
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62927425675
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.852
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -29.0984555772
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -22.734
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.601
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.55334467063
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.7620972072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -21.601
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | -21.115
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.53212905098
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.1888666972
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -21.115
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.797
operation: col(d(18_16)Mean)
inst: col
dD_actual | -20.976
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56786062092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -4.797
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -27.0249180534
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -20.976
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.961
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.032
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1008195987
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -7.961
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0140286254
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -54.032
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.049
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.718
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.199081416
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.049
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.1821297567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -56.718
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.121
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.895
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2794774484
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.121
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5703855391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -57.895
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.063
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.889
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2147139779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.063
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5633086192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -57.889
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.459
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.371
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.42366903467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.459
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.6190081378
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -44.371
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.415
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.374538126
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.415
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.0951113911
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.401
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.86
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.35890556415
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.401
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.8368038157
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.86
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.789
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.3376899445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7530602638
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.789
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.75
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.3801211838
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7070602846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.75
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.814
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.39017068785
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7825474299
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.814
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.415
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.866
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.374538126
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.415
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.8438807355
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.866
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.436
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.933
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.39798696877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.436
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.9229063408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.933
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.434
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.811
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.39575374565
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.434
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.77900897
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.811
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.472
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.913
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.43818498495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.472
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.8993166079
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.913
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.461
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.018
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.42590225779
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.461
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.0231627057
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.018
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.382
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.693
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.3376899445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.382
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6398295458
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.693
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.345
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.695
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.29637531676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.345
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6421885191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.695
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.331
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.28074275491
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.331
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.5077270414
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.334
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.769
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.28409258959
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.334
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7294705309
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.769
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.338
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.28855903583
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.338
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6645987653
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.287
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.555
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.23161184624
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.287
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.4770603886
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.555
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.343
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.29414209363
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.343
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6940859315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.263
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.20481316878
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.263
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.7764453208
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.248
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.218
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.18806399538
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.248
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.0795733889
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.218
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.742
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.273
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9728932276
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.742
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.9136647216
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -63.273
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.905
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.25
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1549009119
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.2455098208
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.25
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.919
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.947
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1705334738
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.919
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0676120131
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.947
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.934
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1872826472
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.934
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.074688933
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -65.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.205
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.44153481964
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.6027000006
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -45.205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.726
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.41920258842
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.8582392511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.726
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.739
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.31647432485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.6940859315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.739
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.348
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.5
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.29972515144
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.348
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.4121886231
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.5
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.446
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.024
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.40915308438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.446
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.0302396256
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.024
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.426
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.019
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.38682085317
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.426
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.0243421924
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.019
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.43
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.815
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.39128729941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.43
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7837269166
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.815
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.402
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.953
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.36002217571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.402
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.9464960737
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.953
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.118
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.04290449249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.118
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.356343399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.044
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.156
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.96027523701
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.044
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6474719247
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.156
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.021
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.121
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.93459317112
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.021
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.6061898921
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.121
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -39.964
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.92454366707
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -49.4210104887
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -39.964
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.33
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.008
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.27962614335
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.33
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.8318811932
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.008
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.65
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.42143581154
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.58911162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.65
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.384
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.109
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.33992316762
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.384
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.9510093445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.109
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.393
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.352
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.34997267166
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.393
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.2376245995
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.352
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.495
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.46386705085
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.495
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.5173676104
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.532
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.643
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.50518167859
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.532
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.7603418595
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.643
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.464
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.53533019073
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.5492137499
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.464
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.567
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.54426308321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.567
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.0563930077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.363
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.29
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.31647432485
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.363
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.9850097813
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.29
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.344
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.29525870519
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.344
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.4908048767
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.372
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.32652382889
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.372
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.6948560664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.391
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.925
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.34773944854
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.391
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.5544971555
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -40.925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.53
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.836
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.50294845547
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.53
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.6290094901
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.836
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.54
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.882
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.51411457108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.54
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.6832658758
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.882
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.98
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.53868002541
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.7988555671
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -41.98
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.618
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.393
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.6012102728
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -6.618
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.2859835519
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -42.393
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.608
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.423617497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 10.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.7190598269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -43.608
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.522
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.5375118762
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 10.522
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.6972636082
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -46.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.614
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.672
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.6402401397
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 10.614
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3330069104
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -46.672
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 10.575
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.702
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 10.5966922889
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: 10.575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3683915098
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -46.702
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:08:47
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.538
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.286
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7451044692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.538
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.4674579861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.286
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.781
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.186
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0164410784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.781
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.3495093215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.186
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.786
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.374
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0220241362
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.571252811
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.374
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8134
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.2821
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.606
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -111.978
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -13.1008181818
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -98.9549090909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.83
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.311
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0711550449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 1.0849564946
measure_array: [ 2.8134 -13.10081818]
inst: fit_type1
actual_mean: -8.83
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.895566583107
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.4969451523
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 1.93644088588
measure_array: [ 17.2821 -98.95490909]
inst: fit_type1
actual_mean: -66.311
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.847826470393
tarray: []
source_col: d(D_H)Mean