## Summary
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
230620_23-083_23-084_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5073 | -76.0049 |  | 24.9144 |  | 0.41235 | 2.9247
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27811 | 3.8007
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37803 | 5.4314
4 |  | EV | -10.1402 | -74.957 |  |  |  | 0.59466 | 0.76994
5 |  | CPER4-3040-2 | -11.1234 | -99.5929 |  |  |  | 0.023258 | 1.2266
6 |  | CPER4-2030-4 | -10.9806 | -97.1154 |  |  |  | 0.020662 | 0.17134
7 |  | CPER4-E-3 | -6.302 | -78.8208 |  |  |  | 0.13357 | 1.0814
8 |  | CPER4-E-5 | -6.2431 | -81.4343 |  |  |  | 0.025461 | 0.26886
9 |  | CPER4-B-1 | -3.9486 | -68.361 |  |  |  | 0.10325 | 0.7989
10 |  | CPER4-2030-3 | -10.6741 | -95.1896 |  |  |  | 0.18872 | 1.4019
11 |  | CPER4-deep-1 | -15.7184 | -127.5074 |  |  |  | 0.13635 | 1.6902
12 |  | CPER4-3040-3 | -10.2337 | -103.081 |  |  |  | 0.17739 | 1.3257
13 |  | EV | -10.293 | -76.0307 |  |  |  | 0.029136 | 1.4045
14 |  | CPER4-1020-1 | -8.7345 | -90.296 |  |  |  | 0.063348 | 0.71269
15 |  | CPER4-3040-1 | -8.5311 | -94.4146 |  |  |  | 0.024576 | 0.2006
16 |  | CPER4-deep-3 | -10.0425 | -97.6774 |  |  |  | 0.029904 | 0.219
17 |  | CPER4-010-2 | -5.5683 | -80.4294 |  |  |  | 0.13145 | 0.94461
18 |  | CPER4-010-3 | -6.0446 | -78.6314 |  |  |  | 0.0069462 | 0.15804
19 |  | CPER4-B-3 | -9.369 | -89.5108 |  |  |  | 0.098181 | 0.50259
20 |  | CPER4-1020-3 | -9.0452 | -90.8328 |  |  |  | 0.017711 | 0.12973
21 |  | CPER4-deep-2 | -10.5806 | -100.969 |  |  |  | 0.031138 | 0.43005
22 |  | EV | -10.297 | -75.9596 |  |  |  | 0.023678 | 1.3763
23 |  | CPER4-1020-2 | -9.7211 | -93.5744 |  |  |  | 0.021777 | 0.88272
24 |  | CPER4-B-5 | -5.0716 | -80.4546 |  |  |  | 0.13961 | 0.81899
25 |  | CPER4-C-5 | -6.6124 | -86.5672 |  |  |  | 0.053804 | 0.43576
26 |  | CPER4-D-2 | 4.0998 | -58.0359 | -5.5818 | 104.4707 | 56.5536 | 0.080297 | 0.94025 |  | 1.0566
27 |  | CPER4-C-2 | -7.2179 | -87.4186 |  |  |  | 0.11966 | 0.86366
28 |  | CPER4-C-4 | -4.9343 | -79.5676 |  | 122.9072 |  | 0.053451 | 0.33518 |  | -5.7076
29 |  | CPER4-D-1 | 1.6315 | -60.2656 | -4.1852 | 83.4725 | 34.288 | 0.040898 | 0.68017 |  | 0.1276
30 |  | CPER4-A-2 | -3.9538 | -64.396 | -1.5417 | 367.9337 |  | 0.027379 | 0.21952 |  | -9.4319
31 |  | EV | -10.1162 | -73.7246 |  |  |  | 0.17327 | 0.64451
32 |  | CPER4-010-1 | -5.0417 | -76.2459 |  |  |  | 0.13634 | 0.16488
33 |  | CPER4-D-4 | 8.9035 | -57.4892 | -7.4733 | 133.1979 | 88.5121 | 0.093991 | 0.095472 |  | 1.7656
34 |  | CPER4-2030-1 | -9.4502 | -94.3724 |  |  |  | 0.27768 | 1.5309
35 |  | CPER4-E-2 | -3.5339 | -70.6816 |  |  |  | 0.17625 | 1.2554
36 |  | CPER4-A-1 | -4.4718 | -71.9036 |  |  |  | 0.014248 | 0.10694
37 |  | CLBJ6-1020-2 | -0.091591 | -17.9613 |  |  |  | 0.12153 | 3.1123
38 |  | CLBJ6-010-3 | 2.8119 | -7.4429 |  |  |  | 0.069604 | 0.60507
39 |  | EV | -9.8685 | -69.9833 |  |  |  | 0.3386 | 3.2622

## Detail
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -10.263 | -76.45 | -11.6773 | -84.7465 | -1.548 | 213.516 | 2023/06/20 14:13:07
1 |  | EV | -9.296 | -69.611 | -10.6699 | -77.2577 | -0.921 | 70.138 | 2023/06/20 14:22:04
1 |  | EV | -9.122 | -68.276 | -10.4886 | -75.7959 | -1.189 | 52.902 | 2023/06/20 14:31:00
1 |  | EV | -9.047 | -67.77 | -10.4105 | -75.2418 | -0.974 | 49.926 | 2023/06/20 14:39:56
1 |  | EV | -8.981 | -67.344 | -10.3418 | -74.7753 | -1.106 | 49.258 | 2023/06/20 14:48:52
1 |  | EV | -8.99 | -67.106 | -10.3511 | -74.5147 | -0.807 | 48.292 | 2023/06/20 14:57:49
1 |  | EV | -8.906 | -67.01 | -10.2636 | -74.4096 | -1.02 | 46.918 | 2023/06/20 15:06:46
1 |  | EV | -8.968 | -67.101 | -10.3282 | -74.5092 | -1.058 | 47.112 | 2023/06/20 15:15:43
1 |  | EV | -8.895 | -66.894 | -10.2522 | -74.2826 | -0.937 | 46.506 | 2023/06/20 15:24:39
1 |  | EV | -8.931 | -67.107 | -10.2897 | -74.5158 | -1.193 | 45.629 | 2023/06/20 15:33:36
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.5073 | -76.0049 |  | 24.9144 |  | 0.41235 | 2.9247
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.061 | 7.559 | 1.1617 | 7.2442 | -1.106 | 41.163 | 2023/06/20 15:43:25
2 |  | PZ | 2.672 | 15.176 | 1.7982 | 15.5849 | -1.093 | 44.215 | 2023/06/20 15:52:22
2 |  | PZ | 2.748 | 16.981 | 1.8774 | 17.5614 | -1.26 | 44.289 | 2023/06/20 16:01:19
2 |  | PZ | 2.833 | 17.871 | 1.9659 | 18.5359 | -1.101 | 44.456 | 2023/06/20 16:10:15
2 |  | PZ | 2.906 | 18.932 | 2.042 | 19.6977 | -0.975 | 44.189 | 2023/06/20 16:19:11
2 |  | PZ | 2.953 | 19.27 | 2.091 | 20.0678 | -1.026 | 43.91 | 2023/06/20 16:28:09
2 |  | PZ | 2.937 | 19.498 | 2.0743 | 20.3175 | -1.023 | 43.548 | 2023/06/20 16:37:05
2 |  | PZ | 2.932 | 19.571 | 2.0691 | 20.3974 | -1.21 | 43.419 | 2023/06/20 16:46:02
2 |  | PZ | 2.962 | 19.767 | 2.1003 | 20.6121 | -1.196 | 43.109 | 2023/06/20 16:54:59
2 |  | PZ | 2.981 | 20.104 | 2.1201 | 20.9811 | -1.059 | 42.758 | 2023/06/20 17:03:55
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.27811 | 3.8007
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.315 | -93.399 | -14.8569 | -103.3058 | -0.715 | 41.241 | 2023/06/20 17:13:43
3 |  | UT2 | -14.148 | -104.413 | -15.7247 | -115.3662 | -0.901 | 42.343 | 2023/06/20 17:22:39
3 |  | UT2 | -14.282 | -107.493 | -15.8643 | -118.7388 | -1.023 | 42.397 | 2023/06/20 17:31:35
3 |  | UT2 | -14.393 | -109.085 | -15.9799 | -120.4821 | -0.991 | 42.311 | 2023/06/20 17:40:33
3 |  | UT2 | -14.45 | -109.784 | -16.0393 | -121.2475 | -0.9 | 42.026 | 2023/06/20 17:49:29
3 |  | UT2 | -14.428 | -110.069 | -16.0164 | -121.5596 | -0.985 | 42.122 | 2023/06/20 17:58:25
3 |  | UT2 | -14.498 | -110.492 | -16.0893 | -122.0228 | -0.801 | 41.566 | 2023/06/20 18:07:22
3 |  | UT2 | -14.544 | -110.988 | -16.1372 | -122.5659 | -0.973 | 41.476 | 2023/06/20 18:16:19
3 |  | UT2 | -14.548 | -111.294 | -16.1414 | -122.901 | -0.922 | 41.338 | 2023/06/20 18:25:15
3 |  | UT2 | -13.981 | -110.298 | -15.5507 | -121.8103 | -0.611 | 30.23 | 2023/06/20 18:34:11
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37803 | 5.4314
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -7.101 | -68.573 | -8.3832 | -76.1211 | -0.761 | 37.675 | 2023/06/20 18:44:00
4 |  | EV | -9.072 | -69.02 | -10.4366 | -76.6105 | -1.139 | 41.691 | 2023/06/20 18:52:56
4 |  | EV | -8.984 | -67.72 | -10.3449 | -75.187 | -1.019 | 41.211 | 2023/06/20 19:01:53
4 |  | EV | -9.007 | -67.462 | -10.3688 | -74.9045 | -0.866 | 40.927 | 2023/06/20 19:10:49
4 |  | EV | -9.018 | -67.599 | -10.3803 | -75.0545 | -0.914 | 40.871 | 2023/06/20 19:19:46
4 |  | EV | -8.913 | -66.988 | -10.2709 | -74.3855 | -0.95 | 41.058 | 2023/06/20 19:28:42
4 |  | EV | -8.926 | -66.826 | -10.2845 | -74.2081 | -1.043 | 41.125 | 2023/06/20 19:37:39
4 |  | EV | -8.988 | -67.367 | -10.349 | -74.8005 | -0.717 | 40.781 | 2023/06/20 19:46:35
4 |  | EV | -8.931 | -66.67 | -10.2897 | -74.0373 | -0.893 | 40.542 | 2023/06/20 19:55:32
4 |  | EV | -8.935 | -66.874 | -10.2938 | -74.2607 | -0.857 | 40.834 | 2023/06/20 20:04:29
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.1402 | -74.957 |  |  |  | 0.59466 | 0.76994
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | CPER4-3040-2 | -9.722 | -88.281 | -11.1137 | -97.7015 | -1.016 | 41.073 | 2023/06/20 20:14:19
5 |  | CPER4-3040-2 | -9.726 | -90.156 | -11.1179 | -99.7547 | -0.856 | 40.703 | 2023/06/20 20:23:15
5 |  | CPER4-3040-2 | -9.712 | -90.437 | -11.1033 | -100.0624 | -0.928 | 41.121 | 2023/06/20 20:32:12
5 |  | CPER4-3040-2 | -9.765 | -91.159 | -11.1585 | -100.853 | -0.986 | 40.922 | 2023/06/20 20:41:10
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | CPER4-3040-2 | -11.1234 | -99.5929 |  |  |  | 0.023258 | 1.2266
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | CPER4-2030-4 | -9.591 | -87.97 | -10.9772 | -97.361 | -0.98 | 38.964 | 2023/06/20 20:50:59
6 |  | CPER4-2030-4 | -9.592 | -87.668 | -10.9783 | -97.0303 | -0.793 | 40.04 | 2023/06/20 20:59:56
6 |  | CPER4-2030-4 | -9.572 | -87.57 | -10.9574 | -96.923 | -0.933 | 40.287 | 2023/06/20 21:08:52
6 |  | CPER4-2030-4 | -9.622 | -87.775 | -11.0095 | -97.1474 | -0.871 | 39.926 | 2023/06/20 21:17:49
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | CPER4-2030-4 | -10.9806 | -97.1154 |  |  |  | 0.020662 | 0.17134
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | CPER4-E-3 | -5.295 | -72.605 | -6.5017 | -80.5362 | -1.025 | 40.021 | 2023/06/20 21:27:40
7 |  | CPER4-E-3 | -5.09 | -70.909 | -6.2882 | -78.679 | -0.933 | 40.796 | 2023/06/20 21:36:36
7 |  | CPER4-E-3 | -5.032 | -70.372 | -6.2277 | -78.091 | -1.055 | 40.973 | 2023/06/20 21:45:32
7 |  | CPER4-E-3 | -4.996 | -70.268 | -6.1902 | -77.9771 | -0.959 | 40.944 | 2023/06/20 21:54:30
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | CPER4-E-3 | -6.302 | -78.8208 |  |  |  | 0.13357 | 1.0814
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | CPER4-E-5 | -5.029 | -73.085 | -6.2246 | -81.0618 | -0.988 | 40.695 | 2023/06/20 22:04:17
8 |  | CPER4-E-5 | -5.083 | -73.619 | -6.2809 | -81.6465 | -0.922 | 40.858 | 2023/06/20 22:13:12
8 |  | CPER4-E-5 | -5.029 | -73.336 | -6.2246 | -81.3366 | -1.027 | 40.925 | 2023/06/20 22:22:08
8 |  | CPER4-E-5 | -5.046 | -73.661 | -6.2423 | -81.6925 | -1.072 | 40.676 | 2023/06/20 22:31:05
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | CPER4-E-5 | -6.2431 | -81.4343 |  |  |  | 0.025461 | 0.26886
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | CPER4-B-1 | -2.991 | -62.632 | -4.1014 | -69.6156 | -0.984 | 40.915 | 2023/06/20 22:40:53
9 |  | CPER4-B-1 | -2.802 | -61.178 | -3.9045 | -68.0235 | -1.042 | 40.973 | 2023/06/20 22:49:49
9 |  | CPER4-B-1 | -2.832 | -61.347 | -3.9358 | -68.2085 | -1.129 | 40.89 | 2023/06/20 22:58:45
9 |  | CPER4-B-1 | -2.752 | -60.788 | -3.8524 | -67.5964 | -0.998 | 41.106 | 2023/06/20 23:07:41
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | CPER4-B-1 | -3.9486 | -68.361 |  |  |  | 0.10325 | 0.7989
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | CPER4-2030-3 | -9.028 | -83.964 | -10.3907 | -92.9744 | -0.922 | 40.545 | 2023/06/20 23:17:29
10 |  | CPER4-2030-3 | -9.328 | -86.133 | -10.7033 | -95.3494 | -1.039 | 40.486 | 2023/06/20 23:26:26
10 |  | CPER4-2030-3 | -9.388 | -86.825 | -10.7658 | -96.1072 | -0.975 | 39.987 | 2023/06/20 23:35:22
10 |  | CPER4-2030-3 | -9.456 | -87.026 | -10.8366 | -96.3273 | -1.036 | 40.011 | 2023/06/20 23:44:18
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | CPER4-2030-3 | -10.6741 | -95.1896 |  |  |  | 0.18872 | 1.4019
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | CPER4-deep-1 | -13.944 | -113.079 | -15.5121 | -124.8556 | -0.845 | 41.913 | 2023/06/20 23:54:06
11 |  | CPER4-deep-1 | -14.16 | -115.602 | -15.7372 | -127.6183 | -1.103 | 41.652 | 2023/06/21 00:03:02
11 |  | CPER4-deep-1 | -14.237 | -116.622 | -15.8174 | -128.7352 | -1.057 | 41.626 | 2023/06/21 00:11:59
11 |  | CPER4-deep-1 | -14.227 | -116.7 | -15.807 | -128.8206 | -0.839 | 41.815 | 2023/06/21 00:20:56
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | CPER4-deep-1 | -15.7184 | -127.5074 |  |  |  | 0.13635 | 1.6902
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | CPER4-3040-3 | -9.136 | -95.044 | -10.5032 | -105.1071 | -0.807 | 39.627 | 2023/06/21 00:30:46
12 |  | CPER4-3040-3 | -8.848 | -93.231 | -10.2032 | -103.1218 | -1.009 | 39.702 | 2023/06/21 00:39:42
12 |  | CPER4-3040-3 | -8.774 | -92.432 | -10.1261 | -102.2469 | -0.99 | 39.734 | 2023/06/21 00:48:38
12 |  | CPER4-3040-3 | -8.751 | -92.068 | -10.1021 | -101.8483 | -0.899 | 39.891 | 2023/06/21 00:57:35
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | CPER4-3040-3 | -10.2337 | -103.081 |  |  |  | 0.17739 | 1.3257
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.891 | -70.504 | -10.248 | -78.2355 | -1.211 | 41.176 | 2023/06/21 01:07:23
13 |  | EV | -8.953 | -68.371 | -10.3126 | -75.8999 | -1.003 | 40.303 | 2023/06/21 01:16:20
13 |  | EV | -8.95 | -67.711 | -10.3095 | -75.1772 | -0.9 | 40.309 | 2023/06/21 01:25:15
13 |  | EV | -8.943 | -67.376 | -10.3022 | -74.8103 | -1.044 | 40.373 | 2023/06/21 01:34:11
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.293 | -76.0307 |  |  |  | 0.029136 | 1.4045
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | CPER4-1020-1 | -7.531 | -80.487 | -8.8312 | -89.167 | -0.875 | 39.587 | 2023/06/21 01:44:01
14 |  | CPER4-1020-1 | -7.415 | -81.6 | -8.7103 | -90.3858 | -0.751 | 39.904 | 2023/06/21 01:52:58
14 |  | CPER4-1020-1 | -7.419 | -82.036 | -8.7145 | -90.8632 | -0.837 | 39.37 | 2023/06/21 02:01:54
14 |  | CPER4-1020-1 | -7.388 | -81.949 | -8.6822 | -90.7679 | -0.922 | 39.305 | 2023/06/21 02:10:51
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | CPER4-1020-1 | -8.7345 | -90.296 |  |  |  | 0.063348 | 0.71269
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | CPER4-3040-1 | -7.272 | -85.093 | -8.5613 | -94.2106 | -1.008 | 40.035 | 2023/06/21 02:20:38
15 |  | CPER4-3040-1 | -7.254 | -85.123 | -8.5426 | -94.2435 | -1.029 | 39.768 | 2023/06/21 02:29:33
15 |  | CPER4-3040-1 | -7.228 | -85.412 | -8.5155 | -94.5599 | -1.014 | 39.511 | 2023/06/21 02:38:29
15 |  | CPER4-3040-1 | -7.218 | -85.489 | -8.5051 | -94.6443 | -1.005 | 39.543 | 2023/06/21 02:47:26
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | CPER4-3040-1 | -8.5311 | -94.4146 |  |  |  | 0.024576 | 0.2006
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | CPER4-deep-3 | -8.65 | -87.993 | -9.9969 | -97.3862 | -0.898 | 39.61 | 2023/06/21 02:57:13
16 |  | CPER4-deep-3 | -8.707 | -88.197 | -10.0563 | -97.6095 | -0.847 | 39.731 | 2023/06/21 03:06:10
16 |  | CPER4-deep-3 | -8.701 | -88.335 | -10.0501 | -97.7606 | -0.934 | 39.424 | 2023/06/21 03:15:05
16 |  | CPER4-deep-3 | -8.717 | -88.511 | -10.0667 | -97.9534 | -0.976 | 39.601 | 2023/06/21 03:24:01
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | CPER4-deep-3 | -10.0425 | -97.6774 |  |  |  | 0.029904 | 0.219
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | CPER4-010-2 | -4.587 | -73.885 | -5.7641 | -81.9378 | -1.007 | 40.111 | 2023/06/21 03:33:48
17 |  | CPER4-010-2 | -4.392 | -72.346 | -5.561 | -80.2525 | -1.125 | 40.865 | 2023/06/21 03:42:46
17 |  | CPER4-010-2 | -4.314 | -71.983 | -5.4797 | -79.8551 | -0.895 | 39.996 | 2023/06/21 03:51:41
17 |  | CPER4-010-2 | -4.303 | -71.816 | -5.4683 | -79.6722 | -1.144 | 40.596 | 2023/06/21 04:00:37
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | CPER4-010-2 | -5.5683 | -80.4294 |  |  |  | 0.13145 | 0.94461
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | CPER4-010-3 | -4.853 | -71.096 | -6.0412 | -78.8838 | -1.131 | 47.434 | 2023/06/21 04:10:25
18 |  | CPER4-010-3 | -4.848 | -70.741 | -6.036 | -78.4951 | -1.05 | 40.045 | 2023/06/21 04:19:21
18 |  | CPER4-010-3 | -4.862 | -70.794 | -6.0506 | -78.5531 | -0.8 | 37.769 | 2023/06/21 04:28:17
18 |  | CPER4-010-3 | -4.862 | -70.831 | -6.0506 | -78.5936 | -0.772 | 37.277 | 2023/06/21 04:37:14
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | CPER4-010-3 | -6.0446 | -78.6314 |  |  |  | 0.0069462 | 0.15804
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | CPER4-B-3 | -7.902 | -80.099 | -9.2177 | -88.7422 | -0.977 | 41.972 | 2023/06/21 04:47:02
19 |  | CPER4-B-3 | -8.073 | -80.776 | -9.3958 | -89.4835 | -0.94 | 41.369 | 2023/06/21 04:55:58
19 |  | CPER4-B-3 | -8.11 | -81.158 | -9.4344 | -89.9018 | -1.032 | 41.32 | 2023/06/21 05:04:53
19 |  | CPER4-B-3 | -8.104 | -81.171 | -9.4281 | -89.916 | -0.846 | 41.36 | 2023/06/21 05:13:50
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | CPER4-B-3 | -9.369 | -89.5108 |  |  |  | 0.098181 | 0.50259
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | CPER4-1020-3 | -7.756 | -81.831 | -9.0656 | -90.6387 | -0.7 | 38.032 | 2023/06/21 05:23:38
20 |  | CPER4-1020-3 | -7.747 | -81.993 | -9.0562 | -90.8161 | -0.842 | 37.99 | 2023/06/21 05:32:35
20 |  | CPER4-1020-3 | -7.722 | -82.092 | -9.0301 | -90.9245 | -0.999 | 38.085 | 2023/06/21 05:41:31
20 |  | CPER4-1020-3 | -7.721 | -82.117 | -9.0291 | -90.9519 | -0.896 | 38.535 | 2023/06/21 05:50:29
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | CPER4-1020-3 | -9.0452 | -90.8328 |  |  |  | 0.017711 | 0.12973
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | CPER4-deep-2 | -9.176 | -90.646 | -10.5449 | -100.2912 | -0.971 | 40.596 | 2023/06/21 06:00:18
21 |  | CPER4-deep-2 | -9.192 | -91.328 | -10.5616 | -101.038 | -0.941 | 40.363 | 2023/06/21 06:09:15
21 |  | CPER4-deep-2 | -9.233 | -91.624 | -10.6043 | -101.3621 | -1.009 | 39.663 | 2023/06/21 06:18:11
21 |  | CPER4-deep-2 | -9.24 | -91.462 | -10.6116 | -101.1847 | -0.957 | 39.9 | 2023/06/21 06:27:07
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | CPER4-deep-2 | -10.5806 | -100.969 |  |  |  | 0.031138 | 0.43005
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.953 | -70.45 | -10.3126 | -78.1764 | -0.955 | 40.838 | 2023/06/21 06:36:52
22 |  | EV | -8.903 | -67.977 | -10.2605 | -75.4684 | -0.957 | 40.179 | 2023/06/21 06:45:49
22 |  | EV | -8.952 | -67.903 | -10.3115 | -75.3874 | -1.065 | 40.055 | 2023/06/21 06:54:44
22 |  | EV | -8.944 | -67.372 | -10.3032 | -74.806 | -0.9 | 39.911 | 2023/06/21 07:03:40
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.297 | -75.9596 |  |  |  | 0.023678 | 1.3763
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | CPER4-1020-2 | -8.408 | -83.204 | -9.7448 | -92.1422 | -0.812 | 39.193 | 2023/06/21 07:13:26
23 |  | CPER4-1020-2 | -8.398 | -84.754 | -9.7344 | -93.8394 | -1.065 | 38.839 | 2023/06/21 07:22:22
23 |  | CPER4-1020-2 | -8.36 | -85.044 | -9.6948 | -94.157 | -1.097 | 38.709 | 2023/06/21 07:31:19
23 |  | CPER4-1020-2 | -8.375 | -85.046 | -9.7104 | -94.1592 | -0.922 | 38.075 | 2023/06/21 07:40:16
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | CPER4-1020-2 | -9.7211 | -93.5744 |  |  |  | 0.021777 | 0.88272
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | CPER4-B-5 | -4.125 | -73.704 | -5.2828 | -81.7396 | -0.979 | 38.892 | 2023/06/21 07:50:02
24 |  | CPER4-B-5 | -3.892 | -72.473 | -5.0401 | -80.3916 | -1.07 | 38.983 | 2023/06/21 07:58:59
24 |  | CPER4-B-5 | -3.864 | -72.04 | -5.0109 | -79.9175 | -1.022 | 38.896 | 2023/06/21 08:07:55
24 |  | CPER4-B-5 | -3.808 | -71.905 | -4.9526 | -79.7696 | -1.042 | 39.099 | 2023/06/21 08:16:51
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | CPER4-B-5 | -5.0716 | -80.4546 |  |  |  | 0.13961 | 0.81899
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | CPER4-C-5 | -5.325 | -77.475 | -6.533 | -85.8688 | -1.162 | 39.923 | 2023/06/21 08:26:39
25 |  | CPER4-C-5 | -5.404 | -78.193 | -6.6153 | -86.6551 | -0.969 | 39.937 | 2023/06/21 08:35:35
25 |  | CPER4-C-5 | -5.447 | -78.415 | -6.6601 | -86.8982 | -0.823 | 39.68 | 2023/06/21 08:44:31
25 |  | CPER4-C-5 | -5.429 | -78.368 | -6.6413 | -86.8467 | -0.888 | 39.606 | 2023/06/21 08:53:26
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | CPER4-C-5 | -6.6124 | -86.5672 |  |  |  | 0.053804 | 0.43576
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | CPER4-D-2 | 4.778 | -53.362 | 3.9922 | -59.4649 | -5.706 | 148.623 | 2023/06/21 09:03:14
26 |  | CPER4-D-2 | 4.86 | -52.12 | 4.0776 | -58.1049 | -5.546 | 145.307 | 2023/06/21 09:12:09
26 |  | CPER4-D-2 | 4.928 | -51.442 | 4.1485 | -57.3625 | -5.564 | 146.16 | 2023/06/21 09:21:05
26 |  | CPER4-D-2 | 4.959 | -51.304 | 4.1808 | -57.2114 | -5.511 | 146.214 | 2023/06/21 09:30:02
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | CPER4-D-2 | 4.0998 | -58.0359 | -5.5818 | 104.4707 | 56.5536 | 0.080297 | 0.94025 |  | 1.0566
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | CPER4-C-2 | -5.81 | -77.628 | -7.0382 | -86.0364 | -0.955 | 43.071 | 2023/06/21 09:39:51
27 |  | CPER4-C-2 | -6 | -79.07 | -7.2362 | -87.6154 | -1.03 | 41.499 | 2023/06/21 09:48:47
27 |  | CPER4-C-2 | -6.039 | -79.317 | -7.2768 | -87.8859 | -0.928 | 41.285 | 2023/06/21 09:57:42
27 |  | CPER4-C-2 | -6.081 | -79.546 | -7.3206 | -88.1366 | -0.845 | 41.088 | 2023/06/21 10:06:38
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | CPER4-C-2 | -7.2179 | -87.4186 |  |  |  | 0.11966 | 0.86366
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | CPER4-C-4 | -3.869 | -72.2 | -5.0161 | -80.0927 | -1.344 | 162.089 | 2023/06/21 10:16:26
28 |  | CPER4-C-4 | -3.779 | -71.692 | -4.9224 | -79.5364 | -1.336 | 163.534 | 2023/06/21 10:25:22
28 |  | CPER4-C-4 | -3.753 | -71.543 | -4.8953 | -79.3733 | -1.424 | 167.378 | 2023/06/21 10:34:18
28 |  | CPER4-C-4 | -3.761 | -71.447 | -4.9036 | -79.2681 | -1.521 | 167.049 | 2023/06/21 10:43:14
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | CPER4-C-4 | -4.9343 | -79.5676 |  | 122.9072 |  | 0.053451 | 0.33518 |  | -5.7076
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | CPER4-D-1 | 2.457 | -55.103 | 1.5742 | -61.3713 | -4.166 | 125.694 | 2023/06/21 10:52:58
29 |  | CPER4-D-1 | 2.527 | -53.845 | 1.6472 | -59.9938 | -4.242 | 124.367 | 2023/06/21 11:01:55
29 |  | CPER4-D-1 | 2.554 | -53.621 | 1.6753 | -59.7485 | -4.11 | 126.409 | 2023/06/21 11:10:52
29 |  | CPER4-D-1 | 2.51 | -53.804 | 1.6294 | -59.9489 | -4.223 | 125.841 | 2023/06/21 11:19:48
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | CPER4-D-1 | 1.6315 | -60.2656 | -4.1852 | 83.4725 | 34.288 | 0.040898 | 0.68017 |  | 0.1276
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | CPER4-A-2 | -2.831 | -58.165 | -3.9348 | -64.7242 | -1.568 | 414.532 | 2023/06/21 11:29:33
30 |  | CPER4-A-2 | -2.839 | -57.824 | -3.9431 | -64.3508 | -1.571 | 408.554 | 2023/06/21 11:38:29
30 |  | CPER4-A-2 | -2.89 | -57.835 | -3.9962 | -64.3629 | -1.385 | 408.218 | 2023/06/21 11:47:25
30 |  | CPER4-A-2 | -2.837 | -57.637 | -3.941 | -64.146 | -1.643 | 408.852 | 2023/06/21 11:56:21
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | CPER4-A-2 | -3.9538 | -64.396 | -1.5417 | 367.9337 |  | 0.027379 | 0.21952 |  | -9.4319
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.512 | -65.481 | -9.8532 | -72.7353 | -0.983 | 45.616 | 2023/06/21 12:06:07
31 |  | EV | -8.794 | -66.368 | -10.1469 | -73.7066 | -1.107 | 42.55 | 2023/06/21 12:15:02
31 |  | EV | -8.859 | -66.788 | -10.2147 | -74.1665 | -0.869 | 41.346 | 2023/06/21 12:23:58
31 |  | EV | -8.893 | -66.901 | -10.2501 | -74.2902 | -0.871 | 41.048 | 2023/06/21 12:32:55
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.1162 | -73.7246 |  |  |  | 0.17327 | 0.64451
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | CPER4-010-1 | -4.087 | -68.511 | -5.2432 | -76.0532 | -0.941 | 40.629 | 2023/06/21 12:42:41
32 |  | CPER4-010-1 | -3.887 | -68.633 | -5.0349 | -76.1868 | -0.975 | 40.68 | 2023/06/21 12:51:36
32 |  | CPER4-010-1 | -3.82 | -68.699 | -4.9651 | -76.259 | -0.935 | 40.958 | 2023/06/21 13:00:33
32 |  | CPER4-010-1 | -3.78 | -68.905 | -4.9234 | -76.4846 | -0.952 | 40.707 | 2023/06/21 13:09:30
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | CPER4-010-1 | -5.0417 | -76.2459 |  |  |  | 0.13634 | 0.16488
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | CPER4-D-4 | 9.367 | -51.641 | 8.773 | -57.5804 | -7.426 | 173.809 | 2023/06/21 13:19:16
33 |  | CPER4-D-4 | 9.475 | -51.452 | 8.8855 | -57.3734 | -7.588 | 177.064 | 2023/06/21 13:28:12
33 |  | CPER4-D-4 | 9.55 | -51.636 | 8.9636 | -57.5749 | -7.474 | 175.424 | 2023/06/21 13:37:08
33 |  | CPER4-D-4 | 9.577 | -51.502 | 8.9918 | -57.4282 | -7.405 | 174.916 | 2023/06/21 13:46:02
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | CPER4-D-4 | 8.9035 | -57.4892 | -7.4733 | 133.1979 | 88.5121 | 0.093991 | 0.095472 |  | 1.7656
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | CPER4-2030-1 | -7.725 | -83.014 | -9.0333 | -91.9341 | -0.812 | 36.673 | 2023/06/21 13:55:50
34 |  | CPER4-2030-1 | -8.152 | -85.457 | -9.4781 | -94.6092 | -0.745 | 36.475 | 2023/06/21 14:04:45
34 |  | CPER4-2030-1 | -8.294 | -86.197 | -9.626 | -95.4195 | -0.834 | 38.809 | 2023/06/21 14:13:40
34 |  | CPER4-2030-1 | -8.33 | -86.295 | -9.6635 | -95.5268 | -0.942 | 38.044 | 2023/06/21 14:22:36
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | CPER4-2030-1 | -9.4502 | -94.3724 |  |  |  | 0.27768 | 1.5309
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | CPER4-E-2 | -2.701 | -65.383 | -3.7993 | -72.628 | -0.976 | 40.583 | 2023/06/21 14:32:22
35 |  | CPER4-E-2 | -2.4 | -63.543 | -3.4857 | -70.6132 | -1.003 | 40.271 | 2023/06/21 14:41:17
35 |  | CPER4-E-2 | -2.389 | -62.965 | -3.4743 | -69.9803 | -1.034 | 40.428 | 2023/06/21 14:50:14
35 |  | CPER4-E-2 | -2.295 | -62.531 | -3.3764 | -69.505 | -1.053 | 40.306 | 2023/06/21 14:59:09
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | CPER4-E-2 | -3.5339 | -70.6816 |  |  |  | 0.17625 | 1.2554
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | CPER4-A-1 | -3.326 | -64.841 | -4.4504 | -72.0345 | -0.902 | 40.106 | 2023/06/21 15:08:52
36 |  | CPER4-A-1 | -3.359 | -64.61 | -4.4848 | -71.7816 | -0.827 | 39.604 | 2023/06/21 15:17:49
36 |  | CPER4-A-1 | -3.35 | -64.779 | -4.4754 | -71.9666 | -0.795 | 39.621 | 2023/06/21 15:26:43
36 |  | CPER4-A-1 | -3.351 | -64.656 | -4.4765 | -71.8319 | -1.012 | 40.053 | 2023/06/21 15:35:39
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | CPER4-A-1 | -4.4718 | -71.9036 |  |  |  | 0.014248 | 0.10694
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | CLBJ6-1020-2 | 0.678 | -19.961 | -0.27911 | -22.8905 | -1.061 | 43.753 | 2023/06/21 15:45:21
37 |  | CLBJ6-1020-2 | 0.907 | -15.051 | -0.040544 | -17.514 | -0.95 | 43.474 | 2023/06/21 15:54:17
37 |  | CLBJ6-1020-2 | 0.902 | -13.768 | -0.045752 | -16.1091 | -0.818 | 43.371 | 2023/06/21 16:03:13
37 |  | CLBJ6-1020-2 | 0.945 | -13.058 | -0.00095555 | -15.3316 | -1.203 | 43.652 | 2023/06/21 16:12:08
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | CLBJ6-1020-2 | -0.091591 | -17.9613 |  |  |  | 0.12153 | 3.1123
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | CLBJ6-010-3 | 3.546 | -6.701 | 2.7087 | -8.3706 | -0.846 | 40.105 | 2023/06/21 16:21:52
38 |  | CLBJ6-010-3 | 3.647 | -5.869 | 2.814 | -7.4596 | -0.81 | 39.937 | 2023/06/21 16:30:48
38 |  | CLBJ6-010-3 | 3.69 | -5.487 | 2.8588 | -7.0413 | -1.049 | 39.815 | 2023/06/21 16:39:42
38 |  | CLBJ6-010-3 | 3.697 | -5.358 | 2.866 | -6.9001 | -1.127 | 39.789 | 2023/06/21 16:48:39
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | CLBJ6-010-3 | 2.8119 | -7.4429 |  |  |  | 0.069604 | 0.60507
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | EV | -8.028 | -58.202 | -9.3489 | -64.7647 | -0.988 | 40.542 | 2023/06/21 16:58:24
39 |  | EV | -8.603 | -63.513 | -9.948 | -70.5803 | -0.948 | 40.295 | 2023/06/21 17:07:20
39 |  | EV | -8.746 | -65.011 | -10.0969 | -72.2206 | -1.036 | 40.159 | 2023/06/21 17:16:15
39 |  | EV | -8.73 | -65.145 | -10.0803 | -72.3674 | -0.909 | 40.194 | 2023/06/21 17:25:11
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | EV | -9.8685 | -69.9833 |  |  |  | 0.3386 | 3.2622

## Source
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-24764 | 2023/06/20 14:13:07 | 1-01 | 1 | -10.263 | -76.45 | 15457 | -1 | 1 | EV | standard | H2O | 1687292292.4 | 0.252 | 0.842 | 89.492 | 0.001 | 0.008 | 2.173 | 213.516 | 31.58 | 2.597 | 72.904 | 0.859 | -1.548 | 0 | 0 | 1 | 44.25 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-24764 | 2023/06/20 14:22:04 | 1-01 | 2 | -9.296 | -69.611 | 12855 | -1 | 1 | EV | standard | H2O | 1687292829.68 | 0.19 | 0.89 | 76.898 | 0 | 0.003 | 1.899 | 70.138 | 29.142 | 2.19 | 63.63 | 0.857 | -0.921 | 0 | 0 | 1 | 43.438 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-24764 | 2023/06/20 14:31:00 | 1-01 | 3 | -9.122 | -68.276 | 12743 | -1 | 1 | EV | standard | H2O | 1687293365.81 | 0.215 | 0.757 | 72.905 | 0 | -0.003 | 1.791 | 52.902 | 29.291 | 2.154 | 63.355 | 0.856 | -1.189 | 0 | 0 | 1 | 43.125 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-24764 | 2023/06/20 14:39:56 | 1-01 | 4 | -9.047 | -67.77 | 12428 | 0 | 1 | EV | standard | H2O | 1687293901.57 | 0.19 | 0.825 | 71.417 | 0.001 | 0.002 | 1.766 | 49.926 | 28.784 | 2.127 | 62.025 | 0.857 | -0.974 | 0 | 0 | 1 | 42.875 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-24764 | 2023/06/20 14:48:52 | 1-01 | 5 | -8.981 | -67.344 | 12682 | 0 | 1 | EV | standard | H2O | 1687294437.6 | 0.204 | 0.741 | 68.412 | 0.001 | 0.001 | 1.688 | 49.258 | 29.366 | 2.154 | 63.651 | 0.856 | -1.106 | 0 | 0 | 1 | 42.875 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-24764 | 2023/06/20 14:57:49 | 1-01 | 6 | -8.99 | -67.106 | 12663 | 0 | 1 | EV | standard | H2O | 1687294974.49 | 0.212 | 0.734 | 67.338 | 0 | -0.001 | 1.663 | 48.292 | 28.965 | 2.174 | 65.293 | 0.856 | -0.807 | 0 | 0 | 1 | 42.812 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-24764 | 2023/06/20 15:06:46 | 1-01 | 7 | -8.906 | -67.01 | 12450 | 0 | 1 | EV | standard | H2O | 1687295511.87 | 0.196 | 0.794 | 68.578 | 0.001 | 0.002 | 1.706 | 46.918 | 28.574 | 2.118 | 62.677 | 0.856 | -1.02 | 0 | 0 | 1 | 42.875 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-24764 | 2023/06/20 15:15:43 | 1-01 | 8 | -8.968 | -67.101 | 12619 | 0 | 1 | EV | standard | H2O | 1687296048.42 | 0.185 | 0.752 | 71.206 | 0.001 | 0 | 1.76 | 47.112 | 29.172 | 2.157 | 62.533 | 0.856 | -1.058 | 0 | 0 | 1 | 42.938 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-24764 | 2023/06/20 15:24:39 | 1-01 | 9 | -8.895 | -66.894 | 12717 | 0 | 1 | EV | standard | H2O | 1687296584.52 | 0.214 | 0.849 | 74.53 | 0.001 | 0.001 | 1.841 | 46.506 | 29.386 | 2.167 | 64.044 | 0.856 | -0.937 | 0 | 0 | 1 | 42.938 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-24764 | 2023/06/20 15:33:36 | 1-01 | 10 | -8.931 | -67.107 | 12504 | 0 | 1 | EV | standard | H2O | 1687297121.42 | 0.21 | 0.718 | 71.152 | 0 | 0 | 1.756 | 45.629 | 29.132 | 2.123 | 61.885 | 0.856 | -1.193 | 0 | 0 | 1 | 43 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-24765 | 2023/06/20 15:43:25 | 1-02 | 1 | 2.061 | 7.559 | 12532 | -1 | 1 | PZ | standard | H2O | 1687297711.23 | 0.206 | 0.808 | 65.026 | 0 | -0.004 | 1.607 | 41.163 | 29.318 | 2.252 | 64.595 | 0.857 | -1.106 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-24765 | 2023/06/20 15:52:22 | 1-02 | 2 | 2.672 | 15.176 | 12734 | -1 | 1 | PZ | standard | H2O | 1687298247.76 | 0.218 | 0.913 | 66.983 | 0 | -0.008 | 1.642 | 44.215 | 29.999 | 2.295 | 66.326 | 0.856 | -1.093 | 0 | 0 | 1 | 43.125 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-24765 | 2023/06/20 16:01:19 | 1-02 | 3 | 2.748 | 16.981 | 12680 | -1 | 1 | PZ | standard | H2O | 1687298784.17 | 0.196 | 0.862 | 68.937 | 0.001 | -0.005 | 1.705 | 44.289 | 29.769 | 2.295 | 65.594 | 0.856 | -1.26 | 0 | 0 | 1 | 43.188 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-24765 | 2023/06/20 16:10:15 | 1-02 | 4 | 2.833 | 17.871 | 12742 | 0 | 1 | PZ | standard | H2O | 1687299321.08 | 0.184 | 0.778 | 66.774 | 0 | -0.005 | 1.643 | 44.456 | 29.961 | 2.302 | 66.246 | 0.857 | -1.101 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-24765 | 2023/06/20 16:19:11 | 1-02 | 5 | 2.906 | 18.932 | 12838 | 0 | 1 | PZ | standard | H2O | 1687299856.76 | 0.204 | 0.874 | 66.245 | 0 | 0.001 | 1.622 | 44.189 | 30.421 | 2.327 | 67.03 | 0.857 | -0.975 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-24765 | 2023/06/20 16:28:09 | 1-02 | 6 | 2.953 | 19.27 | 12811 | 0 | 1 | PZ | standard | H2O | 1687300394.17 | 0.203 | 0.788 | 68.489 | 0 | -0.002 | 1.693 | 43.91 | 29.972 | 2.323 | 66.879 | 0.856 | -1.026 | 0 | 0 | 1 | 43.438 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-24765 | 2023/06/20 16:37:05 | 1-02 | 7 | 2.937 | 19.498 | 12810 | 0 | 1 | PZ | standard | H2O | 1687300930.93 | 0.21 | 0.833 | 64.437 | 0.001 | -0.007 | 1.581 | 43.548 | 30.574 | 2.309 | 65.627 | 0.856 | -1.023 | 0 | 0 | 1 | 43.438 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-24765 | 2023/06/20 16:46:02 | 1-02 | 8 | 2.932 | 19.571 | 12819 | 0 | 1 | PZ | standard | H2O | 1687301468.19 | 0.212 | 0.846 | 63.002 | 0.001 | -0.003 | 1.555 | 43.419 | 30.451 | 2.32 | 66.108 | 0.856 | -1.21 | 0 | 0 | 1 | 43.062 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-24765 | 2023/06/20 16:54:59 | 1-02 | 9 | 2.962 | 19.767 | 12837 | 0 | 1 | PZ | standard | H2O | 1687302004.58 | 0.23 | 0.844 | 67.263 | 0.001 | -0.002 | 1.654 | 43.109 | 30.219 | 2.323 | 66.205 | 0.857 | -1.196 | 0 | 0 | 1 | 42.938 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-24765 | 2023/06/20 17:03:55 | 1-02 | 10 | 2.981 | 20.104 | 12710 | 0 | 1 | PZ | standard | H2O | 1687302540.27 | 0.208 | 0.84 | 68.051 | 0 | 0 | 1.683 | 42.758 | 29.544 | 2.309 | 66.325 | 0.856 | -1.059 | 0 | 0 | 1 | 43 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-24766 | 2023/06/20 17:13:43 | 1-03 | 1 | -13.315 | -93.399 | 12815 | -1 | 1 | UT2 | standard | H2O | 1687303128.77 | 0.2 | 0.808 | 64.88 | 0.002 | -0.003 | 1.593 | 41.241 | 29.936 | 2.147 | 64.332 | 0.856 | -0.715 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-24766 | 2023/06/20 17:22:39 | 1-03 | 2 | -14.148 | -104.413 | 12828 | -1 | 1 | UT2 | standard | H2O | 1687303664.41 | 0.198 | 0.813 | 71.347 | 0.002 | 0.008 | 1.764 | 42.343 | 30.247 | 2.134 | 63.225 | 0.856 | -0.901 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-24766 | 2023/06/20 17:31:35 | 1-03 | 3 | -14.282 | -107.493 | 12809 | -1 | 1 | UT2 | standard | H2O | 1687304200.92 | 0.188 | 0.814 | 72.145 | 0.001 | 0.008 | 1.782 | 42.397 | 30.178 | 2.109 | 61.963 | 0.856 | -1.023 | 0 | 0 | 1 | 43.062 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-24766 | 2023/06/20 17:40:33 | 1-03 | 4 | -14.393 | -109.085 | 12867 | 0 | 1 | UT2 | standard | H2O | 1687304738.61 | 0.219 | 0.732 | 64.094 | 0.001 | 0.005 | 1.573 | 42.311 | 30.028 | 2.115 | 63.671 | 0.856 | -0.991 | 0 | 0 | 1 | 42.875 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-24766 | 2023/06/20 17:49:29 | 1-03 | 5 | -14.45 | -109.784 | 12815 | 0 | 1 | UT2 | standard | H2O | 1687305274.24 | 0.183 | 0.809 | 66.557 | 0.001 | 0.002 | 1.633 | 42.026 | 29.92 | 2.123 | 63.368 | 0.857 | -0.9 | 0 | 0 | 1 | 43 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-24766 | 2023/06/20 17:58:25 | 1-03 | 6 | -14.428 | -110.069 | 12857 | 0 | 1 | UT2 | standard | H2O | 1687305810.68 | 0.194 | 0.773 | 65.6 | 0.001 | 0.005 | 1.611 | 42.122 | 29.99 | 2.123 | 63.048 | 0.856 | -0.985 | 0 | 0 | 1 | 43.125 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-24766 | 2023/06/20 18:07:22 | 1-03 | 7 | -14.498 | -110.492 | 12675 | 0 | 1 | UT2 | standard | H2O | 1687306347.2 | 0.204 | 0.86 | 67.968 | 0.001 | 0.002 | 1.671 | 41.566 | 29.597 | 2.085 | 63.056 | 0.856 | -0.801 | 0 | 0 | 1 | 43.188 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-24766 | 2023/06/20 18:16:19 | 1-03 | 8 | -14.544 | -110.988 | 12654 | 0 | 1 | UT2 | standard | H2O | 1687306884.49 | 0.194 | 0.766 | 65.859 | 0.001 | 0.003 | 1.628 | 41.476 | 29.528 | 2.087 | 62.495 | 0.857 | -0.973 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-24766 | 2023/06/20 18:25:15 | 1-03 | 9 | -14.548 | -111.294 | 12654 | 0 | 1 | UT2 | standard | H2O | 1687307419.98 | 0.188 | 0.748 | 68.455 | 0.001 | 0.002 | 1.691 | 41.338 | 29.674 | 2.073 | 62.118 | 0.857 | -0.922 | 0 | 0 | 1 | 43.25 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-24766 | 2023/06/20 18:34:11 | 1-03 | 10 | -13.981 | -110.298 | 6927 | 0 | 1 | UT2 | standard | H2O | 1687307926.32 | 0.24 | 0.955 | 653.112 | 0.003 | 0.001 | -27.56 | 30.23 | 13.204 | 1.177 | 40.637 | 0.855 | -0.611 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-24767 | 2023/06/20 18:44:00 | 1-04 | 1 | -7.101 | -68.573 | 11044 | -1 | 1 | EV | standard | H2O | 1687308553.81 | 0.421 | 1.018 | 82.874 | 0.01 | 0.019 | 2.306 | 37.675 | 24.895 | 1.888 | 57.625 | 0.855 | -0.761 | 0 | 0 | 1 | 43.312 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-24767 | 2023/06/20 18:52:56 | 1-04 | 2 | -9.072 | -69.02 | 12869 | -1 | 1 | EV | standard | H2O | 1687309082.1 | 0.202 | 0.848 | 67.536 | 0.001 | 0.001 | 1.67 | 41.691 | 30.177 | 2.18 | 64.036 | 0.857 | -1.139 | 0 | 0 | 1 | 43.375 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-24767 | 2023/06/20 19:01:53 | 1-04 | 3 | -8.984 | -67.72 | 12656 | -1 | 1 | EV | standard | H2O | 1687309618.65 | 0.194 | 0.712 | 69.37 | 0 | 0 | 1.716 | 41.211 | 29.701 | 2.152 | 63.65 | 0.856 | -1.019 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-24767 | 2023/06/20 19:10:49 | 1-04 | 4 | -9.007 | -67.462 | 12602 | 0 | 1 | EV | standard | H2O | 1687310154.61 | 0.197 | 0.864 | 64.291 | 0 | 0.001 | 1.59 | 40.927 | 29.444 | 2.14 | 63.468 | 0.855 | -0.866 | 0 | 0 | 1 | 43.5 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-24767 | 2023/06/20 19:19:46 | 1-04 | 5 | -9.018 | -67.599 | 12610 | 0 | 1 | EV | standard | H2O | 1687310691.98 | 0.196 | 0.803 | 71.132 | 0.001 | -0.003 | 1.768 | 40.871 | 29.622 | 2.152 | 63.643 | 0.856 | -0.914 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-24767 | 2023/06/20 19:28:42 | 1-04 | 6 | -8.913 | -66.988 | 12656 | 0 | 1 | EV | standard | H2O | 1687311227.21 | 0.192 | 0.816 | 69.572 | 0.001 | 0 | 1.722 | 41.058 | 29.571 | 2.144 | 63.771 | 0.856 | -0.95 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-24767 | 2023/06/20 19:37:39 | 1-04 | 7 | -8.926 | -66.826 | 12647 | 0 | 1 | EV | standard | H2O | 1687311763.91 | 0.188 | 0.816 | 65.839 | 0 | 0.002 | 1.626 | 41.125 | 29.327 | 2.152 | 63.935 | 0.856 | -1.043 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-24767 | 2023/06/20 19:46:35 | 1-04 | 8 | -8.988 | -67.367 | 12623 | 0 | 1 | EV | standard | H2O | 1687312300.56 | 0.204 | 0.758 | 68.339 | 0.001 | 0.002 | 1.698 | 40.781 | 29.322 | 2.155 | 65.389 | 0.856 | -0.717 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-24767 | 2023/06/20 19:55:32 | 1-04 | 9 | -8.931 | -66.67 | 12478 | 0 | 1 | EV | standard | H2O | 1687312838.19 | 0.2 | 0.753 | 65.325 | 0.001 | 0.001 | 1.615 | 40.542 | 29.075 | 2.125 | 63.292 | 0.855 | -0.893 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-24767 | 2023/06/20 20:04:29 | 1-04 | 10 | -8.935 | -66.874 | 12606 | 0 | 1 | EV | standard | H2O | 1687313374.49 | 0.189 | 0.75 | 68.666 | 0.001 | 0.002 | 1.687 | 40.834 | 29.257 | 2.144 | 64.181 | 0.856 | -0.857 | 0 | 0 | 1 | 43.438 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-24768 | 2023/06/20 20:14:19 | 1-05 | 1 | -9.722 | -88.281 | 12769 | -1 | 1 | CPER4-3040-2 | 23-083 | H2O | 1687313964.19 | 0.199 | 0.715 | 66.169 | 0.002 | -0.003 | 1.648 | 41.073 | 29.952 | 2.134 | 63.654 | 0.856 | -1.016 | 0 | 0 | 1 | 43.438 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-24768 | 2023/06/20 20:23:15 | 1-05 | 2 | -9.726 | -90.156 | 12505 | -1 | 1 | CPER4-3040-2 | 23-083 | H2O | 1687314500.89 | 0.207 | 0.781 | 68.499 | 0.001 | 0.003 | 1.687 | 40.703 | 28.891 | 2.089 | 62.839 | 0.855 | -0.856 | 0 | 0 | 1 | 43.5 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-24768 | 2023/06/20 20:32:12 | 1-05 | 3 | -9.712 | -90.437 | 12755 | -1 | 1 | CPER4-3040-2 | 23-083 | H2O | 1687315037.49 | 0.194 | 0.811 | 68.29 | 0.001 | 0.003 | 1.674 | 41.121 | 29.84 | 2.128 | 63.787 | 0.856 | -0.928 | 0 | 0 | 1 | 43.5 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-24768 | 2023/06/20 20:41:10 | 1-05 | 4 | -9.765 | -91.159 | 12655 | 0 | 1 | CPER4-3040-2 | 23-083 | H2O | 1687315575.32 | 0.189 | 0.811 | 66.665 | 0.001 | 0.003 | 1.65 | 40.922 | 29.501 | 2.126 | 62.889 | 0.855 | -0.986 | 0 | 0 | 1 | 43.562 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-24769 | 2023/06/20 20:50:59 | 1-06 | 1 | -9.591 | -87.97 | 11693 | -1 | 1 | CPER4-2030-4 | 23-083 | H2O | 1687316165.7 | 0.206 | 0.776 | 63.467 | 0.001 | 0.002 | 1.588 | 38.964 | 26.845 | 1.953 | 58.959 | 0.856 | -0.98 | 0 | 0 | 1 | 43.562 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-24769 | 2023/06/20 20:59:56 | 1-06 | 2 | -9.592 | -87.668 | 12273 | -1 | 1 | CPER4-2030-4 | 23-083 | H2O | 1687316701.65 | 0.206 | 0.873 | 66.591 | 0.001 | -0.001 | 1.655 | 40.04 | 28.431 | 2.051 | 61.892 | 0.856 | -0.793 | 0 | 0 | 1 | 43.562 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-24769 | 2023/06/20 21:08:52 | 1-06 | 3 | -9.572 | -87.57 | 12344 | -1 | 1 | CPER4-2030-4 | 23-083 | H2O | 1687317238.27 | 0.19 | 0.791 | 67.103 | 0 | 0.004 | 1.66 | 40.287 | 28.723 | 2.059 | 61.892 | 0.856 | -0.933 | 0 | 0 | 1 | 43.562 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-24769 | 2023/06/20 21:17:49 | 1-06 | 4 | -9.622 | -87.775 | 12205 | 0 | 1 | CPER4-2030-4 | 23-083 | H2O | 1687317774.56 | 0.207 | 0.825 | 68.011 | 0.001 | 0.003 | 1.69 | 39.926 | 28.252 | 2.04 | 61.66 | 0.856 | -0.871 | 0 | 0 | 1 | 43.562 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-24770 | 2023/06/20 21:27:40 | 1-07 | 1 | -5.295 | -72.605 | 12148 | -1 | 1 | CPER4-E-3 | 23-083 | H2O | 1687318365.69 | 0.205 | 0.756 | 64.693 | 0.001 | -0.001 | 1.609 | 40.021 | 28.116 | 2.067 | 61.017 | 0.856 | -1.025 | 0 | 0 | 1 | 43.5 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-24770 | 2023/06/20 21:36:36 | 1-07 | 2 | -5.09 | -70.909 | 12547 | -1 | 1 | CPER4-E-3 | 23-083 | H2O | 1687318901.45 | 0.194 | 0.759 | 64.04 | 0.001 | 0.001 | 1.582 | 40.796 | 29.163 | 2.129 | 63.516 | 0.856 | -0.933 | 0 | 0 | 1 | 43.5 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-24770 | 2023/06/20 21:45:32 | 1-07 | 3 | -5.032 | -70.372 | 12592 | -1 | 1 | CPER4-E-3 | 23-083 | H2O | 1687319438.12 | 0.202 | 0.711 | 65.354 | 0 | 0.001 | 1.618 | 40.973 | 29.201 | 2.122 | 63.2 | 0.855 | -1.055 | 0 | 0 | 1 | 43.562 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-24770 | 2023/06/20 21:54:30 | 1-07 | 4 | -4.996 | -70.268 | 12630 | 0 | 1 | CPER4-E-3 | 23-083 | H2O | 1687319975.3 | 0.201 | 0.735 | 63.284 | 0.001 | 0.001 | 1.556 | 40.944 | 29.403 | 2.138 | 63.674 | 0.856 | -0.959 | 0 | 0 | 1 | 43.562 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-24771 | 2023/06/20 22:04:17 | 1-08 | 1 | -5.029 | -73.085 | 12589 | -1 | 1 | CPER4-E-5 | 23-083 | H2O | 1687320562.7 | 0.182 | 0.758 | 62.416 | 0.001 | -0.003 | 1.537 | 40.695 | 29.447 | 2.121 | 63.204 | 0.855 | -0.988 | 0 | 0 | 1 | 43.688 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-24771 | 2023/06/20 22:13:12 | 1-08 | 2 | -5.083 | -73.619 | 12681 | -1 | 1 | CPER4-E-5 | 23-083 | H2O | 1687321097.56 | 0.184 | 0.866 | 66.253 | 0 | 0 | 1.628 | 40.858 | 29.485 | 2.147 | 63.52 | 0.856 | -0.922 | 0 | 0 | 1 | 43.75 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-24771 | 2023/06/20 22:22:08 | 1-08 | 3 | -5.029 | -73.336 | 12644 | -1 | 1 | CPER4-E-5 | 23-083 | H2O | 1687321633.77 | 0.202 | 0.871 | 65.933 | 0.001 | 0.002 | 1.632 | 40.925 | 29.276 | 2.138 | 63.499 | 0.855 | -1.027 | 0 | 0 | 1 | 43.688 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-24771 | 2023/06/20 22:31:05 | 1-08 | 4 | -5.046 | -73.661 | 12572 | 0 | 1 | CPER4-E-5 | 23-083 | H2O | 1687322170.56 | 0.211 | 0.823 | 64.692 | 0.001 | 0.002 | 1.598 | 40.676 | 29.233 | 2.129 | 62.972 | 0.856 | -1.072 | 0 | 0 | 1 | 43.688 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-24772 | 2023/06/20 22:40:53 | 1-09 | 1 | -2.991 | -62.632 | 12659 | -1 | 1 | CPER4-B-1 | 23-083 | H2O | 1687322758.4 | 0.177 | 0.765 | 70.087 | 0 | 0.003 | 1.73 | 40.915 | 29.428 | 2.156 | 63.904 | 0.856 | -0.984 | 0 | 0 | 1 | 43.625 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-24772 | 2023/06/20 22:49:49 | 1-09 | 2 | -2.802 | -61.178 | 12712 | -1 | 1 | CPER4-B-1 | 23-083 | H2O | 1687323294.88 | 0.199 | 0.781 | 71.642 | 0 | 0 | 1.761 | 40.973 | 29.781 | 2.166 | 63.886 | 0.855 | -1.042 | 0 | 0 | 1 | 43.5 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-24772 | 2023/06/20 22:58:45 | 1-09 | 3 | -2.832 | -61.347 | 12672 | -1 | 1 | CPER4-B-1 | 23-083 | H2O | 1687323830.86 | 0.196 | 0.773 | 63.523 | 0.001 | 0.001 | 1.565 | 40.89 | 29.718 | 2.15 | 63.888 | 0.855 | -1.129 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-24772 | 2023/06/20 23:07:41 | 1-09 | 4 | -2.752 | -60.788 | 12828 | 0 | 1 | CPER4-B-1 | 23-083 | H2O | 1687324366.61 | 0.203 | 0.67 | 66.916 | 0.002 | -0.002 | 1.651 | 41.106 | 30.275 | 2.188 | 64.485 | 0.856 | -0.998 | 0 | 0 | 1 | 43.438 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-24773 | 2023/06/20 23:17:29 | 1-10 | 1 | -9.028 | -83.964 | 12436 | -1 | 1 | CPER4-2030-3 | 23-083 | H2O | 1687324954.67 | 0.192 | 0.818 | 63.287 | 0 | -0.001 | 1.564 | 40.545 | 28.715 | 2.1 | 62.846 | 0.856 | -0.922 | 0 | 0 | 1 | 43.5 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-24773 | 2023/06/20 23:26:26 | 1-10 | 2 | -9.328 | -86.133 | 12378 | -1 | 1 | CPER4-2030-3 | 23-083 | H2O | 1687325491.24 | 0.224 | 0.866 | 67.147 | 0.001 | 0.002 | 1.662 | 40.486 | 28.789 | 2.074 | 61.292 | 0.856 | -1.039 | 0 | 0 | 1 | 43.375 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-24773 | 2023/06/20 23:35:22 | 1-10 | 3 | -9.388 | -86.825 | 12148 | -1 | 1 | CPER4-2030-3 | 23-083 | H2O | 1687326027.9 | 0.206 | 0.817 | 64.59 | 0.002 | 0.002 | 1.598 | 39.987 | 27.85 | 2.043 | 61.602 | 0.855 | -0.975 | 0 | 0 | 1 | 43.25 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-24773 | 2023/06/20 23:44:18 | 1-10 | 4 | -9.456 | -87.026 | 12135 | 0 | 1 | CPER4-2030-3 | 23-083 | H2O | 1687326563.7 | 0.193 | 0.737 | 62.841 | 0.001 | 0.004 | 1.561 | 40.011 | 28.005 | 2.033 | 60.701 | 0.856 | -1.036 | 0 | 0 | 1 | 43.25 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-24774 | 2023/06/20 23:54:06 | 1-11 | 1 | -13.944 | -113.079 | 12181 | -1 | 1 | CPER4-deep-1 | 23-083 | H2O | 1687327151.91 | 0.201 | 0.706 | 64.592 | 0.001 | 0.002 | 1.607 | 41.913 | 28.002 | 2.009 | 60.424 | 0.856 | -0.845 | 0 | 0 | 1 | 43.312 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-24774 | 2023/06/21 00:03:02 | 1-11 | 2 | -14.16 | -115.602 | 12011 | -1 | 1 | CPER4-deep-1 | 23-083 | H2O | 1687327687.61 | 0.202 | 0.787 | 61.796 | 0.001 | 0.004 | 1.534 | 41.652 | 27.552 | 1.969 | 59.491 | 0.855 | -1.103 | 0 | 0 | 1 | 43.25 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-24774 | 2023/06/21 00:11:59 | 1-11 | 3 | -14.237 | -116.622 | 12022 | -1 | 1 | CPER4-deep-1 | 23-083 | H2O | 1687328224.86 | 0.193 | 0.797 | 66.397 | 0.001 | 0.001 | 1.645 | 41.626 | 27.555 | 1.966 | 59.286 | 0.855 | -1.057 | 0 | 0 | 1 | 43.312 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-24774 | 2023/06/21 00:20:56 | 1-11 | 4 | -14.227 | -116.7 | 12094 | 0 | 1 | CPER4-deep-1 | 23-083 | H2O | 1687328761.44 | 0.205 | 0.743 | 65.432 | 0.001 | 0.001 | 1.618 | 41.815 | 27.552 | 1.987 | 60.35 | 0.855 | -0.839 | 0 | 0 | 1 | 43.312 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-24775 | 2023/06/21 00:30:46 | 1-12 | 1 | -9.136 | -95.044 | 12076 | -1 | 1 | CPER4-3040-3 | 23-083 | H2O | 1687329351.87 | 0.184 | 0.779 | 67.645 | 0 | 0.002 | 1.684 | 39.627 | 27.981 | 2.013 | 61.068 | 0.856 | -0.807 | 0 | 0 | 1 | 43.312 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-24775 | 2023/06/21 00:39:42 | 1-12 | 2 | -8.848 | -93.231 | 12069 | -1 | 1 | CPER4-3040-3 | 23-083 | H2O | 1687329887.83 | 0.216 | 0.776 | 65.392 | 0.001 | 0.001 | 1.627 | 39.702 | 27.941 | 2.015 | 60.3 | 0.856 | -1.009 | 0 | 0 | 1 | 43.312 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-24775 | 2023/06/21 00:48:38 | 1-12 | 3 | -8.774 | -92.432 | 12040 | -1 | 1 | CPER4-3040-3 | 23-083 | H2O | 1687330424.23 | 0.225 | 0.877 | 62.466 | 0.001 | 0.003 | 1.555 | 39.734 | 27.768 | 2 | 60.175 | 0.855 | -0.99 | 0 | 0 | 1 | 43.312 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-24775 | 2023/06/21 00:57:35 | 1-12 | 4 | -8.751 | -92.068 | 12143 | 0 | 1 | CPER4-3040-3 | 23-083 | H2O | 1687330960 | 0.208 | 0.793 | 61.599 | 0.001 | 0.002 | 1.531 | 39.891 | 27.8 | 2.031 | 61.58 | 0.855 | -0.899 | 0 | 0 | 1 | 43.375 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-24776 | 2023/06/21 01:07:23 | 1-13 | 1 | -8.891 | -70.504 | 12848 | -1 | 1 | EV | standard | H2O | 1687331548.62 | 0.203 | 0.795 | 68.816 | 0.001 | 0 | 1.704 | 41.176 | 30.522 | 2.169 | 63.491 | 0.856 | -1.211 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-24776 | 2023/06/21 01:16:20 | 1-13 | 2 | -8.953 | -68.371 | 12388 | -1 | 1 | EV | standard | H2O | 1687332085.46 | 0.197 | 0.853 | 63.684 | 0 | 0.003 | 1.575 | 40.303 | 28.9 | 2.106 | 62.207 | 0.856 | -1.003 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-24776 | 2023/06/21 01:25:15 | 1-13 | 3 | -8.95 | -67.711 | 12407 | -1 | 1 | EV | standard | H2O | 1687332621.15 | 0.196 | 0.807 | 62.741 | 0.002 | 0.001 | 1.55 | 40.309 | 28.627 | 2.106 | 62.943 | 0.855 | -0.9 | 0 | 0 | 1 | 43.188 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-24776 | 2023/06/21 01:34:11 | 1-13 | 4 | -8.943 | -67.376 | 12404 | 0 | 1 | EV | standard | H2O | 1687333157.37 | 0.178 | 0.759 | 67.346 | 0.001 | 0.001 | 1.668 | 40.373 | 28.833 | 2.109 | 62.248 | 0.855 | -1.044 | 0 | 0 | 1 | 43.25 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-24777 | 2023/06/21 01:44:01 | 1-14 | 1 | -7.531 | -80.487 | 12085 | -1 | 1 | CPER4-1020-1 | 23-083 | H2O | 1687333747.09 | 0.2 | 0.864 | 66.198 | 0 | 0.002 | 1.649 | 39.587 | 27.943 | 2.046 | 61.292 | 0.855 | -0.875 | 0 | 0 | 1 | 43.375 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-24777 | 2023/06/21 01:52:58 | 1-14 | 2 | -7.415 | -81.6 | 12214 | -1 | 1 | CPER4-1020-1 | 23-083 | H2O | 1687334284.17 | 0.215 | 0.759 | 66.645 | 0.001 | 0.001 | 1.65 | 39.904 | 28.254 | 2.061 | 61.583 | 0.855 | -0.751 | 0 | 0 | 1 | 43.375 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-24777 | 2023/06/21 02:01:54 | 1-14 | 3 | -7.419 | -82.036 | 11953 | -1 | 1 | CPER4-1020-1 | 23-083 | H2O | 1687334820.4 | 0.189 | 0.713 | 63.403 | 0 | 0.004 | 1.579 | 39.37 | 27.397 | 2.024 | 61.231 | 0.855 | -0.837 | 0 | 0 | 1 | 43.438 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-24777 | 2023/06/21 02:10:51 | 1-14 | 4 | -7.388 | -81.949 | 11928 | 0 | 1 | CPER4-1020-1 | 23-083 | H2O | 1687335356.89 | 0.206 | 0.76 | 64.977 | 0.001 | 0.004 | 1.619 | 39.305 | 27.441 | 2.018 | 61.071 | 0.855 | -0.922 | 0 | 0 | 1 | 43.438 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-24778 | 2023/06/21 02:20:38 | 1-15 | 1 | -7.272 | -85.093 | 12314 | -1 | 1 | CPER4-3040-1 | 23-083 | H2O | 1687335943.38 | 0.212 | 0.727 | 63.107 | 0.001 | -0.001 | 1.562 | 40.035 | 28.78 | 2.07 | 61.145 | 0.855 | -1.008 | 0 | 0 | 1 | 43.375 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-24778 | 2023/06/21 02:29:33 | 1-15 | 2 | -7.254 | -85.123 | 12116 | -1 | 1 | CPER4-3040-1 | 23-083 | H2O | 1687336478.89 | 0.217 | 0.881 | 66.664 | 0.001 | 0.005 | 1.659 | 39.768 | 28.117 | 2.026 | 60.497 | 0.855 | -1.029 | 0 | 0 | 1 | 43.375 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-24778 | 2023/06/21 02:38:29 | 1-15 | 3 | -7.228 | -85.412 | 11989 | -1 | 1 | CPER4-3040-1 | 23-083 | H2O | 1687337015.03 | 0.206 | 0.841 | 67.209 | 0.001 | 0.002 | 1.674 | 39.511 | 27.785 | 2.01 | 59.929 | 0.855 | -1.014 | 0 | 0 | 1 | 43.438 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-24778 | 2023/06/21 02:47:26 | 1-15 | 4 | -7.218 | -85.489 | 12038 | 0 | 1 | CPER4-3040-1 | 23-083 | H2O | 1687337551.86 | 0.217 | 0.859 | 65.188 | 0.001 | 0.004 | 1.615 | 39.543 | 27.792 | 2.038 | 60.673 | 0.855 | -1.005 | 0 | 0 | 1 | 43.5 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-24779 | 2023/06/21 02:57:13 | 1-16 | 1 | -8.65 | -87.993 | 12154 | -1 | 1 | CPER4-deep-3 | 23-083 | H2O | 1687338139.26 | 0.189 | 0.742 | 63.862 | 0.001 | 0 | 1.59 | 39.61 | 28.377 | 2.039 | 61.415 | 0.856 | -0.898 | 0 | 0 | 1 | 43.562 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-24779 | 2023/06/21 03:06:10 | 1-16 | 2 | -8.707 | -88.197 | 12180 | -1 | 1 | CPER4-deep-3 | 23-083 | H2O | 1687338674.98 | 0.182 | 0.749 | 62.932 | 0.001 | 0.002 | 1.568 | 39.731 | 27.986 | 2.049 | 62.134 | 0.855 | -0.847 | 0 | 0 | 1 | 43.5 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-24779 | 2023/06/21 03:15:05 | 1-16 | 3 | -8.701 | -88.335 | 12041 | -1 | 1 | CPER4-deep-3 | 23-083 | H2O | 1687339211.17 | 0.192 | 0.836 | 60.566 | 0 | -0.001 | 1.492 | 39.424 | 28.003 | 2.013 | 60.491 | 0.856 | -0.934 | 0 | 0 | 1 | 43.5 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-24779 | 2023/06/21 03:24:01 | 1-16 | 4 | -8.717 | -88.511 | 12053 | 0 | 1 | CPER4-deep-3 | 23-083 | H2O | 1687339747.36 | 0.201 | 0.731 | 65.241 | 0.001 | 0.003 | 1.624 | 39.601 | 27.776 | 2.009 | 60.087 | 0.855 | -0.976 | 0 | 0 | 1 | 43.5 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-24780 | 2023/06/21 03:33:48 | 1-17 | 1 | -4.587 | -73.885 | 12320 | -1 | 1 | CPER4-010-2 | 23-083 | H2O | 1687340333.84 | 0.198 | 0.768 | 64.73 | 0.002 | 0.002 | 1.602 | 40.111 | 28.568 | 2.092 | 62.205 | 0.856 | -1.007 | 0 | 0 | 1 | 43.562 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-24780 | 2023/06/21 03:42:46 | 1-17 | 2 | -4.392 | -72.346 | 12666 | -1 | 1 | CPER4-010-2 | 23-083 | H2O | 1687340871.03 | 0.191 | 0.795 | 65.362 | 0.001 | -0.001 | 1.608 | 40.865 | 29.719 | 2.131 | 63.026 | 0.855 | -1.125 | 0 | 0 | 1 | 43.562 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-24780 | 2023/06/21 03:51:41 | 1-17 | 3 | -4.314 | -71.983 | 12320 | -1 | 1 | CPER4-010-2 | 23-083 | H2O | 1687341407.26 | 0.185 | 0.775 | 65.863 | 0.001 | -0.001 | 1.631 | 39.996 | 28.781 | 2.091 | 62.215 | 0.856 | -0.895 | 0 | 0 | 1 | 43.25 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-24780 | 2023/06/21 04:00:37 | 1-17 | 4 | -4.303 | -71.816 | 12466 | 0 | 1 | CPER4-010-2 | 23-083 | H2O | 1687341942.84 | 0.194 | 0.704 | 65.296 | 0.001 | 0 | 1.626 | 40.596 | 28.886 | 2.108 | 62.09 | 0.856 | -1.144 | 0 | 0 | 1 | 43.312 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-24781 | 2023/06/21 04:10:25 | 1-18 | 1 | -4.853 | -71.096 | 15850 | -1 | 1 | CPER4-010-3 | 23-083 | H2O | 1687342529.74 | 0.212 | 0.675 | 78.488 | 0.001 | 0.001 | 1.903 | 47.434 | 39.173 | 2.634 | 74.776 | 0.857 | -1.131 | 0 | 0 | 1 | 43.438 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-24781 | 2023/06/21 04:19:21 | 1-18 | 2 | -4.848 | -70.741 | 12267 | -1 | 1 | CPER4-010-3 | 23-083 | H2O | 1687343066.97 | 0.201 | 0.784 | 66.662 | 0 | 0.001 | 1.649 | 40.045 | 28.368 | 2.075 | 62.227 | 0.855 | -1.05 | 0 | 0 | 1 | 43.5 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-24781 | 2023/06/21 04:28:17 | 1-18 | 3 | -4.862 | -70.794 | 11171 | -1 | 1 | CPER4-010-3 | 23-083 | H2O | 1687343603.2 | 0.213 | 0.817 | 61.772 | 0.001 | 0.001 | 1.557 | 37.769 | 25.257 | 1.924 | 58.012 | 0.855 | -0.8 | 0 | 0 | 1 | 43.5 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-24781 | 2023/06/21 04:37:14 | 1-18 | 4 | -4.862 | -70.831 | 10893 | 0 | 1 | CPER4-010-3 | 23-083 | H2O | 1687344140.39 | 0.207 | 0.836 | 60.771 | 0 | 0.001 | 1.534 | 37.277 | 24.53 | 1.867 | 56.568 | 0.855 | -0.772 | 0 | 0 | 1 | 43.438 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-24782 | 2023/06/21 04:47:02 | 1-19 | 1 | -7.902 | -80.099 | 12460 | -1 | 1 | CPER4-B-3 | 23-083 | H2O | 1687344727.03 | 0.202 | 0.728 | 65.962 | 0.001 | 0.002 | 1.63 | 41.972 | 28.902 | 2.102 | 62.121 | 0.856 | -0.977 | 0 | 0 | 1 | 43.375 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-24782 | 2023/06/21 04:55:58 | 1-19 | 2 | -8.073 | -80.776 | 12164 | -1 | 1 | CPER4-B-3 | 23-083 | H2O | 1687345263.34 | 0.197 | 0.796 | 61.64 | 0.001 | 0.002 | 1.532 | 41.369 | 27.96 | 2.058 | 61.664 | 0.856 | -0.94 | 0 | 0 | 1 | 43.438 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-24782 | 2023/06/21 05:04:53 | 1-19 | 3 | -8.11 | -81.158 | 12130 | -1 | 1 | CPER4-B-3 | 23-083 | H2O | 1687345799.26 | 0.197 | 0.818 | 64.444 | 0.001 | -0.002 | 1.602 | 41.32 | 27.89 | 2.046 | 61.453 | 0.856 | -1.032 | 0 | 0 | 1 | 43.5 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-24782 | 2023/06/21 05:13:50 | 1-19 | 4 | -8.104 | -81.171 | 12206 | 0 | 1 | CPER4-B-3 | 23-083 | H2O | 1687346335.39 | 0.189 | 0.785 | 63.623 | 0.001 | -0.002 | 1.575 | 41.36 | 28.044 | 2.075 | 62.076 | 0.855 | -0.846 | 0 | 0 | 1 | 43.562 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-24783 | 2023/06/21 05:23:38 | 1-20 | 1 | -7.756 | -81.831 | 11287 | -1 | 1 | CPER4-1020-3 | 23-083 | H2O | 1687346925.27 | 0.185 | 0.82 | 62.485 | 0.002 | 0 | 1.576 | 38.032 | 25.419 | 1.926 | 58.8 | 0.855 | -0.7 | 0 | 0 | 1 | 43.562 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-24783 | 2023/06/21 05:32:35 | 1-20 | 2 | -7.747 | -81.993 | 11220 | -1 | 1 | CPER4-1020-3 | 23-083 | H2O | 1687347461.62 | 0.195 | 0.864 | 62.579 | 0.001 | 0.004 | 1.565 | 37.99 | 25.343 | 1.908 | 57.104 | 0.855 | -0.842 | 0 | 0 | 1 | 43.5 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-24783 | 2023/06/21 05:41:31 | 1-20 | 3 | -7.722 | -82.092 | 11291 | -1 | 1 | CPER4-1020-3 | 23-083 | H2O | 1687347997.78 | 0.197 | 0.83 | 65.24 | 0.002 | 0.003 | 1.644 | 38.085 | 25.881 | 1.918 | 57.05 | 0.855 | -0.999 | 0 | 0 | 1 | 43.375 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-24783 | 2023/06/21 05:50:29 | 1-20 | 4 | -7.721 | -82.117 | 11464 | 0 | 1 | CPER4-1020-3 | 23-083 | H2O | 1687348534.6 | 0.203 | 0.791 | 60.088 | 0.001 | -0.002 | 1.494 | 38.535 | 25.71 | 1.941 | 59.512 | 0.856 | -0.896 | 0 | 0 | 1 | 43.5 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-24784 | 2023/06/21 06:00:18 | 1-21 | 1 | -9.176 | -90.646 | 12580 | -1 | 1 | CPER4-deep-2 | 23-083 | H2O | 1687349122.97 | 0.215 | 0.806 | 64.295 | 0.001 | 0 | 1.589 | 40.596 | 29.209 | 2.124 | 63.049 | 0.855 | -0.971 | 0 | 0 | 1 | 43.5 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-24784 | 2023/06/21 06:09:15 | 1-21 | 2 | -9.192 | -91.328 | 12440 | -1 | 1 | CPER4-deep-2 | 23-083 | H2O | 1687349659.96 | 0.18 | 0.853 | 69.462 | 0.001 | 0.004 | 1.72 | 40.363 | 28.619 | 2.087 | 62.738 | 0.855 | -0.941 | 0 | 0 | 1 | 43.5 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-24784 | 2023/06/21 06:18:11 | 1-21 | 3 | -9.233 | -91.624 | 12071 | -1 | 1 | CPER4-deep-2 | 23-083 | H2O | 1687350196.27 | 0.184 | 0.815 | 65.631 | 0.001 | 0.003 | 1.634 | 39.663 | 27.714 | 2.018 | 60.686 | 0.855 | -1.009 | 0 | 0 | 1 | 43.438 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-24784 | 2023/06/21 06:27:07 | 1-21 | 4 | -9.24 | -91.462 | 12196 | 0 | 1 | CPER4-deep-2 | 23-083 | H2O | 1687350732.12 | 0.18 | 0.719 | 62.724 | 0 | 0 | 1.552 | 39.9 | 28.066 | 2.038 | 61.532 | 0.855 | -0.957 | 0 | 0 | 1 | 43.562 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-24785 | 2023/06/21 06:36:52 | 1-22 | 1 | -8.953 | -70.45 | 12718 | -1 | 1 | EV | standard | H2O | 1687351317.69 | 0.205 | 0.852 | 64.28 | 0.001 | 0.002 | 1.581 | 40.838 | 29.605 | 2.16 | 63.818 | 0.856 | -0.955 | 0 | 0 | 1 | 43.562 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-24785 | 2023/06/21 06:45:49 | 1-22 | 2 | -8.903 | -67.977 | 12372 | -1 | 1 | EV | standard | H2O | 1687351854.5 | 0.205 | 0.789 | 64.243 | 0.001 | -0.001 | 1.589 | 40.179 | 28.682 | 2.1 | 62.868 | 0.856 | -0.957 | 0 | 0 | 1 | 43.625 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-24785 | 2023/06/21 06:54:44 | 1-22 | 3 | -8.952 | -67.903 | 12323 | -1 | 1 | EV | standard | H2O | 1687352389.39 | 0.217 | 0.827 | 63.675 | 0 | 0.002 | 1.574 | 40.055 | 28.831 | 2.103 | 61.754 | 0.856 | -1.065 | 0 | 0 | 1 | 43.562 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-24785 | 2023/06/21 07:03:40 | 1-22 | 4 | -8.944 | -67.372 | 12290 | 0 | 1 | EV | standard | H2O | 1687352925.35 | 0.184 | 0.797 | 65.607 | 0.001 | 0.003 | 1.621 | 39.911 | 28.524 | 2.089 | 62.398 | 0.856 | -0.9 | 0 | 0 | 1 | 43.625 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-24786 | 2023/06/21 07:13:26 | 1-23 | 1 | -8.408 | -83.204 | 11914 | -1 | 1 | CPER4-1020-2 | 23-083 | H2O | 1687353511.27 | 0.205 | 0.793 | 61.361 | 0.001 | -0.001 | 1.524 | 39.193 | 27.604 | 2.018 | 60.377 | 0.856 | -0.812 | 0 | 0 | 1 | 43.375 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-24786 | 2023/06/21 07:22:22 | 1-23 | 2 | -8.398 | -84.754 | 11704 | -1 | 1 | CPER4-1020-2 | 23-083 | H2O | 1687354048.28 | 0.197 | 0.784 | 64.591 | 0 | 0.004 | 1.608 | 38.839 | 27.118 | 1.981 | 58.493 | 0.855 | -1.065 | 0 | 0 | 1 | 43.125 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-24786 | 2023/06/21 07:31:19 | 1-23 | 3 | -8.36 | -85.044 | 11575 | -1 | 1 | CPER4-1020-2 | 23-083 | H2O | 1687354585.22 | 0.197 | 0.81 | 60.521 | 0.001 | -0.001 | 1.509 | 38.709 | 26.692 | 1.948 | 57.43 | 0.855 | -1.097 | 0 | 0 | 1 | 43.25 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-24786 | 2023/06/21 07:40:16 | 1-23 | 4 | -8.375 | -85.046 | 11306 | 0 | 1 | CPER4-1020-2 | 23-083 | H2O | 1687355121.85 | 0.204 | 0.859 | 62.036 | 0.001 | 0.004 | 1.563 | 38.075 | 25.714 | 1.921 | 57.752 | 0.856 | -0.922 | 0 | 0 | 1 | 43.312 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-24787 | 2023/06/21 07:50:02 | 1-24 | 1 | -4.125 | -73.704 | 11614 | -1 | 1 | CPER4-B-5 | 23-083 | H2O | 1687355707.79 | 0.2 | 0.831 | 65.767 | 0.001 | 0.003 | 1.647 | 38.892 | 26.361 | 2.002 | 59.514 | 0.856 | -0.979 | 0 | 0 | 1 | 43.312 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-24787 | 2023/06/21 07:58:59 | 1-24 | 2 | -3.892 | -72.473 | 11662 | -1 | 1 | CPER4-B-5 | 23-083 | H2O | 1687356244.95 | 0.207 | 0.845 | 61.279 | 0 | -0.001 | 1.527 | 38.983 | 26.73 | 1.994 | 59.436 | 0.855 | -1.07 | 0 | 0 | 1 | 43.312 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-24787 | 2023/06/21 08:07:55 | 1-24 | 3 | -3.864 | -72.04 | 11624 | -1 | 1 | CPER4-B-5 | 23-083 | H2O | 1687356780.74 | 0.194 | 0.793 | 62.756 | 0.001 | -0.002 | 1.561 | 38.896 | 26.727 | 1.988 | 59.202 | 0.855 | -1.022 | 0 | 0 | 1 | 43.25 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-24787 | 2023/06/21 08:16:51 | 1-24 | 4 | -3.808 | -71.905 | 11697 | 0 | 1 | CPER4-B-5 | 23-083 | H2O | 1687357317.29 | 0.201 | 0.779 | 64.601 | 0.001 | -0.002 | 1.606 | 39.099 | 26.723 | 2.001 | 59.856 | 0.856 | -1.042 | 0 | 0 | 1 | 43.312 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-24788 | 2023/06/21 08:26:39 | 1-25 | 1 | -5.325 | -77.475 | 12108 | -1 | 1 | CPER4-C-5 | 23-083 | H2O | 1687357905.05 | 0.193 | 0.756 | 61.262 | 0 | -0.004 | 1.525 | 39.923 | 28.09 | 2.05 | 60.151 | 0.855 | -1.162 | 0 | 0 | 1 | 43.188 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-24788 | 2023/06/21 08:35:35 | 1-25 | 2 | -5.404 | -78.193 | 12129 | -1 | 1 | CPER4-C-5 | 23-083 | H2O | 1687358440.47 | 0.207 | 0.802 | 63.573 | 0.001 | 0.004 | 1.572 | 39.937 | 27.713 | 2.06 | 61.913 | 0.856 | -0.969 | 0 | 0 | 1 | 43.438 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-24788 | 2023/06/21 08:44:31 | 1-25 | 3 | -5.447 | -78.415 | 12018 | -1 | 1 | CPER4-C-5 | 23-083 | H2O | 1687358976.32 | 0.184 | 0.759 | 63.512 | 0.001 | 0.002 | 1.592 | 39.68 | 27.357 | 2.052 | 61.617 | 0.855 | -0.823 | 0 | 0 | 1 | 43.5 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-24788 | 2023/06/21 08:53:26 | 1-25 | 4 | -5.429 | -78.368 | 12004 | 0 | 1 | CPER4-C-5 | 23-083 | H2O | 1687359511.65 | 0.205 | 0.747 | 66.286 | 0.001 | -0.003 | 1.662 | 39.606 | 27.386 | 2.046 | 61.779 | 0.855 | -0.888 | 0 | 0 | 1 | 43.562 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-24789 | 2023/06/21 09:03:14 | 1-26 | 1 | 4.778 | -53.362 | 12575 | -1 | 1 | CPER4-D-2 | 23-083 | H2O | 1687360099.46 | 0.216 | 0.883 | 59.573 | -0.001 | -0.001 | 1.461 | 148.623 | 16.958 | 3.444 | 65.641 | 0.856 | -5.706 | 0 | 0 | 1 | 43.375 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-24789 | 2023/06/21 09:12:09 | 1-26 | 2 | 4.86 | -52.12 | 12189 | -1 | 1 | CPER4-D-2 | 23-083 | H2O | 1687360635.1 | 0.215 | 1.056 | 62.274 | 0 | 0 | 1.546 | 145.307 | 16.418 | 3.349 | 64.365 | 0.857 | -5.546 | 0 | 0 | 1 | 43.062 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-24789 | 2023/06/21 09:21:05 | 1-26 | 3 | 4.928 | -51.442 | 12225 | -1 | 1 | CPER4-D-2 | 23-083 | H2O | 1687361170.69 | 0.191 | 0.935 | 61.181 | 0 | 0 | 1.522 | 146.16 | 16.241 | 3.374 | 63.801 | 0.856 | -5.564 | 0 | 0 | 1 | 42.875 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-24789 | 2023/06/21 09:30:02 | 1-26 | 4 | 4.959 | -51.304 | 12222 | 0 | 1 | CPER4-D-2 | 23-083 | H2O | 1687361707.92 | 0.2 | 1.031 | 63.033 | 0 | -0.001 | 1.568 | 146.214 | 16.247 | 3.368 | 64.189 | 0.856 | -5.511 | 0 | 0 | 1 | 43 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-24790 | 2023/06/21 09:39:51 | 1-27 | 1 | -5.81 | -77.628 | 12728 | -1 | 1 | CPER4-C-2 | 23-083 | H2O | 1687362296.08 | 0.207 | 0.761 | 67.374 | 0.002 | 0.001 | 1.678 | 43.071 | 29.825 | 2.168 | 63.652 | 0.857 | -0.955 | 0 | 0 | 1 | 43.062 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-24790 | 2023/06/21 09:48:47 | 1-27 | 2 | -6 | -79.07 | 12358 | -1 | 1 | CPER4-C-2 | 23-083 | H2O | 1687362832.1 | 0.225 | 0.829 | 64.76 | 0.001 | -0.003 | 1.602 | 41.499 | 28.685 | 2.116 | 61.646 | 0.856 | -1.03 | 0 | 0 | 1 | 43.125 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-24790 | 2023/06/21 09:57:42 | 1-27 | 3 | -6.039 | -79.317 | 12365 | -1 | 1 | CPER4-C-2 | 23-083 | H2O | 1687363367.92 | 0.211 | 0.762 | 63.291 | 0.002 | 0.002 | 1.567 | 41.285 | 28.505 | 2.092 | 62.72 | 0.856 | -0.928 | 0 | 0 | 1 | 43 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-24790 | 2023/06/21 10:06:38 | 1-27 | 4 | -6.081 | -79.546 | 12393 | 0 | 1 | CPER4-C-2 | 23-083 | H2O | 1687363903.57 | 0.197 | 0.789 | 63.681 | 0.001 | 0.004 | 1.579 | 41.088 | 28.76 | 2.107 | 62.635 | 0.856 | -0.845 | 0 | 0 | 1 | 42.938 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-24791 | 2023/06/21 10:16:26 | 1-28 | 1 | -3.869 | -72.2 | 11683 | -1 | 1 | CPER4-C-4 | 23-083 | H2O | 1687364492.14 | 0.227 | 0.799 | 61.52 | 0 | 0.001 | 1.54 | 162.089 | 22.178 | 2.079 | 57.45 | 0.856 | -1.344 | 0 | 0 | 1 | 43.125 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-24791 | 2023/06/21 10:25:22 | 1-28 | 2 | -3.779 | -71.692 | 11681 | -1 | 1 | CPER4-C-4 | 23-083 | H2O | 1687365028.1 | 0.216 | 0.856 | 56.868 | 0 | 0.003 | 1.42 | 163.534 | 22.338 | 2.066 | 57.052 | 0.856 | -1.336 | 0 | 0 | 1 | 43.188 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-24791 | 2023/06/21 10:34:18 | 1-28 | 3 | -3.753 | -71.543 | 11967 | -1 | 1 | CPER4-C-4 | 23-083 | H2O | 1687365563.71 | 0.203 | 0.832 | 63.438 | 0.002 | -0.002 | 1.577 | 167.378 | 23.007 | 2.119 | 58.626 | 0.856 | -1.424 | 0 | 0 | 1 | 43.312 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-24791 | 2023/06/21 10:43:14 | 1-28 | 4 | -3.761 | -71.447 | 11933 | 0 | 1 | CPER4-C-4 | 23-083 | H2O | 1687366099.95 | 0.218 | 0.902 | 64.154 | 0 | 0.001 | 1.594 | 167.049 | 22.9 | 2.097 | 57.811 | 0.856 | -1.521 | 0 | 0 | 1 | 43.375 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-24792 | 2023/06/21 10:52:58 | 1-29 | 1 | 2.457 | -55.103 | 12017 | -1 | 1 | CPER4-D-1 | 23-083 | H2O | 1687366683.75 | 0.201 | 0.875 | 64.413 | 0 | -0.002 | 1.603 | 125.694 | 18.465 | 2.883 | 63.372 | 0.856 | -4.166 | 0 | 0 | 1 | 43.438 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-24792 | 2023/06/21 11:01:55 | 1-29 | 2 | 2.527 | -53.845 | 11923 | -1 | 1 | CPER4-D-1 | 23-083 | H2O | 1687367221.03 | 0.225 | 0.924 | 63.183 | 0 | -0.001 | 1.567 | 124.367 | 18.672 | 2.874 | 62.392 | 0.857 | -4.242 | 0 | 0 | 1 | 43.25 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-24792 | 2023/06/21 11:10:52 | 1-29 | 3 | 2.554 | -53.621 | 12134 | -1 | 1 | CPER4-D-1 | 23-083 | H2O | 1687367757.46 | 0.207 | 0.978 | 63.603 | 0 | 0.003 | 1.566 | 126.409 | 18.687 | 2.939 | 64.218 | 0.856 | -4.11 | 0 | 0 | 1 | 43.062 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-24792 | 2023/06/21 11:19:48 | 1-29 | 4 | 2.51 | -53.804 | 12080 | 0 | 1 | CPER4-D-1 | 23-083 | H2O | 1687368293.64 | 0.191 | 0.856 | 62.18 | 0 | -0.004 | 1.547 | 125.841 | 18.657 | 2.914 | 64.298 | 0.856 | -4.223 | 0 | 0 | 1 | 43.062 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-24793 | 2023/06/21 11:29:33 | 1-30 | 1 | -2.831 | -58.165 | 12020 | -1 | 1 | CPER4-A-2 | 23-083 | H2O | 1687368879.43 | 0.223 | 0.882 | 64.403 | 0.001 | 0.001 | 1.599 | 414.532 | 15.301 | 2.179 | 54.623 | 0.857 | -1.568 | 0 | 0 | 1 | 43.062 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-24793 | 2023/06/21 11:38:29 | 1-30 | 2 | -2.839 | -57.824 | 11778 | -1 | 1 | CPER4-A-2 | 23-083 | H2O | 1687369415.2 | 0.239 | 1.104 | 60.535 | 0.001 | -0.001 | 1.51 | 408.554 | 14.692 | 2.126 | 53.52 | 0.857 | -1.571 | 0 | 0 | 1 | 43.125 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-24793 | 2023/06/21 11:47:25 | 1-30 | 3 | -2.89 | -57.835 | 11759 | -1 | 1 | CPER4-A-2 | 23-083 | H2O | 1687369950.69 | 0.221 | 0.962 | 65.327 | 0.002 | 0.002 | 1.634 | 408.218 | 14.412 | 2.118 | 54.652 | 0.857 | -1.385 | 0 | 0 | 1 | 43.25 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-24793 | 2023/06/21 11:56:21 | 1-30 | 4 | -2.837 | -57.637 | 11763 | 0 | 1 | CPER4-A-2 | 23-083 | H2O | 1687370487.33 | 0.242 | 1.019 | 60.99 | 0.001 | 0 | 1.525 | 408.852 | 14.713 | 2.108 | 53.247 | 0.857 | -1.643 | 0 | 0 | 1 | 43.25 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-24794 | 2023/06/21 12:06:07 | 1-31 | 1 | -8.512 | -65.481 | 12638 | -1 | 1 | EV | standard | H2O | 1687371072.27 | 0.196 | 0.785 | 64.077 | 0.002 | -0.001 | 1.574 | 45.616 | 29.612 | 2.155 | 63.2 | 0.856 | -0.983 | 0 | 0 | 1 | 43.375 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-24794 | 2023/06/21 12:15:02 | 1-31 | 2 | -8.794 | -66.368 | 12284 | -1 | 1 | EV | standard | H2O | 1687371607.77 | 0.197 | 0.771 | 60.781 | 0.001 | 0.002 | 1.5 | 42.55 | 28.41 | 2.086 | 60.842 | 0.856 | -1.107 | 0 | 0 | 1 | 43.25 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-24794 | 2023/06/21 12:23:58 | 1-31 | 3 | -8.859 | -66.788 | 12233 | -1 | 1 | EV | standard | H2O | 1687372143.85 | 0.206 | 0.826 | 67.929 | 0.001 | 0 | 1.68 | 41.346 | 27.875 | 2.094 | 62.895 | 0.856 | -0.869 | 0 | 0 | 1 | 43.188 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-24794 | 2023/06/21 12:32:55 | 1-31 | 4 | -8.893 | -66.901 | 12269 | 0 | 1 | EV | standard | H2O | 1687372680.2 | 0.185 | 0.716 | 66.144 | 0.001 | 0.002 | 1.636 | 41.048 | 28.506 | 2.097 | 62.641 | 0.856 | -0.871 | 0 | 0 | 1 | 43.312 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-24795 | 2023/06/21 12:42:41 | 1-32 | 1 | -4.087 | -68.511 | 12271 | -1 | 1 | CPER4-010-1 | 23-083 | H2O | 1687373266.62 | 0.232 | 0.747 | 69.45 | 0 | 0 | 1.715 | 40.629 | 28.494 | 2.106 | 61.839 | 0.856 | -0.941 | 0 | 0 | 1 | 43.375 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-24795 | 2023/06/21 12:51:36 | 1-32 | 2 | -3.887 | -68.633 | 12192 | -1 | 1 | CPER4-010-1 | 23-083 | H2O | 1687373802.15 | 0.189 | 0.783 | 57.464 | 0.001 | 0 | 1.428 | 40.68 | 28.251 | 2.088 | 61.549 | 0.857 | -0.975 | 0 | 0 | 1 | 43.438 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-24795 | 2023/06/21 13:00:33 | 1-32 | 3 | -3.82 | -68.699 | 12326 | -1 | 1 | CPER4-010-1 | 23-083 | H2O | 1687374338.72 | 0.197 | 0.741 | 61.027 | 0 | 0.003 | 1.518 | 40.958 | 28.401 | 2.102 | 62.28 | 0.855 | -0.935 | 0 | 0 | 1 | 43.438 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-24795 | 2023/06/21 13:09:30 | 1-32 | 4 | -3.78 | -68.905 | 12274 | 0 | 1 | CPER4-010-1 | 23-083 | H2O | 1687374875.19 | 0.204 | 0.864 | 66.267 | 0 | 0.001 | 1.635 | 40.707 | 28.316 | 2.089 | 61.941 | 0.856 | -0.952 | 0 | 0 | 1 | 43.5 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-24796 | 2023/06/21 13:19:16 | 1-33 | 1 | 9.367 | -51.641 | 11699 | -1 | 1 | CPER4-D-4 | 23-083 | H2O | 1687375461.41 | 0.212 | 1.103 | 65.146 | -0.001 | 0.002 | 1.629 | 173.809 | 9.282 | 4.046 | 65.555 | 0.856 | -7.426 | 0 | 0 | 1 | 43.5 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-24796 | 2023/06/21 13:28:12 | 1-33 | 2 | 9.475 | -51.452 | 11845 | -1 | 1 | CPER4-D-4 | 23-083 | H2O | 1687375997.16 | 0.213 | 1.126 | 65.565 | 0 | -0.003 | 1.631 | 177.064 | 9.761 | 4.115 | 64.827 | 0.856 | -7.588 | 0 | 0 | 1 | 43.5 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-24796 | 2023/06/21 13:37:08 | 1-33 | 3 | 9.55 | -51.636 | 11710 | -1 | 1 | CPER4-D-4 | 23-083 | H2O | 1687376533.45 | 0.216 | 1.056 | 64.419 | -0.002 | 0 | 1.599 | 175.424 | 9.586 | 4.073 | 65.054 | 0.856 | -7.474 | 0 | 0 | 1 | 43.5 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-24796 | 2023/06/21 13:46:02 | 1-33 | 4 | 9.577 | -51.502 | 11653 | 0 | 1 | CPER4-D-4 | 23-083 | H2O | 1687377068.14 | 0.206 | 1.064 | 62.704 | -0.001 | 0.001 | 1.57 | 174.916 | 9.206 | 4.064 | 65.396 | 0.856 | -7.405 | 0 | 0 | 1 | 43.375 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-24797 | 2023/06/21 13:55:50 | 1-34 | 1 | -7.725 | -83.014 | 9483 | -1 | 1 | CPER4-2030-1 | 23-083 | H2O | 1687377657.68 | 0.209 | 0.875 | 60.345 | 0.002 | 0.003 | 1.573 | 36.673 | 20.268 | 1.649 | 51.165 | 0.854 | -0.812 | 0 | 0 | 1 | 43.438 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-24797 | 2023/06/21 14:04:45 | 1-34 | 2 | -8.152 | -85.457 | 9877 | -1 | 1 | CPER4-2030-1 | 23-083 | H2O | 1687378193.14 | 0.19 | 0.874 | 58.408 | 0.001 | 0.006 | 1.513 | 36.475 | 21.242 | 1.699 | 52.86 | 0.855 | -0.745 | 0 | 0 | 1 | 43.438 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-24797 | 2023/06/21 14:13:40 | 1-34 | 3 | -8.294 | -86.197 | 11233 | -1 | 1 | CPER4-2030-1 | 23-083 | H2O | 1687378726.49 | 0.189 | 0.902 | 59.362 | 0.001 | 0 | 1.486 | 38.809 | 25.205 | 1.902 | 58.13 | 0.856 | -0.834 | 0 | 0 | 1 | 43.562 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-24797 | 2023/06/21 14:22:36 | 1-34 | 4 | -8.33 | -86.295 | 10948 | 0 | 1 | CPER4-2030-1 | 23-083 | H2O | 1687379261.89 | 0.201 | 0.771 | 59.511 | 0.001 | 0.001 | 1.5 | 38.044 | 24.533 | 1.861 | 56.728 | 0.855 | -0.942 | 0 | 0 | 1 | 43.375 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-24798 | 2023/06/21 14:32:22 | 1-35 | 1 | -2.701 | -65.383 | 12347 | -1 | 1 | CPER4-E-2 | 23-083 | H2O | 1687379847.6 | 0.21 | 0.827 | 66.144 | 0.001 | 0 | 1.637 | 40.583 | 28.619 | 2.108 | 62.577 | 0.855 | -0.976 | 0 | 0 | 1 | 43.375 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-24798 | 2023/06/21 14:41:17 | 1-35 | 2 | -2.4 | -63.543 | 12150 | -1 | 1 | CPER4-E-2 | 23-083 | H2O | 1687380382.64 | 0.2 | 0.843 | 63.604 | 0.001 | 0.001 | 1.584 | 40.271 | 28.189 | 2.081 | 61.686 | 0.856 | -1.003 | 0 | 0 | 1 | 43.312 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-24798 | 2023/06/21 14:50:14 | 1-35 | 3 | -2.389 | -62.965 | 12222 | -1 | 1 | CPER4-E-2 | 23-083 | H2O | 1687380919.32 | 0.202 | 0.758 | 65.802 | 0.001 | 0.001 | 1.624 | 40.428 | 28.382 | 2.102 | 61.631 | 0.856 | -1.034 | 0 | 0 | 1 | 43.188 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-24798 | 2023/06/21 14:59:09 | 1-35 | 4 | -2.295 | -62.531 | 12165 | 0 | 1 | CPER4-E-2 | 23-083 | H2O | 1687381455.13 | 0.198 | 0.798 | 61.2 | 0.001 | -0.003 | 1.511 | 40.306 | 28.3 | 2.083 | 61.149 | 0.856 | -1.053 | 0 | 0 | 1 | 43.25 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-24799 | 2023/06/21 15:08:52 | 1-36 | 1 | -3.326 | -64.841 | 12250 | -1 | 1 | CPER4-A-1 | 23-083 | H2O | 1687382038 | 0.192 | 0.811 | 66.083 | 0.001 | -0.005 | 1.637 | 40.106 | 28.611 | 2.096 | 61.598 | 0.856 | -0.902 | 0 | 0 | 1 | 43.312 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-24799 | 2023/06/21 15:17:49 | 1-36 | 2 | -3.359 | -64.61 | 11964 | -1 | 1 | CPER4-A-1 | 23-083 | H2O | 1687382574.51 | 0.221 | 0.737 | 63.108 | 0.001 | 0.001 | 1.555 | 39.604 | 27.477 | 2.054 | 62.004 | 0.856 | -0.827 | 0 | 0 | 1 | 43.438 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-24799 | 2023/06/21 15:26:43 | 1-36 | 3 | -3.35 | -64.779 | 11992 | -1 | 1 | CPER4-A-1 | 23-083 | H2O | 1687383108.86 | 0.201 | 0.905 | 64.834 | 0.001 | -0.001 | 1.609 | 39.621 | 27.395 | 2.061 | 62.181 | 0.856 | -0.795 | 0 | 0 | 1 | 43.5 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-24799 | 2023/06/21 15:35:39 | 1-36 | 4 | -3.351 | -64.656 | 12123 | 0 | 1 | CPER4-A-1 | 23-083 | H2O | 1687383644.4 | 0.184 | 0.743 | 65.615 | 0.001 | 0 | 1.631 | 40.053 | 27.978 | 2.063 | 61.362 | 0.855 | -1.012 | 0 | 0 | 1 | 43.625 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-24800 | 2023/06/21 15:45:21 | 1-37 | 1 | 0.678 | -19.961 | 12283 | -1 | 1 | CLBJ6-1020-2 | 23-084 | H2O | 1687384227.09 | 0.211 | 0.811 | 66.26 | 0.001 | 0.002 | 1.651 | 43.753 | 28.282 | 2.167 | 63.229 | 0.856 | -1.061 | 0 | 0 | 1 | 43.688 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-24800 | 2023/06/21 15:54:17 | 1-37 | 2 | 0.907 | -15.051 | 12137 | -1 | 1 | CLBJ6-1020-2 | 23-084 | H2O | 1687384762.12 | 0.214 | 0.754 | 67.361 | 0.001 | -0.001 | 1.675 | 43.474 | 27.976 | 2.141 | 63.116 | 0.856 | -0.95 | 0 | 0 | 1 | 43.688 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-24800 | 2023/06/21 16:03:13 | 1-37 | 3 | 0.902 | -13.768 | 12128 | -1 | 1 | CLBJ6-1020-2 | 23-084 | H2O | 1687385299.27 | 0.199 | 0.847 | 62.43 | 0.001 | -0.006 | 1.551 | 43.371 | 27.601 | 2.154 | 64.231 | 0.855 | -0.818 | 0 | 0 | 1 | 43.688 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-24800 | 2023/06/21 16:12:08 | 1-37 | 4 | 0.945 | -13.058 | 12179 | 0 | 1 | CLBJ6-1020-2 | 23-084 | H2O | 1687385834.41 | 0.19 | 0.738 | 64.947 | 0.001 | 0 | 1.607 | 43.652 | 28.217 | 2.147 | 62.343 | 0.856 | -1.203 | 0 | 0 | 1 | 43.625 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-24801 | 2023/06/21 16:21:52 | 1-38 | 1 | 3.546 | -6.701 | 12359 | -1 | 1 | CLBJ6-010-3 | 23-084 | H2O | 1687386418.11 | 0.208 | 0.834 | 67.768 | 0.001 | 0 | 1.672 | 40.105 | 28.811 | 2.213 | 65.137 | 0.856 | -0.846 | 0 | 0 | 1 | 43.438 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-24801 | 2023/06/21 16:30:48 | 1-38 | 2 | 3.647 | -5.869 | 12261 | -1 | 1 | CLBJ6-010-3 | 23-084 | H2O | 1687386953.23 | 0.205 | 0.941 | 65.356 | 0.001 | -0.001 | 1.617 | 39.937 | 28.519 | 2.194 | 64.695 | 0.856 | -0.81 | 0 | 0 | 1 | 43.375 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-24801 | 2023/06/21 16:39:42 | 1-38 | 3 | 3.69 | -5.487 | 12178 | -1 | 1 | CLBJ6-010-3 | 23-084 | H2O | 1687387488.53 | 0.201 | 0.837 | 64.821 | 0.001 | -0.001 | 1.611 | 39.815 | 28.251 | 2.175 | 63.551 | 0.856 | -1.049 | 0 | 0 | 1 | 43.188 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-24801 | 2023/06/21 16:48:39 | 1-38 | 4 | 3.697 | -5.358 | 12147 | 0 | 1 | CLBJ6-010-3 | 23-084 | H2O | 1687388024.33 | 0.196 | 0.868 | 65.365 | 0.001 | -0.002 | 1.627 | 39.789 | 28.318 | 2.17 | 63.693 | 0.856 | -1.127 | 0 | 0 | 1 | 43.188 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-24802 | 2023/06/21 16:58:24 | 1-39 | 1 | -8.028 | -58.202 | 12568 | -1 | 1 | EV | standard | H2O | 1687388609.43 | 0.21 | 0.729 | 65.835 | 0.002 | -0.002 | 1.626 | 40.542 | 29.592 | 2.146 | 63.091 | 0.856 | -0.988 | 0 | 0 | 1 | 43.25 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-24802 | 2023/06/21 17:07:20 | 1-39 | 2 | -8.603 | -63.513 | 12428 | -1 | 1 | EV | standard | H2O | 1687389145.04 | 0.199 | 0.781 | 64.65 | 0.001 | 0.003 | 1.606 | 40.295 | 28.934 | 2.126 | 62.853 | 0.857 | -0.948 | 0 | 0 | 1 | 43.25 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-24802 | 2023/06/21 17:16:15 | 1-39 | 3 | -8.746 | -65.011 | 12336 | -1 | 1 | EV | standard | H2O | 1687389681.23 | 0.221 | 0.78 | 64.756 | 0.001 | -0.001 | 1.59 | 40.159 | 28.874 | 2.095 | 61.671 | 0.856 | -1.036 | 0 | 0 | 1 | 43.188 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-24802 | 2023/06/21 17:25:11 | 1-39 | 4 | -8.73 | -65.145 | 12393 | 0 | 1 | EV | standard | H2O | 1687390217.02 | 0.196 | 0.704 | 62.141 | 0.001 | -0.001 | 1.538 | 40.194 | 29.047 | 2.109 | 62.292 | 0.856 | -0.909 | 0 | 0 | 1 | 43.188 | 1 | 39 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 42.1053
varray: [41.163, 44.215, 44.289, 44.456, 44.189, 43.91, 43.548, 43.419, 43.109, 42.758, 41.241, 42.343, 42.397, 42.311, 42.026, 42.122, 41.566, 41.476, 41.338, 30.23]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
baseline_shift_STD | 3.00620687911
varray: [41.163, 44.215, 44.289, 44.456, 44.189, 43.91, 43.548, 43.419, 43.109, 42.758, 41.241, 42.343, 42.397, 42.311, 42.026, 42.122, 41.566, 41.476, 41.338, 30.23]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
ch4_ppm_MEAN | -0.99355
varray: [-1.106, -1.093, -1.26, -1.101, -0.975, -1.026, -1.023, -1.21, -1.196, -1.059, -0.715, -0.901, -1.023, -0.991, -0.9, -0.985, -0.801, -0.973, -0.922, -0.611]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.159650432277
varray: [-1.106, -1.093, -1.26, -1.101, -0.975, -1.026, -1.023, -1.21, -1.196, -1.059, -0.715, -0.901, -1.023, -0.991, -0.9, -0.985, -0.801, -0.973, -0.922, -0.611]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
residuals_MEAN | 2.1614
varray: [2.252, 2.295, 2.295, 2.302, 2.327, 2.323, 2.309, 2.32, 2.323, 2.309, 2.147, 2.134, 2.109, 2.115, 2.123, 2.123, 2.085, 2.087, 2.073, 1.177]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
residuals_STD | 0.252262269536
varray: [2.252, 2.295, 2.295, 2.302, 2.327, 2.323, 2.309, 2.32, 2.323, 2.309, 2.147, 2.134, 2.109, 2.115, 2.123, 2.123, 2.085, 2.087, 2.073, 1.177]
inst: std_standards
operation: std_standards(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
slope_shift_MEAN | 29.1265
varray: [29.318, 29.999, 29.769, 29.961, 30.421, 29.972, 30.574, 30.451, 30.219, 29.544, 29.936, 30.247, 30.178, 30.028, 29.92, 29.99, 29.597, 29.528, 29.674, 13.204]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
slope_shift_STD | 3.76219364201
varray: [29.318, 29.999, 29.769, 29.961, 30.421, 29.972, 30.574, 30.451, 30.219, 29.544, 29.936, 30.247, 30.178, 30.028, 29.92, 29.99, 29.597, 29.528, 29.674, 13.204]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
baseline_curvature_MEAN | 63.4424
varray: [64.595, 66.326, 65.594, 66.246, 67.03, 66.879, 65.627, 66.108, 66.205, 66.325, 64.332, 63.225, 61.963, 63.671, 63.368, 63.048, 63.056, 62.495, 62.118, 40.637]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
baseline_curvature_STD | 5.62307895992
varray: [64.595, 66.326, 65.594, 66.246, 67.03, 66.879, 65.627, 66.108, 66.205, 66.325, 64.332, 63.225, 61.963, 63.671, 63.368, 63.048, 63.056, 62.495, 62.118, 40.637]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.1614
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 42.1053
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 0.99355
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 29.1265
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.378393404304
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.5397934043
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.504524539072
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.66592453907
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.756786808609
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.91818680861
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.21614
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.478951296832
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.47250129683
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 1.50310343955
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 43.6084034396
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 45.1115068791
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 9.01862063733
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 51.1239206373
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 54.111723824
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 96.217023824
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 18.8109682101
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 47.9374682101
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 16.8692368798
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 80.3116368798
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -6.95025986094
actual_array: [ 1.93 -15.84]
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931, 2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981, -13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981, -7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935, -9.722, -9.726, -9.712, -9.765, -9.591, -9.592, -9.572, -9.622, -5.295, -5.09, -5.032, -4.996, -5.029, -5.083, -5.029, -5.046, -2.991, -2.802, -2.832, -2.752, -9.028, -9.328, -9.388, -9.456, -13.944, -14.16, -14.237, -14.227, -9.136, -8.848, -8.774, -8.751, -8.891, -8.953, -8.95, -8.943, -7.531, -7.415, -7.419, -7.388, -7.272, -7.254, -7.228, -7.218, -8.65, -8.707, -8.701, -8.717, -4.587, -4.392, -4.314, -4.303, -4.853, -4.848, -4.862, -4.862, -7.902, -8.073, -8.11, -8.104, -7.756, -7.747, -7.722, -7.721, -9.176, -9.192, -9.233, -9.24, -8.953, -8.903, -8.952, -8.944, -8.408, -8.398, -8.36, -8.375, -4.125, -3.892, -3.864, -3.808, -5.325, -5.404, -5.447, -5.429, 4.778, 4.86, 4.928, 4.959, -5.81, -6.0, -6.039, -6.081, -3.869, -3.779, -3.753, -3.761, 2.457, 2.527, 2.554, 2.51, -2.831, -2.839, -2.89, -2.837, -8.512, -8.794, -8.859, -8.893, -4.087, -3.887, -3.82, -3.78, 9.367, 9.475, 9.55, 9.577, -7.725, -8.152, -8.294, -8.33, -2.701, -2.4, -2.389, -2.295, -3.326, -3.359, -3.35, -3.351, 0.678, 0.907, 0.902, 0.945, 3.546, 3.647, 3.69, 3.697, -8.028, -8.603, -8.746, -8.73]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.72555
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0, 5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0, 16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0, 21673.0, 22209.0, 22746.0, 23284.0, 23873.0, 24410.0, 24946.0, 25483.0, 26074.0, 26610.0, 27146.0, 27684.0, 28271.0, 28806.0, 29342.0, 29879.0, 30467.0, 31003.0, 31539.0, 32075.0, 32663.0, 33200.0, 33736.0, 34272.0, 34860.0, 35396.0, 35933.0, 36470.0, 37060.0, 37596.0, 38132.0, 38669.0, 39257.0, 39794.0, 40329.0, 40865.0, 41455.0, 41992.0, 42528.0, 43065.0, 43652.0, 44187.0, 44723.0, 45260.0, 45847.0, 46384.0, 46919.0, 47455.0, 48042.0, 48580.0, 49115.0, 49651.0, 50239.0, 50775.0, 51311.0, 51848.0, 52436.0, 52972.0, 53507.0, 54044.0, 54632.0, 55169.0, 55705.0, 56243.0, 56832.0, 57369.0, 57905.0, 58441.0, 59026.0, 59563.0, 60098.0, 60634.0, 61220.0, 61756.0, 62293.0, 62830.0, 63416.0, 63953.0, 64489.0, 65025.0, 65613.0, 66149.0, 66685.0, 67220.0, 67808.0, 68343.0, 68879.0, 69416.0, 70005.0, 70541.0, 71076.0, 71612.0, 72200.0, 72736.0, 73272.0, 73808.0, 74392.0, 74929.0, 75466.0, 76002.0, 76587.0, 77123.0, 77659.0, 78195.0, 78781.0, 79316.0, 79852.0, 80389.0, 80975.0, 81510.0, 82047.0, 82584.0, 83170.0, 83706.0, 84242.0, 84776.0, 85364.0, 85899.0, 86434.0, 86970.0, 87556.0, 88091.0, 88628.0, 89163.0, 89746.0, 90283.0, 90817.0, 91353.0, 91935.0, 92471.0, 93007.0, 93542.0, 94126.0, 94662.0, 95196.0, 95733.0, 96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -75.6798570178
actual_array: [ 18.1 -119. ]
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107, 7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104, -93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298, -68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874, -88.281, -90.156, -90.437, -91.159, -87.97, -87.668, -87.57, -87.775, -72.605, -70.909, -70.372, -70.268, -73.085, -73.619, -73.336, -73.661, -62.632, -61.178, -61.347, -60.788, -83.964, -86.133, -86.825, -87.026, -113.079, -115.602, -116.622, -116.7, -95.044, -93.231, -92.432, -92.068, -70.504, -68.371, -67.711, -67.376, -80.487, -81.6, -82.036, -81.949, -85.093, -85.123, -85.412, -85.489, -87.993, -88.197, -88.335, -88.511, -73.885, -72.346, -71.983, -71.816, -71.096, -70.741, -70.794, -70.831, -80.099, -80.776, -81.158, -81.171, -81.831, -81.993, -82.092, -82.117, -90.646, -91.328, -91.624, -91.462, -70.45, -67.977, -67.903, -67.372, -83.204, -84.754, -85.044, -85.046, -73.704, -72.473, -72.04, -71.905, -77.475, -78.193, -78.415, -78.368, -53.362, -52.12, -51.442, -51.304, -77.628, -79.07, -79.317, -79.546, -72.2, -71.692, -71.543, -71.447, -55.103, -53.845, -53.621, -53.804, -58.165, -57.824, -57.835, -57.637, -65.481, -66.368, -66.788, -66.901, -68.511, -68.633, -68.699, -68.905, -51.641, -51.452, -51.636, -51.502, -83.014, -85.457, -86.197, -86.295, -65.383, -63.543, -62.965, -62.531, -64.841, -64.61, -64.779, -64.656, -19.961, -15.051, -13.768, -13.058, -6.701, -5.869, -5.487, -5.358, -58.202, -63.513, -65.011, -65.145]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.1700666667
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0, 5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0, 10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0, 16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0, 21673.0, 22209.0, 22746.0, 23284.0, 23873.0, 24410.0, 24946.0, 25483.0, 26074.0, 26610.0, 27146.0, 27684.0, 28271.0, 28806.0, 29342.0, 29879.0, 30467.0, 31003.0, 31539.0, 32075.0, 32663.0, 33200.0, 33736.0, 34272.0, 34860.0, 35396.0, 35933.0, 36470.0, 37060.0, 37596.0, 38132.0, 38669.0, 39257.0, 39794.0, 40329.0, 40865.0, 41455.0, 41992.0, 42528.0, 43065.0, 43652.0, 44187.0, 44723.0, 45260.0, 45847.0, 46384.0, 46919.0, 47455.0, 48042.0, 48580.0, 49115.0, 49651.0, 50239.0, 50775.0, 51311.0, 51848.0, 52436.0, 52972.0, 53507.0, 54044.0, 54632.0, 55169.0, 55705.0, 56243.0, 56832.0, 57369.0, 57905.0, 58441.0, 59026.0, 59563.0, 60098.0, 60634.0, 61220.0, 61756.0, 62293.0, 62830.0, 63416.0, 63953.0, 64489.0, 65025.0, 65613.0, 66149.0, 66685.0, 67220.0, 67808.0, 68343.0, 68879.0, 69416.0, 70005.0, 70541.0, 71076.0, 71612.0, 72200.0, 72736.0, 73272.0, 73808.0, 74392.0, 74929.0, 75466.0, 76002.0, 76587.0, 77123.0, 77659.0, 78195.0, 78781.0, 79316.0, 79852.0, 80389.0, 80975.0, 81510.0, 82047.0, 82584.0, 83170.0, 83706.0, 84242.0, 84776.0, 85364.0, 85899.0, 86434.0, 86970.0, 87556.0, 88091.0, 88628.0, 89163.0, 89746.0, 90283.0, 90817.0, 91353.0, 91935.0, 92471.0, 93007.0, 93542.0, 94126.0, 94662.0, 95196.0, 95733.0, 96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1961
varray: [2.597, 2.19, 2.154, 2.127, 2.154, 2.174, 2.118, 2.157, 2.167, 2.123]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_residuals_STD | 0.142744877316
varray: [2.597, 2.19, 2.154, 2.127, 2.154, 2.174, 2.118, 2.157, 2.167, 2.123]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0753
varray: [-1.548, -0.921, -1.189, -0.974, -1.106, -0.807, -1.02, -1.058, -0.937, -1.193]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.205401530883
varray: [-1.548, -0.921, -1.189, -0.974, -1.106, -0.807, -1.02, -1.058, -0.937, -1.193]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.0197
varray: [213.516, 70.138, 52.902, 49.926, 49.258, 48.292, 46.918, 47.112, 46.506, 45.629]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 51.9694004756
varray: [213.516, 70.138, 52.902, 49.926, 49.258, 48.292, 46.918, 47.112, 46.506, 45.629]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.3392
varray: [31.58, 29.142, 29.291, 28.784, 29.366, 28.965, 28.574, 29.172, 29.386, 29.132]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.827982527996
varray: [31.58, 29.142, 29.291, 28.784, 29.366, 28.965, 28.574, 29.172, 29.386, 29.132]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: slope_shift
SAMP_d18O_STD | 0.412346254176
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.92471269054
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.1997
varray: [72.904, 63.63, 63.355, 62.025, 63.651, 65.293, 62.677, 62.533, 64.044, 61.885]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.22330393127
varray: [72.904, 63.63, 63.355, 62.025, 63.651, 65.293, 62.677, 62.533, 64.044, 61.885]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1961
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0753
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.0197
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.3392
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0347
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0347
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.60544091792
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.7573
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.1399
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_mean | -68.4669
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
d18O_stdev | 0.412346254176
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
dD_stdev | 2.92471269054
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.1399
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.4669
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.412346254176
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.92471269054
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.5072914664
actual_array: [ 1.93 -15.84]
varray: [-10.263, -9.296, -9.122, -9.047, -8.981, -8.99, -8.906, -8.968, -8.895, -8.931]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.1399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.0048923201
actual_array: [ 18.1 -119. ]
varray: [-76.45, -69.611, -68.276, -67.77, -67.344, -67.106, -67.01, -67.101, -66.894, -67.107]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.4669
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [1.0, 538.0, 1074.0, 1610.0, 2146.0, 2683.0, 3220.0, 3757.0, 4293.0, 4830.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9092
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 24.9144
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3055
varray: [2.252, 2.295, 2.295, 2.302, 2.327, 2.323, 2.309, 2.32, 2.323, 2.309]
inst: mean_all
operation: mean_all(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: residuals
SAMP_residuals_STD | 0.0221221357217
varray: [2.252, 2.295, 2.295, 2.302, 2.327, 2.323, 2.309, 2.32, 2.323, 2.309]
inst: std_all
operation: std_all(residuals)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.1049
varray: [-1.106, -1.093, -1.26, -1.101, -0.975, -1.026, -1.023, -1.21, -1.196, -1.059]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0915744384519
varray: [-1.106, -1.093, -1.26, -1.101, -0.975, -1.026, -1.023, -1.21, -1.196, -1.059]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.5056
varray: [41.163, 44.215, 44.289, 44.456, 44.189, 43.91, 43.548, 43.419, 43.109, 42.758]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.992419512549
varray: [41.163, 44.215, 44.289, 44.456, 44.189, 43.91, 43.548, 43.419, 43.109, 42.758]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.0228
varray: [29.318, 29.999, 29.769, 29.961, 30.421, 29.972, 30.574, 30.451, 30.219, 29.544]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.405737682089
varray: [29.318, 29.999, 29.769, 29.961, 30.421, 29.972, 30.574, 30.451, 30.219, 29.544]
inst: std_all
operation: std_all(slope_shift)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: slope_shift
SAMP_d18O_STD | 0.278112387354
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.80071322108
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.0935
varray: [64.595, 66.326, 65.594, 66.246, 67.03, 66.879, 65.627, 66.108, 66.205, 66.325]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.695238927756
varray: [64.595, 66.326, 65.594, 66.246, 67.03, 66.879, 65.627, 66.108, 66.205, 66.325]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3055
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.1049
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.5056
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.0228
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.1441
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1441
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.66697510873
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.6511
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.7985
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.4729
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.7985
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
dD_mean | 17.4729
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
d18O_stdev | 0.278112387354
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
dD_stdev | 3.80071322108
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.7985
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.4729
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.278112387354
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.80071322108
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.061, 2.672, 2.748, 2.833, 2.906, 2.953, 2.937, 2.932, 2.962, 2.981]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.7985
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [7.559, 15.176, 16.981, 17.871, 18.932, 19.27, 19.498, 19.571, 19.767, 20.104]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 17.4729
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [5419.0, 5956.0, 6493.0, 7029.0, 7565.0, 8103.0, 8639.0, 9176.0, 9713.0, 10249.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.7998
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.4003
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0173
varray: [2.147, 2.134, 2.109, 2.115, 2.123, 2.123, 2.085, 2.087, 2.073, 1.177]
inst: mean_all
operation: mean_all(residuals)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
SAMP_residuals_STD | 0.296158384803
varray: [2.147, 2.134, 2.109, 2.115, 2.123, 2.123, 2.085, 2.087, 2.073, 1.177]
inst: std_all
operation: std_all(residuals)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8822
varray: [-0.715, -0.901, -1.023, -0.991, -0.9, -0.985, -0.801, -0.973, -0.922, -0.611]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.133677721904
varray: [-0.715, -0.901, -1.023, -0.991, -0.9, -0.985, -0.801, -0.973, -0.922, -0.611]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.705
varray: [41.241, 42.343, 42.397, 42.311, 42.026, 42.122, 41.566, 41.476, 41.338, 30.23]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.70626265784
varray: [41.241, 42.343, 42.397, 42.311, 42.026, 42.122, 41.566, 41.476, 41.338, 30.23]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.2302
varray: [29.936, 30.247, 30.178, 30.028, 29.92, 29.99, 29.597, 29.528, 29.674, 13.204]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
SAMP_slope_shift_STD | 5.28498121094
varray: [29.936, 30.247, 30.178, 30.028, 29.92, 29.99, 29.597, 29.528, 29.674, 13.204]
inst: std_all
operation: std_all(slope_shift)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: slope_shift
SAMP_d18O_STD | 0.378026174873
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.43141056469
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.7913
varray: [64.332, 63.225, 61.963, 63.671, 63.368, 63.048, 63.056, 62.495, 62.118, 40.637]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 7.1168382813
varray: [64.332, 63.225, 61.963, 63.671, 63.368, 63.048, 63.056, 62.495, 62.118, 40.637]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0173
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8822
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.705
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.2302
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1441
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1441
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.66697510873
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.6511
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.2587
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
dD_mean | -107.7315
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
d18O_stdev | 0.378026174873
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
dD_stdev | 5.43141056469
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.2587
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.7315
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.378026174873
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.43141056469
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.315, -14.148, -14.282, -14.393, -14.45, -14.428, -14.498, -14.544, -14.548, -13.981]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.2587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-93.399, -104.413, -107.493, -109.085, -109.784, -110.069, -110.492, -110.988, -111.294, -110.298]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -107.7315
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [10837.0, 11373.0, 11909.0, 12447.0, 12983.0, 13519.0, 14056.0, 14593.0, 15129.0, 15665.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.088
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4003
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1232
varray: [1.888, 2.18, 2.152, 2.14, 2.152, 2.144, 2.152, 2.155, 2.125, 2.144]
inst: mean_all
operation: mean_all(residuals)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: residuals
SAMP_residuals_STD | 0.0837957835057
varray: [1.888, 2.18, 2.152, 2.14, 2.152, 2.144, 2.152, 2.155, 2.125, 2.144]
inst: std_all
operation: std_all(residuals)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9159
varray: [-0.761, -1.139, -1.019, -0.866, -0.914, -0.95, -1.043, -0.717, -0.893, -0.857]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.128133133888
varray: [-0.761, -1.139, -1.019, -0.866, -0.914, -0.95, -1.043, -0.717, -0.893, -0.857]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.6715
varray: [37.675, 41.691, 41.211, 40.927, 40.871, 41.058, 41.125, 40.781, 40.542, 40.834]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.09669931562
varray: [37.675, 41.691, 41.211, 40.927, 40.871, 41.058, 41.125, 40.781, 40.542, 40.834]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.0391
varray: [24.895, 30.177, 29.701, 29.444, 29.622, 29.571, 29.327, 29.322, 29.075, 29.257]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.48716041502
varray: [24.895, 30.177, 29.701, 29.444, 29.622, 29.571, 29.327, 29.322, 29.075, 29.257]
inst: std_all
operation: std_all(slope_shift)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.594663910681
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.769940755297
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.299
varray: [57.625, 64.036, 63.65, 63.468, 63.643, 63.771, 63.935, 65.389, 63.292, 64.181]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.07581031032
varray: [57.625, 64.036, 63.65, 63.468, 63.643, 63.771, 63.935, 65.389, 63.292, 64.181]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1232
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9159
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.6715
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.0391
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0382
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0382
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.76737299898
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.1434
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.7875
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
dD_mean | -67.5099
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.594663910681
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
dD_stdev | 0.769940755297
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.7875
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.5099
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.594663910681
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.769940755297
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1401650916
actual_array: [ 1.93 -15.84]
varray: [-7.101, -9.072, -8.984, -9.007, -9.018, -8.913, -8.926, -8.988, -8.931, -8.935]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.7875
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.9569682855
actual_array: [ 18.1 -119. ]
varray: [-68.573, -69.02, -67.72, -67.462, -67.599, -66.988, -66.826, -67.367, -66.67, -66.874]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.5099
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [16254.0, 16790.0, 17327.0, 17863.0, 18400.0, 18936.0, 19473.0, 20009.0, 20546.0, 21083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9821
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4338
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.11925
varray: [2.134, 2.089, 2.128, 2.126]
inst: mean_all
operation: mean_all(residuals)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: residuals
SAMP_residuals_STD | 0.0204511613362
varray: [2.134, 2.089, 2.128, 2.126]
inst: std_all
operation: std_all(residuals)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9465
varray: [-1.016, -0.856, -0.928, -0.986]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0705289538653
varray: [-1.016, -0.856, -0.928, -0.986]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.95475
varray: [41.073, 40.703, 41.121, 40.922]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.188036122416
varray: [41.073, 40.703, 41.121, 40.922]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.546
varray: [29.952, 28.891, 29.84, 29.501]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.476907398419
varray: [29.952, 28.891, 29.84, 29.501]
inst: std_all
operation: std_all(slope_shift)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0232576152403
varray: [-9.722, -9.726, -9.712, -9.765]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.22655000578
varray: [-88.281, -90.156, -90.437, -91.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.29225
varray: [63.654, 62.839, 63.787, 62.889]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.497891136026
varray: [63.654, 62.839, 63.787, 62.889]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.11925
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9465
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.95475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.546
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04215
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04215
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.95012491903
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.15015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-9.722, -9.726, -9.712, -9.765]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-88.281, -90.156, -90.437, -91.159]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.73125
varray: [-9.722, -9.726, -9.712, -9.765]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
dD_mean | -90.00825
varray: [-88.281, -90.156, -90.437, -91.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0232576152403
varray: [-9.722, -9.726, -9.712, -9.765]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
dD_stdev | 1.22655000578
varray: [-88.281, -90.156, -90.437, -91.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.73125
varray: [-9.722, -9.726, -9.712, -9.765]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
measured_dD_mean | -90.00825
varray: [-88.281, -90.156, -90.437, -91.159]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0232576152403
varray: [-9.722, -9.726, -9.712, -9.765]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.22655000578
varray: [-88.281, -90.156, -90.437, -91.159]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -11.123353276
actual_array: [ 1.93 -15.84]
varray: [-9.722, -9.726, -9.712, -9.765]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.73125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -99.5928739325
actual_array: [ 18.1 -119. ]
varray: [-88.281, -90.156, -90.437, -91.159]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -90.00825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [21673.0, 22209.0, 22746.0, 23284.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.98605
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.15055
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.02575
varray: [1.953, 2.051, 2.059, 2.04]
inst: mean_all
operation: mean_all(residuals)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: residuals
SAMP_residuals_STD | 0.0491214481328
varray: [1.953, 2.051, 2.059, 2.04]
inst: std_all
operation: std_all(residuals)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89425
varray: [-0.98, -0.793, -0.933, -0.871]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0809253771487
varray: [-0.98, -0.793, -0.933, -0.871]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.80425
varray: [38.964, 40.04, 40.287, 39.926]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.580077221871
varray: [38.964, 40.04, 40.287, 39.926]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.06275
varray: [26.845, 28.431, 28.723, 28.252]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.834719264184
varray: [26.845, 28.431, 28.723, 28.252]
inst: std_all
operation: std_all(slope_shift)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0206619618301
varray: [-9.591, -9.592, -9.572, -9.622]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.17134443868
varray: [-87.97, -87.668, -87.57, -87.775]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.10075
varray: [58.959, 61.892, 61.892, 61.66]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.43201568311
varray: [58.959, 61.892, 61.892, 61.66]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.02575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.80425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.06275
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13565
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13565
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.27602479874
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.34165
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-9.591, -9.592, -9.572, -9.622]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-87.97, -87.668, -87.57, -87.775]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.59425
varray: [-9.591, -9.592, -9.572, -9.622]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
dD_mean | -87.74575
varray: [-87.97, -87.668, -87.57, -87.775]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0206619618301
varray: [-9.591, -9.592, -9.572, -9.622]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
dD_stdev | 0.17134443868
varray: [-87.97, -87.668, -87.57, -87.775]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.59425
varray: [-9.591, -9.592, -9.572, -9.622]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.74575
varray: [-87.97, -87.668, -87.57, -87.775]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0206619618301
varray: [-9.591, -9.592, -9.572, -9.622]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.17134443868
varray: [-87.97, -87.668, -87.57, -87.775]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.9806282098
actual_array: [ 1.93 -15.84]
varray: [-9.591, -9.592, -9.572, -9.622]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.59425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.1154150733
actual_array: [ 18.1 -119. ]
varray: [-87.97, -87.668, -87.57, -87.775]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.74575
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [23873.0, 24410.0, 24946.0, 25483.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.07955
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.30105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.114
varray: [2.067, 2.129, 2.122, 2.138]
inst: mean_all
operation: mean_all(residuals)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: residuals
SAMP_residuals_STD | 0.0320104149718
varray: [2.067, 2.129, 2.122, 2.138]
inst: std_all
operation: std_all(residuals)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.993
varray: [-1.025, -0.933, -1.055, -0.959]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.056639209034
varray: [-1.025, -0.933, -1.055, -0.959]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.6835
varray: [40.021, 40.796, 40.973, 40.944]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.448416844762
varray: [40.021, 40.796, 40.973, 40.944]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.97075
varray: [28.116, 29.163, 29.201, 29.403]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.579486194831
varray: [28.116, 29.163, 29.201, 29.403]
inst: std_all
operation: std_all(slope_shift)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: slope_shift
SAMP_d18O_STD | 0.133569894313
varray: [-5.295, -5.09, -5.032, -4.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.08144733267
varray: [-72.605, -70.909, -70.372, -70.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.85175
varray: [61.017, 63.516, 63.2, 63.674]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23893889411
varray: [61.017, 63.516, 63.2, 63.674]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.114
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.993
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.6835
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.97075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0474
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0474
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.19302304062
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.59065
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-5.295, -5.09, -5.032, -4.996]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-72.605, -70.909, -70.372, -70.268]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.10325
varray: [-5.295, -5.09, -5.032, -4.996]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
dD_mean | -71.0385
varray: [-72.605, -70.909, -70.372, -70.268]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
d18O_stdev | 0.133569894313
varray: [-5.295, -5.09, -5.032, -4.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
dD_stdev | 1.08144733267
varray: [-72.605, -70.909, -70.372, -70.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.10325
varray: [-5.295, -5.09, -5.032, -4.996]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.0385
varray: [-72.605, -70.909, -70.372, -70.268]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.133569894313
varray: [-5.295, -5.09, -5.032, -4.996]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.08144733267
varray: [-72.605, -70.909, -70.372, -70.268]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.30195468776
actual_array: [ 1.93 -15.84]
varray: [-5.295, -5.09, -5.032, -4.996]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.10325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.8208185974
actual_array: [ 18.1 -119. ]
varray: [-72.605, -70.909, -70.372, -70.268]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.0385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [26074.0, 26610.0, 27146.0, 27684.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9913
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.4218
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13375
varray: [2.121, 2.147, 2.138, 2.129]
inst: mean_all
operation: mean_all(residuals)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: residuals
SAMP_residuals_STD | 0.0112361025271
varray: [2.121, 2.147, 2.138, 2.129]
inst: std_all
operation: std_all(residuals)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00225
varray: [-0.988, -0.922, -1.027, -1.072]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0635629609128
varray: [-0.988, -0.922, -1.027, -1.072]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.7885
varray: [40.695, 40.858, 40.925, 40.676]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.122285185802
varray: [40.695, 40.858, 40.925, 40.676]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.36025
varray: [29.447, 29.485, 29.276, 29.233]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.124336572791
varray: [29.447, 29.485, 29.276, 29.233]
inst: std_all
operation: std_all(slope_shift)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0254607541129
varray: [-5.029, -5.083, -5.029, -5.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.26885730416
varray: [-73.085, -73.619, -73.336, -73.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.29875
varray: [63.204, 63.52, 63.499, 62.972]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.261275556964
varray: [63.204, 63.52, 63.499, 62.972]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.7885
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.36025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02765
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02765
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.27926344036
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.14365
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-5.029, -5.083, -5.029, -5.046]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-73.085, -73.619, -73.336, -73.661]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.04675
varray: [-5.029, -5.083, -5.029, -5.046]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
dD_mean | -73.42525
varray: [-73.085, -73.619, -73.336, -73.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0254607541129
varray: [-5.029, -5.083, -5.029, -5.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
dD_stdev | 0.26885730416
varray: [-73.085, -73.619, -73.336, -73.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.04675
varray: [-5.029, -5.083, -5.029, -5.046]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.42525
varray: [-73.085, -73.619, -73.336, -73.661]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0254607541129
varray: [-5.029, -5.083, -5.029, -5.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.26885730416
varray: [-73.085, -73.619, -73.336, -73.661]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.24309362029
actual_array: [ 1.93 -15.84]
varray: [-5.029, -5.083, -5.029, -5.046]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.04675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.4343323797
actual_array: [ 18.1 -119. ]
varray: [-73.085, -73.619, -73.336, -73.661]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.42525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [28271.0, 28806.0, 29342.0, 29879.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.97155
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.3168
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.165
varray: [2.156, 2.166, 2.15, 2.188]
inst: mean_all
operation: mean_all(residuals)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: residuals
SAMP_residuals_STD | 0.0166933120341
varray: [2.156, 2.166, 2.15, 2.188]
inst: std_all
operation: std_all(residuals)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.03825
varray: [-0.984, -1.042, -1.129, -0.998]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0653522506626
varray: [-0.984, -1.042, -1.129, -0.998]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.971
varray: [40.915, 40.973, 40.89, 41.106]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0964814317196
varray: [40.915, 40.973, 40.89, 41.106]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.8005
varray: [29.428, 29.781, 29.718, 30.275]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.35170679076
varray: [29.428, 29.781, 29.718, 30.275]
inst: std_all
operation: std_all(slope_shift)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: slope_shift
SAMP_d18O_STD | 0.103248486672
varray: [-2.991, -2.802, -2.832, -2.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.798896061241
varray: [-62.632, -61.178, -61.347, -60.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.04075
varray: [63.904, 63.886, 63.888, 64.485]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.296276194341
varray: [63.904, 63.886, 63.888, 64.485]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.165
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.03825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.971
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.8005
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0036
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0036
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.166558711946
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.59835
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-2.991, -2.802, -2.832, -2.752]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-62.632, -61.178, -61.347, -60.788]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.84425
varray: [-2.991, -2.802, -2.832, -2.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
dD_mean | -61.48625
varray: [-62.632, -61.178, -61.347, -60.788]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
d18O_stdev | 0.103248486672
varray: [-2.991, -2.802, -2.832, -2.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
dD_stdev | 0.798896061241
varray: [-62.632, -61.178, -61.347, -60.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.84425
varray: [-2.991, -2.802, -2.832, -2.752]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
measured_dD_mean | -61.48625
varray: [-62.632, -61.178, -61.347, -60.788]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.103248486672
varray: [-2.991, -2.802, -2.832, -2.752]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.798896061241
varray: [-62.632, -61.178, -61.347, -60.788]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.94855377788
actual_array: [ 1.93 -15.84]
varray: [-2.991, -2.802, -2.832, -2.752]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.84425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -68.3610146688
actual_array: [ 18.1 -119. ]
varray: [-62.632, -61.178, -61.347, -60.788]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -61.48625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [30467.0, 31003.0, 31539.0, 32075.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9403
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.1343
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0625
varray: [2.1, 2.074, 2.043, 2.033]
inst: mean_all
operation: mean_all(residuals)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: residuals
SAMP_residuals_STD | 0.0304904356589
varray: [2.1, 2.074, 2.043, 2.033]
inst: std_all
operation: std_all(residuals)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.993
varray: [-0.922, -1.039, -0.975, -1.036]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0557673739744
varray: [-0.922, -1.039, -0.975, -1.036]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.25725
varray: [40.545, 40.486, 39.987, 40.011]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.299333008537
varray: [40.545, 40.486, 39.987, 40.011]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.33975
varray: [28.715, 28.789, 27.85, 28.005]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.48116204824
varray: [28.715, 28.789, 27.85, 28.005]
inst: std_all
operation: std_all(slope_shift)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: slope_shift
SAMP_d18O_STD | 0.188722017793
varray: [-9.028, -9.328, -9.388, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.40185710161
varray: [-83.964, -86.133, -86.825, -87.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.61025
varray: [62.846, 61.292, 61.602, 60.701]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.904648136018
varray: [62.846, 61.292, 61.602, 60.701]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.993
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.25725
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.33975
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0989
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0989
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.57573794763
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.83215
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-9.028, -9.328, -9.388, -9.456]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-83.964, -86.133, -86.825, -87.026]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.3
varray: [-9.028, -9.328, -9.388, -9.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
dD_mean | -85.987
varray: [-83.964, -86.133, -86.825, -87.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
d18O_stdev | 0.188722017793
varray: [-9.028, -9.328, -9.388, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
dD_stdev | 1.40185710161
varray: [-83.964, -86.133, -86.825, -87.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.3
varray: [-9.028, -9.328, -9.388, -9.456]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.987
varray: [-83.964, -86.133, -86.825, -87.026]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.188722017793
varray: [-9.028, -9.328, -9.388, -9.456]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.40185710161
varray: [-83.964, -86.133, -86.825, -87.026]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.674081854
actual_array: [ 1.93 -15.84]
varray: [-9.028, -9.328, -9.388, -9.456]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.3
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -95.1895672197
actual_array: [ 18.1 -119. ]
varray: [-83.964, -86.133, -86.825, -87.026]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.987
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [32663.0, 33200.0, 33736.0, 34272.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0428
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.84805
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.98275
varray: [2.009, 1.969, 1.966, 1.987]
inst: mean_all
operation: mean_all(residuals)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: residuals
SAMP_residuals_STD | 0.0198053023203
varray: [2.009, 1.969, 1.966, 1.987]
inst: std_all
operation: std_all(residuals)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.961
varray: [-0.845, -1.103, -1.057, -0.839]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.138708327075
varray: [-0.845, -1.103, -1.057, -0.839]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.7515
varray: [41.913, 41.652, 41.626, 41.815]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.136339038185
varray: [41.913, 41.652, 41.626, 41.815]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.66525
varray: [28.002, 27.552, 27.555, 27.552]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.224504454299
varray: [28.002, 27.552, 27.555, 27.552]
inst: std_all
operation: std_all(slope_shift)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: slope_shift
SAMP_d18O_STD | 0.136354929015
varray: [-13.944, -14.16, -14.237, -14.227]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.69021958633
varray: [-113.079, -115.602, -116.622, -116.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.88775
varray: [60.424, 59.491, 59.286, 60.35]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.583310309069
varray: [60.424, 59.491, 59.286, 60.35]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.98275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.961
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.7515
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.66525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.17865
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17865
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.26547608032
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.55465
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-13.944, -14.16, -14.237, -14.227]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-113.079, -115.602, -116.622, -116.7]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.142
varray: [-13.944, -14.16, -14.237, -14.227]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
dD_mean | -115.50075
varray: [-113.079, -115.602, -116.622, -116.7]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
d18O_stdev | 0.136354929015
varray: [-13.944, -14.16, -14.237, -14.227]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
dD_stdev | 1.69021958633
varray: [-113.079, -115.602, -116.622, -116.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.142
varray: [-13.944, -14.16, -14.237, -14.227]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
measured_dD_mean | -115.50075
varray: [-113.079, -115.602, -116.622, -116.7]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.136354929015
varray: [-13.944, -14.16, -14.237, -14.227]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.69021958633
varray: [-113.079, -115.602, -116.622, -116.7]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.7184232465
actual_array: [ 1.93 -15.84]
varray: [-13.944, -14.16, -14.237, -14.227]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -127.507402096
actual_array: [ 18.1 -119. ]
varray: [-113.079, -115.602, -116.622, -116.7]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -115.50075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [34860.0, 35396.0, 35933.0, 36470.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.12255
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.3538
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.01475
varray: [2.013, 2.015, 2.0, 2.031]
inst: mean_all
operation: mean_all(residuals)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: residuals
SAMP_residuals_STD | 0.0127115433104
varray: [2.013, 2.015, 2.0, 2.031]
inst: std_all
operation: std_all(residuals)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.92625
varray: [-0.807, -1.009, -0.99, -0.899]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0928704294524
varray: [-0.807, -1.009, -0.99, -0.899]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.7385
varray: [39.627, 39.702, 39.734, 39.891]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.111117055397
varray: [39.627, 39.702, 39.734, 39.891]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.8725
varray: [27.981, 27.941, 27.768, 27.8]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.104308836315
varray: [27.981, 27.941, 27.768, 27.8]
inst: std_all
operation: std_all(slope_shift)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: slope_shift
SAMP_d18O_STD | 0.177394804509
varray: [-9.136, -8.848, -8.774, -8.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.32569337204
varray: [-95.044, -93.231, -92.432, -92.068]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.78075
varray: [61.068, 60.3, 60.175, 61.58]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.663165828935
varray: [61.068, 60.3, 60.175, 61.58]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.01475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.92625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.7385
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.8725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.14665
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.14665
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.78495419635
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.66165
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-9.136, -8.848, -8.774, -8.751]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-95.044, -93.231, -92.432, -92.068]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.87725
varray: [-9.136, -8.848, -8.774, -8.751]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
dD_mean | -93.19375
varray: [-95.044, -93.231, -92.432, -92.068]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
d18O_stdev | 0.177394804509
varray: [-9.136, -8.848, -8.774, -8.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
dD_stdev | 1.32569337204
varray: [-95.044, -93.231, -92.432, -92.068]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.87725
varray: [-9.136, -8.848, -8.774, -8.751]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
measured_dD_mean | -93.19375
varray: [-95.044, -93.231, -92.432, -92.068]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.177394804509
varray: [-9.136, -8.848, -8.774, -8.751]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.32569337204
varray: [-95.044, -93.231, -92.432, -92.068]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2336656368
actual_array: [ 1.93 -15.84]
varray: [-9.136, -8.848, -8.774, -8.751]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.87725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -103.081026505
actual_array: [ 18.1 -119. ]
varray: [-95.044, -93.231, -92.432, -92.068]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -93.19375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [37060.0, 37596.0, 38132.0, 38669.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.09055
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.3668
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1225
varray: [2.169, 2.106, 2.106, 2.109]
inst: mean_all
operation: mean_all(residuals)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: residuals
SAMP_residuals_STD | 0.0310322412984
varray: [2.169, 2.106, 2.106, 2.109]
inst: std_all
operation: std_all(residuals)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0395
varray: [-1.211, -1.003, -0.9, -1.044]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129389592575
varray: [-1.211, -1.003, -0.9, -1.044]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.54025
varray: [41.176, 40.303, 40.309, 40.373]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.425015587949
varray: [41.176, 40.303, 40.309, 40.373]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.2205
varray: [30.522, 28.9, 28.627, 28.833]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.875408666471
varray: [30.522, 28.9, 28.627, 28.833]
inst: std_all
operation: std_all(slope_shift)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0291361745373
varray: [-8.891, -8.953, -8.95, -8.943]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.40453942629
varray: [-70.504, -68.371, -67.711, -67.376]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.72225
varray: [63.491, 62.207, 62.943, 62.248]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.613759657086
varray: [63.491, 62.207, 62.943, 62.248]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0395
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.54025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.2205
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0389
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0389
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.79975941519
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.72015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.891, -8.953, -8.95, -8.943]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-70.504, -68.371, -67.711, -67.376]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.93425
varray: [-8.891, -8.953, -8.95, -8.943]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
dD_mean | -68.4905
varray: [-70.504, -68.371, -67.711, -67.376]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0291361745373
varray: [-8.891, -8.953, -8.95, -8.943]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
dD_stdev | 1.40453942629
varray: [-70.504, -68.371, -67.711, -67.376]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.93425
varray: [-8.891, -8.953, -8.95, -8.943]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.4905
varray: [-70.504, -68.371, -67.711, -67.376]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0291361745373
varray: [-8.891, -8.953, -8.95, -8.943]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.40453942629
varray: [-70.504, -68.371, -67.711, -67.376]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2930475987
actual_array: [ 1.93 -15.84]
varray: [-8.891, -8.953, -8.95, -8.943]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.93425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.0307345429
actual_array: [ 18.1 -119. ]
varray: [-70.504, -68.371, -67.711, -67.376]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.4905
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [39257.0, 39794.0, 40329.0, 40865.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9828
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.56505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.03725
varray: [2.046, 2.061, 2.024, 2.018]
inst: mean_all
operation: mean_all(residuals)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: residuals
SAMP_residuals_STD | 0.0198892768429
varray: [2.046, 2.061, 2.024, 2.018]
inst: std_all
operation: std_all(residuals)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.84625
varray: [-0.875, -0.751, -0.837, -0.922]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0723941756405
varray: [-0.875, -0.751, -0.837, -0.922]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.5415
varray: [39.587, 39.904, 39.37, 39.305]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.27007468103
varray: [39.587, 39.904, 39.37, 39.305]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.75875
varray: [27.943, 28.254, 27.397, 27.441]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.412734276906
varray: [27.943, 28.254, 27.397, 27.441]
inst: std_all
operation: std_all(slope_shift)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0633475861155
varray: [-7.531, -7.415, -7.419, -7.388]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.712687402255
varray: [-80.487, -81.6, -82.036, -81.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.29425
varray: [61.292, 61.583, 61.231, 61.071]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.2138712619
varray: [61.292, 61.583, 61.231, 61.071]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.03725
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.84625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.5415
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.75875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12415
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12415
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.74396224669
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.14815
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.531, -7.415, -7.419, -7.388]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-80.487, -81.6, -82.036, -81.949]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.43825
varray: [-7.531, -7.415, -7.419, -7.388]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
dD_mean | -81.518
varray: [-80.487, -81.6, -82.036, -81.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0633475861155
varray: [-7.531, -7.415, -7.419, -7.388]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
dD_stdev | 0.712687402255
varray: [-80.487, -81.6, -82.036, -81.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.43825
varray: [-7.531, -7.415, -7.419, -7.388]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
measured_dD_mean | -81.518
varray: [-80.487, -81.6, -82.036, -81.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0633475861155
varray: [-7.531, -7.415, -7.419, -7.388]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.712687402255
varray: [-80.487, -81.6, -82.036, -81.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.7345315468
actual_array: [ 1.93 -15.84]
varray: [-7.531, -7.415, -7.419, -7.388]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.43825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.2959700298
actual_array: [ 18.1 -119. ]
varray: [-80.487, -81.6, -82.036, -81.949]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.518
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [41455.0, 41992.0, 42528.0, 43065.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.06805
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.5638
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.036
varray: [2.07, 2.026, 2.01, 2.038]
inst: mean_all
operation: mean_all(residuals)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: residuals
SAMP_residuals_STD | 0.0254034118443
varray: [2.07, 2.026, 2.01, 2.038]
inst: std_all
operation: std_all(residuals)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.014
varray: [-1.008, -1.029, -1.014, -1.005]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.010677078252
varray: [-1.008, -1.029, -1.014, -1.005]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.71425
varray: [40.035, 39.768, 39.511, 39.543]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.242491752437
varray: [40.035, 39.768, 39.511, 39.543]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.1185
varray: [28.78, 28.117, 27.785, 27.792]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.467407388331
varray: [28.78, 28.117, 27.785, 27.792]
inst: std_all
operation: std_all(slope_shift)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0245764114549
varray: [-7.272, -7.254, -7.228, -7.218]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.200599725822
varray: [-85.093, -85.123, -85.412, -85.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.561
varray: [61.145, 60.497, 59.929, 60.673]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.502367727201
varray: [61.145, 60.497, 59.929, 60.673]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.036
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.014
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.71425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.1185
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1254
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1254
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.80179513278
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.8814
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.272, -7.254, -7.228, -7.218]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-85.093, -85.123, -85.412, -85.489]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.243
varray: [-7.272, -7.254, -7.228, -7.218]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
dD_mean | -85.27925
varray: [-85.093, -85.123, -85.412, -85.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0245764114549
varray: [-7.272, -7.254, -7.228, -7.218]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
dD_stdev | 0.200599725822
varray: [-85.093, -85.123, -85.412, -85.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.243
varray: [-7.272, -7.254, -7.228, -7.218]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.27925
varray: [-85.093, -85.123, -85.412, -85.489]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0245764114549
varray: [-7.272, -7.254, -7.228, -7.218]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.200599725822
varray: [-85.093, -85.123, -85.412, -85.489]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.53112228267
actual_array: [ 1.93 -15.84]
varray: [-7.272, -7.254, -7.228, -7.218]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.243
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.4145742881
actual_array: [ 18.1 -119. ]
varray: [-85.093, -85.123, -85.412, -85.489]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.27925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [43652.0, 44187.0, 44723.0, 45260.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0693
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.39105
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0275
varray: [2.039, 2.049, 2.013, 2.009]
inst: mean_all
operation: mean_all(residuals)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: residuals
SAMP_residuals_STD | 0.0195533458348
varray: [2.039, 2.049, 2.013, 2.009]
inst: std_all
operation: std_all(residuals)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.91375
varray: [-0.898, -0.847, -0.934, -0.976]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0547380123863
varray: [-0.898, -0.847, -0.934, -0.976]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.5915
varray: [39.61, 39.731, 39.424, 39.601]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.126423890147
varray: [39.61, 39.731, 39.424, 39.601]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.0355
varray: [28.377, 27.986, 28.003, 27.776]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.249979332479
varray: [28.377, 27.986, 28.003, 27.776]
inst: std_all
operation: std_all(slope_shift)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0299040131086
varray: [-8.65, -8.707, -8.701, -8.717]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.218997716883
varray: [-87.993, -88.197, -88.335, -88.511]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.03175
varray: [61.415, 62.134, 60.491, 60.087]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.921375556799
varray: [61.415, 62.134, 60.491, 60.087]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.91375
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.5915
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.0355
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1339
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1339
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -6.19505875821
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.41065
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.65, -8.707, -8.701, -8.717]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-87.993, -88.197, -88.335, -88.511]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.69375
varray: [-8.65, -8.707, -8.701, -8.717]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
dD_mean | -88.259
varray: [-87.993, -88.197, -88.335, -88.511]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0299040131086
varray: [-8.65, -8.707, -8.701, -8.717]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
dD_stdev | 0.218997716883
varray: [-87.993, -88.197, -88.335, -88.511]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.69375
varray: [-8.65, -8.707, -8.701, -8.717]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
measured_dD_mean | -88.259
varray: [-87.993, -88.197, -88.335, -88.511]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0299040131086
varray: [-8.65, -8.707, -8.701, -8.717]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.218997716883
varray: [-87.993, -88.197, -88.335, -88.511]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0424973911
actual_array: [ 1.93 -15.84]
varray: [-8.65, -8.707, -8.701, -8.717]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.69375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.6774286686
actual_array: [ 18.1 -119. ]
varray: [-87.993, -88.197, -88.335, -88.511]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -88.259
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [45847.0, 46384.0, 46919.0, 47455.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0778
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.5138
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1055
varray: [2.092, 2.131, 2.091, 2.108]
inst: mean_all
operation: mean_all(residuals)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: residuals
SAMP_residuals_STD | 0.0186993761037
varray: [2.092, 2.131, 2.091, 2.108]
inst: std_all
operation: std_all(residuals)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04275
varray: [-1.007, -1.125, -0.895, -1.144]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.11564997478
varray: [-1.007, -1.125, -0.895, -1.144]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.392
varray: [40.111, 40.865, 39.996, 40.596]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.408706088365
varray: [40.111, 40.865, 39.996, 40.596]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.9885
varray: [28.568, 29.719, 28.781, 28.886]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.504649383236
varray: [28.568, 29.719, 28.781, 28.886]
inst: std_all
operation: std_all(slope_shift)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: slope_shift
SAMP_d18O_STD | 0.131445806323
varray: [-4.587, -4.392, -4.314, -4.303]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.944609443103
varray: [-73.885, -72.346, -71.983, -71.816]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.384
varray: [62.205, 63.026, 62.215, 62.09]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.431741434966
varray: [62.205, 63.026, 62.215, 62.09]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1055
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04275
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.392
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.9885
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0559
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0559
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.58628666605
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0584
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-4.587, -4.392, -4.314, -4.303]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-73.885, -72.346, -71.983, -71.816]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.399
varray: [-4.587, -4.392, -4.314, -4.303]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
dD_mean | -72.5075
varray: [-73.885, -72.346, -71.983, -71.816]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
d18O_stdev | 0.131445806323
varray: [-4.587, -4.392, -4.314, -4.303]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
dD_stdev | 0.944609443103
varray: [-73.885, -72.346, -71.983, -71.816]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.399
varray: [-4.587, -4.392, -4.314, -4.303]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.5075
varray: [-73.885, -72.346, -71.983, -71.816]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.131445806323
varray: [-4.587, -4.392, -4.314, -4.303]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.944609443103
varray: [-73.885, -72.346, -71.983, -71.816]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.56827492203
actual_array: [ 1.93 -15.84]
varray: [-4.587, -4.392, -4.314, -4.303]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.399
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.4293874656
actual_array: [ 18.1 -119. ]
varray: [-73.885, -72.346, -71.983, -71.816]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.5075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [48042.0, 48580.0, 49115.0, 49651.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9998
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.7133
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.125
varray: [2.634, 2.075, 1.924, 1.867]
inst: mean_all
operation: mean_all(residuals)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: residuals
SAMP_residuals_STD | 0.350497741315
varray: [2.634, 2.075, 1.924, 1.867]
inst: std_all
operation: std_all(residuals)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.93825
varray: [-1.131, -1.05, -0.8, -0.772]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.17925098791
varray: [-1.131, -1.05, -0.8, -0.772]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.63125
varray: [47.434, 40.045, 37.769, 37.277]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.69270940751
varray: [47.434, 40.045, 37.769, 37.277]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.332
varray: [39.173, 28.368, 25.257, 24.53]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: slope_shift
SAMP_slope_shift_STD | 6.76854011832
varray: [39.173, 28.368, 25.257, 24.53]
inst: std_all
operation: std_all(slope_shift)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: slope_shift
SAMP_d18O_STD | 0.00694622199473
varray: [-4.853, -4.848, -4.862, -4.862]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.158043243028
varray: [-71.096, -70.741, -70.794, -70.831]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.89575
varray: [74.776, 62.227, 58.012, 56.568]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.27604979746
varray: [74.776, 62.227, 58.012, 56.568]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.125
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.93825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.63125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.332
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0364
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0364
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.68409364301
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.54665
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-4.853, -4.848, -4.862, -4.862]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-71.096, -70.741, -70.794, -70.831]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.85625
varray: [-4.853, -4.848, -4.862, -4.862]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
dD_mean | -70.8655
varray: [-71.096, -70.741, -70.794, -70.831]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
d18O_stdev | 0.00694622199473
varray: [-4.853, -4.848, -4.862, -4.862]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
dD_stdev | 0.158043243028
varray: [-71.096, -70.741, -70.794, -70.831]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.85625
varray: [-4.853, -4.848, -4.862, -4.862]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
measured_dD_mean | -70.8655
varray: [-71.096, -70.741, -70.794, -70.831]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.00694622199473
varray: [-4.853, -4.848, -4.862, -4.862]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.158043243028
varray: [-71.096, -70.741, -70.794, -70.831]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.04463285299
actual_array: [ 1.93 -15.84]
varray: [-4.853, -4.848, -4.862, -4.862]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.85625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -78.6313819642
actual_array: [ 18.1 -119. ]
varray: [-71.096, -70.741, -70.794, -70.831]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.8655
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [50239.0, 50775.0, 51311.0, 51848.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9803
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.47405
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.07025
varray: [2.102, 2.058, 2.046, 2.075]
inst: mean_all
operation: mean_all(residuals)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: residuals
SAMP_residuals_STD | 0.0242813371406
varray: [2.102, 2.058, 2.046, 2.075]
inst: std_all
operation: std_all(residuals)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.94875
varray: [-0.977, -0.94, -1.032, -0.846]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0782362874034
varray: [-0.977, -0.94, -1.032, -0.846]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.50525
varray: [41.972, 41.369, 41.32, 41.36]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.311894613612
varray: [41.972, 41.369, 41.32, 41.36]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.199
varray: [28.902, 27.96, 27.89, 28.044]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.472876305179
varray: [28.902, 27.96, 27.89, 28.044]
inst: std_all
operation: std_all(slope_shift)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0981813797689
varray: [-7.902, -8.073, -8.11, -8.104]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.502585979377
varray: [-80.099, -80.776, -81.158, -81.171]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.8285
varray: [62.121, 61.664, 61.453, 62.076]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.323971706584
varray: [62.121, 61.664, 61.453, 62.076]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.07025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.94875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.50525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.199
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.21717405385
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.6139
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.902, -8.073, -8.11, -8.104]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-80.099, -80.776, -81.158, -81.171]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.04725
varray: [-7.902, -8.073, -8.11, -8.104]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
dD_mean | -80.801
varray: [-80.099, -80.776, -81.158, -81.171]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0981813797689
varray: [-7.902, -8.073, -8.11, -8.104]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
dD_stdev | 0.502585979377
varray: [-80.099, -80.776, -81.158, -81.171]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.04725
varray: [-7.902, -8.073, -8.11, -8.104]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
measured_dD_mean | -80.801
varray: [-80.099, -80.776, -81.158, -81.171]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0981813797689
varray: [-7.902, -8.073, -8.11, -8.104]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.502585979377
varray: [-80.099, -80.776, -81.158, -81.171]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.36898092887
actual_array: [ 1.93 -15.84]
varray: [-7.902, -8.073, -8.11, -8.104]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.04725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -89.5108482609
actual_array: [ 18.1 -119. ]
varray: [-80.099, -80.776, -81.158, -81.171]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -80.801
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [52436.0, 52972.0, 53507.0, 54044.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.03505
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.60005
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.92325
varray: [1.926, 1.908, 1.918, 1.941]
inst: mean_all
operation: mean_all(residuals)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: residuals
SAMP_residuals_STD | 0.0139373598648
varray: [1.926, 1.908, 1.918, 1.941]
inst: std_all
operation: std_all(residuals)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.85925
varray: [-0.7, -0.842, -0.999, -0.896]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.124550859759
varray: [-0.7, -0.842, -0.999, -0.896]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 38.1605
varray: [38.032, 37.99, 38.085, 38.535]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.25267436224
varray: [38.032, 37.99, 38.085, 38.535]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.58825
varray: [25.419, 25.343, 25.881, 25.71]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.251209573862
varray: [25.419, 25.343, 25.881, 25.71]
inst: std_all
operation: std_all(slope_shift)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: slope_shift
SAMP_d18O_STD | 0.017710637105
varray: [-7.756, -7.747, -7.722, -7.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.129731453395
varray: [-81.831, -81.993, -82.092, -82.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.1165
varray: [58.8, 57.104, 57.05, 59.512]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.23520187284
varray: [58.8, 57.104, 57.05, 59.512]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.92325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.85925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 38.1605
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.58825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.23815
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.23815
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -11.0183214583
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.3259
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.756, -7.747, -7.722, -7.721]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-81.831, -81.993, -82.092, -82.117]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.7365
varray: [-7.756, -7.747, -7.722, -7.721]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
dD_mean | -82.00825
varray: [-81.831, -81.993, -82.092, -82.117]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
d18O_stdev | 0.017710637105
varray: [-7.756, -7.747, -7.722, -7.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
dD_stdev | 0.129731453395
varray: [-81.831, -81.993, -82.092, -82.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.7365
varray: [-7.756, -7.747, -7.722, -7.721]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
measured_dD_mean | -82.00825
varray: [-81.831, -81.993, -82.092, -82.117]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.017710637105
varray: [-7.756, -7.747, -7.722, -7.721]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.129731453395
varray: [-81.831, -81.993, -82.092, -82.117]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.04524505781
actual_array: [ 1.93 -15.84]
varray: [-7.756, -7.747, -7.722, -7.721]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.7365
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -90.832798408
actual_array: [ 18.1 -119. ]
varray: [-81.831, -81.993, -82.092, -82.117]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -82.00825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [54632.0, 55169.0, 55705.0, 56243.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.18205
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.9448
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06675
varray: [2.124, 2.087, 2.018, 2.038]
inst: mean_all
operation: mean_all(residuals)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: residuals
SAMP_residuals_STD | 0.0479261584802
varray: [2.124, 2.087, 2.018, 2.038]
inst: std_all
operation: std_all(residuals)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9695
varray: [-0.971, -0.941, -1.009, -0.957]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0290459406229
varray: [-0.971, -0.941, -1.009, -0.957]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.1305
varray: [40.596, 40.363, 39.663, 39.9]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.425218767225
varray: [40.596, 40.363, 39.663, 39.9]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.402
varray: [29.209, 28.619, 27.714, 28.066]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.654364322581
varray: [29.209, 28.619, 27.714, 28.066]
inst: std_all
operation: std_all(slope_shift)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0311381331061
varray: [-9.176, -9.192, -9.233, -9.24]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.430046509113
varray: [-90.646, -91.328, -91.624, -91.462]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.00125
varray: [63.049, 62.738, 60.686, 61.532]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.09402296594
varray: [63.049, 62.738, 60.686, 61.532]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06675
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9695
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.1305
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.402
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09465
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09465
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.37910613491
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.44115
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-9.176, -9.192, -9.233, -9.24]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-90.646, -91.328, -91.624, -91.462]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.21025
varray: [-9.176, -9.192, -9.233, -9.24]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
dD_mean | -91.265
varray: [-90.646, -91.328, -91.624, -91.462]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0311381331061
varray: [-9.176, -9.192, -9.233, -9.24]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
dD_stdev | 0.430046509113
varray: [-90.646, -91.328, -91.624, -91.462]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.21025
varray: [-9.176, -9.192, -9.233, -9.24]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
measured_dD_mean | -91.265
varray: [-90.646, -91.328, -91.624, -91.462]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0311381331061
varray: [-9.176, -9.192, -9.233, -9.24]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.430046509113
varray: [-90.646, -91.328, -91.624, -91.462]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.5805813088
actual_array: [ 1.93 -15.84]
varray: [-9.176, -9.192, -9.233, -9.24]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.21025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -100.969027047
actual_array: [ 18.1 -119. ]
varray: [-90.646, -91.328, -91.624, -91.462]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -91.265
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [56832.0, 57369.0, 57905.0, 58441.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.03855
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.9748
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.113
varray: [2.16, 2.1, 2.103, 2.089]
inst: mean_all
operation: mean_all(residuals)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: residuals
SAMP_residuals_STD | 0.0319061122671
varray: [2.16, 2.1, 2.103, 2.089]
inst: std_all
operation: std_all(residuals)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.96925
varray: [-0.955, -0.957, -1.065, -0.9]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.069081473638
varray: [-0.955, -0.957, -1.065, -0.9]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.24575
varray: [40.838, 40.179, 40.055, 39.911]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.409739246351
varray: [40.838, 40.179, 40.055, 39.911]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.9105
varray: [29.605, 28.682, 28.831, 28.524]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.47966828816
varray: [29.605, 28.682, 28.831, 28.524]
inst: std_all
operation: std_all(slope_shift)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0236784008469
varray: [-8.953, -8.903, -8.952, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.37630192908
varray: [-70.45, -67.977, -67.903, -67.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.7095
varray: [63.818, 62.868, 61.754, 62.398]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.868697684276
varray: [63.818, 62.868, 61.754, 62.398]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.113
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.96925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.24575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.9105
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0484
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0484
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.2392893495
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.7329
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.953, -8.903, -8.952, -8.944]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-70.45, -67.977, -67.903, -67.372]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.938
varray: [-8.953, -8.903, -8.952, -8.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
dD_mean | -68.4255
varray: [-70.45, -67.977, -67.903, -67.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0236784008469
varray: [-8.953, -8.903, -8.952, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
dD_stdev | 1.37630192908
varray: [-70.45, -67.977, -67.903, -67.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.938
varray: [-8.953, -8.903, -8.952, -8.944]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.4255
varray: [-70.45, -67.977, -67.903, -67.372]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0236784008469
varray: [-8.953, -8.903, -8.952, -8.944]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.37630192908
varray: [-70.45, -67.977, -67.903, -67.372]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2969543067
actual_array: [ 1.93 -15.84]
varray: [-8.953, -8.903, -8.952, -8.944]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.938
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.9595589292
actual_array: [ 18.1 -119. ]
varray: [-70.45, -67.977, -67.903, -67.372]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.4255
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [59026.0, 59563.0, 60098.0, 60634.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9923
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.85955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.967
varray: [2.018, 1.981, 1.948, 1.921]
inst: mean_all
operation: mean_all(residuals)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: residuals
SAMP_residuals_STD | 0.0419285105865
varray: [2.018, 1.981, 1.948, 1.921]
inst: std_all
operation: std_all(residuals)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.974
varray: [-0.812, -1.065, -1.097, -0.922]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132108541233
varray: [-0.812, -1.065, -1.097, -0.922]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 38.704
varray: [39.193, 38.839, 38.709, 38.075]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.466551890647
varray: [39.193, 38.839, 38.709, 38.075]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.782
varray: [27.604, 27.118, 26.692, 25.714]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.803596913882
varray: [27.604, 27.118, 26.692, 25.714]
inst: std_all
operation: std_all(slope_shift)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0217772817404
varray: [-8.408, -8.398, -8.36, -8.375]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.882724570105
varray: [-83.204, -84.754, -85.044, -85.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.513
varray: [60.377, 58.493, 57.43, 57.752]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.31996287827
varray: [60.377, 58.493, 57.43, 57.752]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.967
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.974
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 38.704
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.782
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1944
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1944
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -8.99417044508
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.9294
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.408, -8.398, -8.36, -8.375]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-83.204, -84.754, -85.044, -85.046]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.38525
varray: [-8.408, -8.398, -8.36, -8.375]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
dD_mean | -84.512
varray: [-83.204, -84.754, -85.044, -85.046]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0217772817404
varray: [-8.408, -8.398, -8.36, -8.375]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
dD_stdev | 0.882724570105
varray: [-83.204, -84.754, -85.044, -85.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.38525
varray: [-8.408, -8.398, -8.36, -8.375]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
measured_dD_mean | -84.512
varray: [-83.204, -84.754, -85.044, -85.046]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0217772817404
varray: [-8.408, -8.398, -8.36, -8.375]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.882724570105
varray: [-83.204, -84.754, -85.044, -85.046]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.72110554487
actual_array: [ 1.93 -15.84]
varray: [-8.408, -8.398, -8.36, -8.375]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.38525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -93.5744282949
actual_array: [ 18.1 -119. ]
varray: [-83.204, -84.754, -85.044, -85.046]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -84.512
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [61220.0, 61756.0, 62293.0, 62830.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.1383
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.4013
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.99625
varray: [2.002, 1.994, 1.988, 2.001]
inst: mean_all
operation: mean_all(residuals)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: residuals
SAMP_residuals_STD | 0.00655108133568
varray: [2.002, 1.994, 1.988, 2.001]
inst: std_all
operation: std_all(residuals)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.02825
varray: [-0.979, -1.07, -1.022, -1.042]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.038282937888
varray: [-0.979, -1.07, -1.022, -1.042]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 38.9675
varray: [38.892, 38.983, 38.896, 39.099]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0972025376898
varray: [38.892, 38.983, 38.896, 39.099]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.63525
varray: [26.361, 26.73, 26.727, 26.723]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.182855817518
varray: [26.361, 26.73, 26.727, 26.723]
inst: std_all
operation: std_all(slope_shift)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: slope_shift
SAMP_d18O_STD | 0.139605097806
varray: [-4.125, -3.892, -3.864, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.818993080964
varray: [-73.704, -72.473, -72.04, -71.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.502
varray: [59.514, 59.436, 59.202, 59.856]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.270688012295
varray: [59.514, 59.436, 59.202, 59.856]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.99625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.02825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 38.9675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.63525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.16515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.64088091052
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.9404
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-4.125, -3.892, -3.864, -3.808]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-73.704, -72.473, -72.04, -71.905]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.92225
varray: [-4.125, -3.892, -3.864, -3.808]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
dD_mean | -72.5305
varray: [-73.704, -72.473, -72.04, -71.905]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
d18O_stdev | 0.139605097806
varray: [-4.125, -3.892, -3.864, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
dD_stdev | 0.818993080964
varray: [-73.704, -72.473, -72.04, -71.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.92225
varray: [-4.125, -3.892, -3.864, -3.808]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
measured_dD_mean | -72.5305
varray: [-73.704, -72.473, -72.04, -71.905]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.139605097806
varray: [-4.125, -3.892, -3.864, -3.808]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.818993080964
varray: [-73.704, -72.473, -72.04, -71.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.07160210937
actual_array: [ 1.93 -15.84]
varray: [-4.125, -3.892, -3.864, -3.808]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.92225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -80.4545726827
actual_array: [ 18.1 -119. ]
varray: [-73.704, -72.473, -72.04, -71.905]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.5305
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [63416.0, 63953.0, 64489.0, 65025.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.10905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -3.1378
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.052
varray: [2.05, 2.06, 2.052, 2.046]
inst: mean_all
operation: mean_all(residuals)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: residuals
SAMP_residuals_STD | 0.00588784057755
varray: [2.05, 2.06, 2.052, 2.046]
inst: std_all
operation: std_all(residuals)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9605
varray: [-1.162, -0.969, -0.823, -0.888]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.147011337431
varray: [-1.162, -0.969, -0.823, -0.888]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.7865
varray: [39.923, 39.937, 39.68, 39.606]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.168527940315
varray: [39.923, 39.937, 39.68, 39.606]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.6365
varray: [28.09, 27.713, 27.357, 27.386]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.342726810545
varray: [28.09, 27.713, 27.357, 27.386]
inst: std_all
operation: std_all(slope_shift)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0538044298052
varray: [-5.325, -5.404, -5.447, -5.429]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.435764749224
varray: [-77.475, -78.193, -78.415, -78.368]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.365
varray: [60.151, 61.913, 61.617, 61.779]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.818331636115
varray: [60.151, 61.913, 61.617, 61.779]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.052
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9605
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.7865
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.6365
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1094
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1094
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.0615341908
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.0774
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-5.325, -5.404, -5.447, -5.429]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-77.475, -78.193, -78.415, -78.368]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.40125
varray: [-5.325, -5.404, -5.447, -5.429]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
dD_mean | -78.11275
varray: [-77.475, -78.193, -78.415, -78.368]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0538044298052
varray: [-5.325, -5.404, -5.447, -5.429]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
dD_stdev | 0.435764749224
varray: [-77.475, -78.193, -78.415, -78.368]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.40125
varray: [-5.325, -5.404, -5.447, -5.429]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.11275
varray: [-77.475, -78.193, -78.415, -78.368]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0538044298052
varray: [-5.325, -5.404, -5.447, -5.429]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.435764749224
varray: [-77.475, -78.193, -78.415, -78.368]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.61240775157
actual_array: [ 1.93 -15.84]
varray: [-5.325, -5.404, -5.447, -5.429]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.40125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.5671891323
actual_array: [ 18.1 -119. ]
varray: [-77.475, -78.193, -78.415, -78.368]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -78.11275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [65613.0, 66149.0, 66685.0, 67220.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0533
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.3188
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 3.38375
varray: [3.444, 3.349, 3.374, 3.368]
inst: mean_all
operation: mean_all(residuals)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: residuals
SAMP_residuals_STD | 0.0415561868639
varray: [3.444, 3.349, 3.374, 3.368]
inst: std_all
operation: std_all(residuals)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -5.58175
varray: [-5.706, -5.546, -5.564, -5.511]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0857063786035
varray: [-5.706, -5.546, -5.564, -5.511]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 146.576
varray: [148.623, 145.307, 146.16, 146.214]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.42649570627
varray: [148.623, 145.307, 146.16, 146.214]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 16.466
varray: [16.958, 16.418, 16.241, 16.247]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.338109449735
varray: [16.958, 16.418, 16.241, 16.247]
inst: std_all
operation: std_all(slope_shift)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: slope_shift
SAMP_d18O_STD | 0.080296845102
varray: [4.778, 4.86, 4.928, 4.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.9402531574
varray: [-53.362, -52.12, -51.442, -51.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.499
varray: [65.641, 64.365, 63.801, 64.189]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.796957548012
varray: [65.641, 64.365, 63.801, 64.189]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 3.38375
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 5.58175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 146.576
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.466
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.22235
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.22235
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 56.553622652
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.0566
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [4.778, 4.86, 4.928, 4.959]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-53.362, -52.12, -51.442, -51.304]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 4.88125
varray: [4.778, 4.86, 4.928, 4.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
dD_mean | -52.057
varray: [-53.362, -52.12, -51.442, -51.304]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
d18O_stdev | 0.080296845102
varray: [4.778, 4.86, 4.928, 4.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
dD_stdev | 0.9402531574
varray: [-53.362, -52.12, -51.442, -51.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
measured_d18O_mean | 4.88125
varray: [4.778, 4.86, 4.928, 4.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
measured_dD_mean | -52.057
varray: [-53.362, -52.12, -51.442, -51.304]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.080296845102
varray: [4.778, 4.86, 4.928, 4.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.9402531574
varray: [-53.362, -52.12, -51.442, -51.304]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.09978563305
actual_array: [ 1.93 -15.84]
varray: [4.778, 4.86, 4.928, 4.959]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 4.88125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.0358969014
actual_array: [ 18.1 -119. ]
varray: [-53.362, -52.12, -51.442, -51.304]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -52.057
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [67808.0, 68343.0, 68879.0, 69416.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 38.72155
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 104.4707
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12075
varray: [2.168, 2.116, 2.092, 2.107]
inst: mean_all
operation: mean_all(residuals)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: residuals
SAMP_residuals_STD | 0.0330189339622
varray: [2.168, 2.116, 2.092, 2.107]
inst: std_all
operation: std_all(residuals)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9395
varray: [-0.955, -1.03, -0.928, -0.845]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0763609848025
varray: [-0.955, -1.03, -0.928, -0.845]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.73575
varray: [43.071, 41.499, 41.285, 41.088]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.905851119114
varray: [43.071, 41.499, 41.285, 41.088]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.94375
varray: [29.825, 28.685, 28.505, 28.76]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.597165178154
varray: [29.825, 28.685, 28.505, 28.76]
inst: std_all
operation: std_all(slope_shift)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: slope_shift
SAMP_d18O_STD | 0.119662024051
varray: [-5.81, -6.0, -6.039, -6.081]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.86365671228
varray: [-77.628, -79.07, -79.317, -79.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.66325
varray: [63.652, 61.646, 62.72, 62.635]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.819846072951
varray: [63.652, 61.646, 62.72, 62.635]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9395
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.73575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.94375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04065
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04065
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.88072545572
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.77915
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-5.81, -6.0, -6.039, -6.081]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-77.628, -79.07, -79.317, -79.546]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.9825
varray: [-5.81, -6.0, -6.039, -6.081]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
dD_mean | -78.89025
varray: [-77.628, -79.07, -79.317, -79.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
d18O_stdev | 0.119662024051
varray: [-5.81, -6.0, -6.039, -6.081]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
dD_stdev | 0.86365671228
varray: [-77.628, -79.07, -79.317, -79.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.9825
varray: [-5.81, -6.0, -6.039, -6.081]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
measured_dD_mean | -78.89025
varray: [-77.628, -79.07, -79.317, -79.546]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.119662024051
varray: [-5.81, -6.0, -6.039, -6.081]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.86365671228
varray: [-77.628, -79.07, -79.317, -79.546]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.21794749431
actual_array: [ 1.93 -15.84]
varray: [-5.81, -6.0, -6.039, -6.081]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.9825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.4185589724
actual_array: [ 18.1 -119. ]
varray: [-77.628, -79.07, -79.317, -79.546]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -78.89025
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [70005.0, 70541.0, 71076.0, 71612.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.98455
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.36955
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09025
varray: [2.079, 2.066, 2.119, 2.097]
inst: mean_all
operation: mean_all(residuals)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: residuals
SAMP_residuals_STD | 0.0229981883344
varray: [2.079, 2.066, 2.119, 2.097]
inst: std_all
operation: std_all(residuals)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.40625
varray: [-1.344, -1.336, -1.424, -1.521]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0862027648435
varray: [-1.344, -1.336, -1.424, -1.521]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 165.0125
varray: [162.089, 163.534, 167.378, 167.049]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.61251711318
varray: [162.089, 163.534, 167.378, 167.049]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.60575
varray: [22.178, 22.338, 23.007, 22.9]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.409163679555
varray: [22.178, 22.338, 23.007, 22.9]
inst: std_all
operation: std_all(slope_shift)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0534509120596
varray: [-3.869, -3.779, -3.753, -3.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.335182040092
varray: [-72.2, -71.692, -71.543, -71.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 57.73475
varray: [57.45, 57.052, 58.626, 57.811]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.670166832463
varray: [57.45, 57.052, 58.626, 57.811]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.40625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 165.0125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.60575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.29184787638
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.70765
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-3.869, -3.779, -3.753, -3.761]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-72.2, -71.692, -71.543, -71.447]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7905
varray: [-3.869, -3.779, -3.753, -3.761]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
dD_mean | -71.7205
varray: [-72.2, -71.692, -71.543, -71.447]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0534509120596
varray: [-3.869, -3.779, -3.753, -3.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
dD_stdev | 0.335182040092
varray: [-72.2, -71.692, -71.543, -71.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7905
varray: [-3.869, -3.779, -3.753, -3.761]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.7205
varray: [-72.2, -71.692, -71.543, -71.447]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0534509120596
varray: [-3.869, -3.779, -3.753, -3.761]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.335182040092
varray: [-72.2, -71.692, -71.543, -71.447]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.93434643435
actual_array: [ 1.93 -15.84]
varray: [-3.869, -3.779, -3.753, -3.761]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.7905
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.5676150359
actual_array: [ 18.1 -119. ]
varray: [-72.2, -71.692, -71.543, -71.447]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.7205
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [72200.0, 72736.0, 73272.0, 73808.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.01505
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 122.9072
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.9025
varray: [2.883, 2.874, 2.939, 2.914]
inst: mean_all
operation: mean_all(residuals)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: residuals
SAMP_residuals_STD | 0.0297601523294
varray: [2.883, 2.874, 2.939, 2.914]
inst: std_all
operation: std_all(residuals)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -4.18525
varray: [-4.166, -4.242, -4.11, -4.223]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0596622437839
varray: [-4.166, -4.242, -4.11, -4.223]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 125.57775
varray: [125.694, 124.367, 126.409, 125.841]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.864042196115
varray: [125.694, 124.367, 126.409, 125.841]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 18.62025
varray: [18.465, 18.672, 18.687, 18.657]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.104222118574
varray: [18.465, 18.672, 18.687, 18.657]
inst: std_all
operation: std_all(slope_shift)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0408982477212
varray: [2.457, 2.527, 2.554, 2.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.680173691053
varray: [-55.103, -53.845, -53.621, -53.804]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.57
varray: [63.372, 62.392, 64.218, 64.298]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.890089134113
varray: [63.372, 62.392, 64.218, 64.298]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.9025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 4.18525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 125.57775
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 18.62025
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.7411
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.7411
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 34.2879615064
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.1276
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [2.457, 2.527, 2.554, 2.51]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-55.103, -53.845, -53.621, -53.804]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.512
varray: [2.457, 2.527, 2.554, 2.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
dD_mean | -54.09325
varray: [-55.103, -53.845, -53.621, -53.804]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0408982477212
varray: [2.457, 2.527, 2.554, 2.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
dD_stdev | 0.680173691053
varray: [-55.103, -53.845, -53.621, -53.804]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.512
varray: [2.457, 2.527, 2.554, 2.51]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.09325
varray: [-55.103, -53.845, -53.621, -53.804]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0408982477212
varray: [2.457, 2.527, 2.554, 2.51]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.680173691053
varray: [-55.103, -53.845, -53.621, -53.804]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.63152750745
actual_array: [ 1.93 -15.84]
varray: [2.457, 2.527, 2.554, 2.51]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.512
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -60.2656098747
actual_array: [ 18.1 -119. ]
varray: [-55.103, -53.845, -53.621, -53.804]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -54.09325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [74392.0, 74929.0, 75466.0, 76002.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2028
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 83.47245
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13275
varray: [2.179, 2.126, 2.118, 2.108]
inst: mean_all
operation: mean_all(residuals)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: residuals
SAMP_residuals_STD | 0.0317004206071
varray: [2.179, 2.126, 2.118, 2.108]
inst: std_all
operation: std_all(residuals)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.54175
varray: [-1.568, -1.571, -1.385, -1.643]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.110101089913
varray: [-1.568, -1.571, -1.385, -1.643]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 410.039
varray: [414.532, 408.554, 408.218, 408.852]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.00650871721
varray: [414.532, 408.554, 408.218, 408.852]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 14.7795
varray: [15.301, 14.692, 14.412, 14.713]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.373763294078
varray: [15.301, 14.692, 14.412, 14.713]
inst: std_all
operation: std_all(slope_shift)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0273785195607
varray: [-2.831, -2.839, -2.89, -2.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.219518222478
varray: [-58.165, -57.824, -57.835, -57.637]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.0105
varray: [54.623, 53.52, 54.652, 53.247]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.732621093517
varray: [54.623, 53.52, 54.652, 53.247]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.54175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 410.039
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 14.7795
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02865
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02865
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.32552974924
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -9.4319
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-2.831, -2.839, -2.89, -2.837]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-58.165, -57.824, -57.835, -57.637]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.84925
varray: [-2.831, -2.839, -2.89, -2.837]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
dD_mean | -57.86525
varray: [-58.165, -57.824, -57.835, -57.637]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0273785195607
varray: [-2.831, -2.839, -2.89, -2.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
dD_stdev | 0.219518222478
varray: [-58.165, -57.824, -57.835, -57.637]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.84925
varray: [-2.831, -2.839, -2.89, -2.837]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
measured_dD_mean | -57.86525
varray: [-58.165, -57.824, -57.835, -57.637]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0273785195607
varray: [-2.831, -2.839, -2.89, -2.837]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.219518222478
varray: [-58.165, -57.824, -57.835, -57.637]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.9537627219
actual_array: [ 1.93 -15.84]
varray: [-2.831, -2.839, -2.89, -2.837]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.84925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -64.3959854845
actual_array: [ 18.1 -119. ]
varray: [-58.165, -57.824, -57.835, -57.637]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -57.86525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [76587.0, 77123.0, 77659.0, 78195.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.97255
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 367.9337
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.108
varray: [2.155, 2.086, 2.094, 2.097]
inst: mean_all
operation: mean_all(residuals)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: residuals
SAMP_residuals_STD | 0.0316754373819
varray: [2.155, 2.086, 2.094, 2.097]
inst: std_all
operation: std_all(residuals)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9575
varray: [-0.983, -1.107, -0.869, -0.871]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113011798794
varray: [-0.983, -1.107, -0.869, -0.871]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.64
varray: [45.616, 42.55, 41.346, 41.048]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.08754848247
varray: [45.616, 42.55, 41.346, 41.048]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.60075
varray: [29.612, 28.41, 27.875, 28.506]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.729086814675
varray: [29.612, 28.41, 27.875, 28.506]
inst: std_all
operation: std_all(slope_shift)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: slope_shift
SAMP_d18O_STD | 0.173271463317
varray: [-8.512, -8.794, -8.859, -8.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.644507305363
varray: [-65.481, -66.368, -66.788, -66.901]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.3945
varray: [63.2, 60.842, 62.895, 62.641]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.05992908568
varray: [63.2, 60.842, 62.895, 62.641]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.108
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9575
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.64
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.60075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0534
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0534
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.47062089387
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.0479
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.512, -8.794, -8.859, -8.893]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-65.481, -66.368, -66.788, -66.901]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.7645
varray: [-8.512, -8.794, -8.859, -8.893]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
dD_mean | -66.3845
varray: [-65.481, -66.368, -66.788, -66.901]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
d18O_stdev | 0.173271463317
varray: [-8.512, -8.794, -8.859, -8.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
dD_stdev | 0.644507305363
varray: [-65.481, -66.368, -66.788, -66.901]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.7645
varray: [-8.512, -8.794, -8.859, -8.893]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.3845
varray: [-65.481, -66.368, -66.788, -66.901]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.173271463317
varray: [-8.512, -8.794, -8.859, -8.893]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.644507305363
varray: [-65.481, -66.368, -66.788, -66.901]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1162039491
actual_array: [ 1.93 -15.84]
varray: [-8.512, -8.794, -8.859, -8.893]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.7645
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.7246446611
actual_array: [ 18.1 -119. ]
varray: [-65.481, -66.368, -66.788, -66.901]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.3845
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [78781.0, 79316.0, 79852.0, 80389.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9973
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.5347
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.09625
varray: [2.106, 2.088, 2.102, 2.089]
inst: mean_all
operation: mean_all(residuals)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: residuals
SAMP_residuals_STD | 0.00910585891977
varray: [2.106, 2.088, 2.102, 2.089]
inst: std_all
operation: std_all(residuals)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95075
varray: [-0.941, -0.975, -0.935, -0.952]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0176328292303
varray: [-0.941, -0.975, -0.935, -0.952]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.7435
varray: [40.629, 40.68, 40.958, 40.707]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.146611732136
varray: [40.629, 40.68, 40.958, 40.707]
inst: std_all
operation: std_all(baseline_shift)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3655
varray: [28.494, 28.251, 28.401, 28.316]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.105408728291
varray: [28.494, 28.251, 28.401, 28.316]
inst: std_all
operation: std_all(slope_shift)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: slope_shift
SAMP_d18O_STD | 0.136343927869
varray: [-4.087, -3.887, -3.82, -3.78]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.164883797466
varray: [-68.511, -68.633, -68.699, -68.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.90225
varray: [61.839, 61.549, 62.28, 61.941]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.301652531897
varray: [61.839, 61.549, 62.28, 61.941]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.09625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95075
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.7435
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3655
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06515
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06515
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.01425002313
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.54015
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-4.087, -3.887, -3.82, -3.78]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-68.511, -68.633, -68.699, -68.905]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.8935
varray: [-4.087, -3.887, -3.82, -3.78]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
dD_mean | -68.687
varray: [-68.511, -68.633, -68.699, -68.905]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
d18O_stdev | 0.136343927869
varray: [-4.087, -3.887, -3.82, -3.78]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
dD_stdev | 0.164883797466
varray: [-68.511, -68.633, -68.699, -68.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.8935
varray: [-4.087, -3.887, -3.82, -3.78]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.687
varray: [-68.511, -68.633, -68.699, -68.905]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.136343927869
varray: [-4.087, -3.887, -3.82, -3.78]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.164883797466
varray: [-68.511, -68.633, -68.699, -68.905]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.04165068124
actual_array: [ 1.93 -15.84]
varray: [-4.087, -3.887, -3.82, -3.78]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.8935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.2459038979
actual_array: [ 18.1 -119. ]
varray: [-68.511, -68.633, -68.699, -68.905]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.687
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [80975.0, 81510.0, 82047.0, 82584.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.00905
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.3618
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 4.0745
varray: [4.046, 4.115, 4.073, 4.064]
inst: mean_all
operation: mean_all(residuals)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: residuals
SAMP_residuals_STD | 0.0292403830344
varray: [4.046, 4.115, 4.073, 4.064]
inst: std_all
operation: std_all(residuals)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -7.47325
varray: [-7.426, -7.588, -7.474, -7.405]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0817694930888
varray: [-7.426, -7.588, -7.474, -7.405]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 175.30325
varray: [173.809, 177.064, 175.424, 174.916]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.35370685527
varray: [173.809, 177.064, 175.424, 174.916]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 9.45875
varray: [9.282, 9.761, 9.586, 9.206]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.259917134231
varray: [9.282, 9.761, 9.586, 9.206]
inst: std_all
operation: std_all(slope_shift)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0939906910284
varray: [9.367, 9.475, 9.55, 9.577]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0954720727054
varray: [-51.641, -51.452, -51.636, -51.502]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.208
varray: [65.555, 64.827, 65.054, 65.396]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.328952884164
varray: [65.555, 64.827, 65.054, 65.396]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 4.0745
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 7.47325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 175.30325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.45875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 1.9131
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 1.9131
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 88.5120755066
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.7656
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [9.367, 9.475, 9.55, 9.577]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-51.641, -51.452, -51.636, -51.502]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 9.49225
varray: [9.367, 9.475, 9.55, 9.577]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
dD_mean | -51.55775
varray: [-51.641, -51.452, -51.636, -51.502]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0939906910284
varray: [9.367, 9.475, 9.55, 9.577]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
dD_stdev | 0.0954720727054
varray: [-51.641, -51.452, -51.636, -51.502]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
measured_d18O_mean | 9.49225
varray: [9.367, 9.475, 9.55, 9.577]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.55775
varray: [-51.641, -51.452, -51.636, -51.502]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0939906910284
varray: [9.367, 9.475, 9.55, 9.577]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0954720727054
varray: [-51.641, -51.452, -51.636, -51.502]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 8.90347381165
actual_array: [ 1.93 -15.84]
varray: [9.367, 9.475, 9.55, 9.577]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 9.49225
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.4892134382
actual_array: [ 18.1 -119. ]
varray: [-51.641, -51.452, -51.636, -51.502]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.55775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [83170.0, 83706.0, 84242.0, 84776.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 38.0308
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 133.19795
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.77775
varray: [1.649, 1.699, 1.902, 1.861]
inst: mean_all
operation: mean_all(residuals)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: residuals
SAMP_residuals_STD | 0.122674026047
varray: [1.649, 1.699, 1.902, 1.861]
inst: std_all
operation: std_all(residuals)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.83325
varray: [-0.812, -0.745, -0.834, -0.942]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.081785797463
varray: [-0.812, -0.745, -0.834, -0.942]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 37.50025
varray: [36.673, 36.475, 38.809, 38.044]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.11713483967
varray: [36.673, 36.475, 38.809, 38.044]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 22.812
varray: [20.268, 21.242, 25.205, 24.533]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.42384859263
varray: [20.268, 21.242, 25.205, 24.533]
inst: std_all
operation: std_all(slope_shift)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: slope_shift
SAMP_d18O_STD | 0.277677720868
varray: [-7.725, -8.152, -8.294, -8.33]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.53090787334
varray: [-83.014, -85.457, -86.197, -86.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 54.72075
varray: [51.165, 52.86, 58.13, 56.728]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.25360245215
varray: [51.165, 52.86, 58.13, 56.728]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.77775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.83325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 37.50025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 22.812
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.38365
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.38365
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -17.7500693995
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.72165
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-7.725, -8.152, -8.294, -8.33]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-83.014, -85.457, -86.197, -86.295]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.12525
varray: [-7.725, -8.152, -8.294, -8.33]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
dD_mean | -85.24075
varray: [-83.014, -85.457, -86.197, -86.295]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
d18O_stdev | 0.277677720868
varray: [-7.725, -8.152, -8.294, -8.33]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
dD_stdev | 1.53090787334
varray: [-83.014, -85.457, -86.197, -86.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.12525
varray: [-7.725, -8.152, -8.294, -8.33]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
measured_dD_mean | -85.24075
varray: [-83.014, -85.457, -86.197, -86.295]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.277677720868
varray: [-7.725, -8.152, -8.294, -8.33]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.53090787334
varray: [-83.014, -85.457, -86.197, -86.295]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.45024045564
actual_array: [ 1.93 -15.84]
varray: [-7.725, -8.152, -8.294, -8.33]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.12525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -94.3724164247
actual_array: [ 18.1 -119. ]
varray: [-83.014, -85.457, -86.197, -86.295]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.24075
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [85364.0, 85899.0, 86434.0, 86970.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.32755
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -4.60505
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0935
varray: [2.108, 2.081, 2.102, 2.083]
inst: mean_all
operation: mean_all(residuals)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: residuals
SAMP_residuals_STD | 0.0135277492585
varray: [2.108, 2.081, 2.102, 2.083]
inst: std_all
operation: std_all(residuals)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0165
varray: [-0.976, -1.003, -1.034, -1.053]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0339656689418
varray: [-0.976, -1.003, -1.034, -1.053]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.397
varray: [40.583, 40.271, 40.428, 40.306]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.141083899389
varray: [40.583, 40.271, 40.428, 40.306]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.3725
varray: [28.619, 28.189, 28.382, 28.3]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.182374157526
varray: [28.619, 28.189, 28.382, 28.3]
inst: std_all
operation: std_all(slope_shift)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: slope_shift
SAMP_d18O_STD | 0.176248640657
varray: [-2.701, -2.4, -2.389, -2.295]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.25541533632
varray: [-65.383, -63.543, -62.965, -62.531]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.76075
varray: [62.577, 61.686, 61.631, 61.149]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.595237977619
varray: [62.577, 61.686, 61.631, 61.149]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0935
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0165
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.397
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.3725
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0679
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0679
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.14148237254
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.68165
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-2.701, -2.4, -2.389, -2.295]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-65.383, -63.543, -62.965, -62.531]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.44625
varray: [-2.701, -2.4, -2.389, -2.295]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
dD_mean | -63.6055
varray: [-65.383, -63.543, -62.965, -62.531]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
d18O_stdev | 0.176248640657
varray: [-2.701, -2.4, -2.389, -2.295]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
dD_stdev | 1.25541533632
varray: [-65.383, -63.543, -62.965, -62.531]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.44625
varray: [-2.701, -2.4, -2.389, -2.295]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.6055
varray: [-65.383, -63.543, -62.965, -62.531]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.176248640657
varray: [-2.701, -2.4, -2.389, -2.295]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.25541533632
varray: [-65.383, -63.543, -62.965, -62.531]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.5339218336
actual_array: [ 1.93 -15.84]
varray: [-2.701, -2.4, -2.389, -2.295]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.44625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.6816134257
actual_array: [ 18.1 -119. ]
varray: [-65.383, -63.543, -62.965, -62.531]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -63.6055
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [87556.0, 88091.0, 88628.0, 89163.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0118
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.7083
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0685
varray: [2.096, 2.054, 2.061, 2.063]
inst: mean_all
operation: mean_all(residuals)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: residuals
SAMP_residuals_STD | 0.0187349939952
varray: [2.096, 2.054, 2.061, 2.063]
inst: std_all
operation: std_all(residuals)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.884
varray: [-0.902, -0.827, -0.795, -1.012]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0963984785495
varray: [-0.902, -0.827, -0.795, -1.012]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.846
varray: [40.106, 39.604, 39.621, 40.053]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.27057839283
varray: [40.106, 39.604, 39.621, 40.053]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.86525
varray: [28.611, 27.477, 27.395, 27.978]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.559978794241
varray: [28.611, 27.477, 27.395, 27.978]
inst: std_all
operation: std_all(slope_shift)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0142478068488
varray: [-3.326, -3.359, -3.35, -3.351]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.106940793588
varray: [-64.841, -64.61, -64.779, -64.656]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.78625
varray: [61.598, 62.004, 62.181, 61.362]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.373572282162
varray: [61.598, 62.004, 62.181, 61.362]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0685
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.884
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.846
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.86525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0929
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0929
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.29814009438
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.65615
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-3.326, -3.359, -3.35, -3.351]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-64.841, -64.61, -64.779, -64.656]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.3465
varray: [-3.326, -3.359, -3.35, -3.351]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
dD_mean | -64.7215
varray: [-64.841, -64.61, -64.779, -64.656]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0142478068488
varray: [-3.326, -3.359, -3.35, -3.351]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
dD_stdev | 0.106940793588
varray: [-64.841, -64.61, -64.779, -64.656]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.3465
varray: [-3.326, -3.359, -3.35, -3.351]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.7215
varray: [-64.841, -64.61, -64.779, -64.656]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0142478068488
varray: [-3.326, -3.359, -3.35, -3.351]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.106940793588
varray: [-64.841, -64.61, -64.779, -64.656]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.47179220505
actual_array: [ 1.93 -15.84]
varray: [-3.326, -3.359, -3.35, -3.351]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.3465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.9036439614
actual_array: [ 18.1 -119. ]
varray: [-64.841, -64.61, -64.779, -64.656]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -64.7215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [89746.0, 90283.0, 90817.0, 91353.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 40.0368
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.2593
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.15225
varray: [2.167, 2.141, 2.154, 2.147]
inst: mean_all
operation: mean_all(residuals)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: residuals
SAMP_residuals_STD | 0.0111766124862
varray: [2.167, 2.141, 2.154, 2.147]
inst: std_all
operation: std_all(residuals)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.008
varray: [-1.061, -0.95, -0.818, -1.203]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.163603178453
varray: [-1.061, -0.95, -0.818, -1.203]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.5625
varray: [43.753, 43.474, 43.371, 43.652]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.172051348924
varray: [43.753, 43.474, 43.371, 43.652]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.019
varray: [28.282, 27.976, 27.601, 28.217]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.308191498909
varray: [28.282, 27.976, 27.601, 28.217]
inst: std_all
operation: std_all(slope_shift)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: slope_shift
SAMP_d18O_STD | 0.121526403167
varray: [0.678, 0.907, 0.902, 0.945]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.11227403464
varray: [-19.961, -15.051, -13.768, -13.058]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.22975
varray: [63.229, 63.116, 64.231, 62.343]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.77497672008
varray: [63.229, 63.116, 64.231, 62.343]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.15225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.008
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.5625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.019
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.00915
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00915
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.423336726196
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.21265
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [0.678, 0.907, 0.902, 0.945]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-19.961, -15.051, -13.768, -13.058]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 0.858
varray: [0.678, 0.907, 0.902, 0.945]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
dD_mean | -15.4595
varray: [-19.961, -15.051, -13.768, -13.058]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
d18O_stdev | 0.121526403167
varray: [0.678, 0.907, 0.902, 0.945]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
dD_stdev | 3.11227403464
varray: [-19.961, -15.051, -13.768, -13.058]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
measured_d18O_mean | 0.858
varray: [0.678, 0.907, 0.902, 0.945]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
measured_dD_mean | -15.4595
varray: [-19.961, -15.051, -13.768, -13.058]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.121526403167
varray: [0.678, 0.907, 0.902, 0.945]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.11227403464
varray: [-19.961, -15.051, -13.768, -13.058]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -0.091591175574
actual_array: [ 1.93 -15.84]
varray: [0.678, 0.907, 0.902, 0.945]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 0.858
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -17.9612889004
actual_array: [ 18.1 -119. ]
varray: [-19.961, -15.051, -13.768, -13.058]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -15.4595
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [91935.0, 92471.0, 93007.0, 93542.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.95305
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.4572
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.188
varray: [2.213, 2.194, 2.175, 2.17]
inst: mean_all
operation: mean_all(residuals)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: residuals
SAMP_residuals_STD | 0.0196129209111
varray: [2.213, 2.194, 2.175, 2.17]
inst: std_all
operation: std_all(residuals)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.958
varray: [-0.846, -0.81, -1.049, -1.127]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.15415360305
varray: [-0.846, -0.81, -1.049, -1.127]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.9115
varray: [40.105, 39.937, 39.815, 39.789]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.144234762338
varray: [40.105, 39.937, 39.815, 39.789]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.47475
varray: [28.811, 28.519, 28.251, 28.318]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.251433722215
varray: [28.811, 28.519, 28.251, 28.318]
inst: std_all
operation: std_all(slope_shift)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0696036397516
varray: [3.546, 3.647, 3.69, 3.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.60507265404
varray: [-6.701, -5.869, -5.487, -5.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.269
varray: [65.137, 64.695, 63.551, 63.693]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.770757203448
varray: [65.137, 64.695, 63.551, 63.693]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.188
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.958
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.9115
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.47475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0266
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0266
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.23068381605
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.8266
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [3.546, 3.647, 3.69, 3.697]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-6.701, -5.869, -5.487, -5.358]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 3.645
varray: [3.546, 3.647, 3.69, 3.697]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
dD_mean | -5.85375
varray: [-6.701, -5.869, -5.487, -5.358]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0696036397516
varray: [3.546, 3.647, 3.69, 3.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
dD_stdev | 0.60507265404
varray: [-6.701, -5.869, -5.487, -5.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
measured_d18O_mean | 3.645
varray: [3.546, 3.647, 3.69, 3.697]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
measured_dD_mean | -5.85375
varray: [-6.701, -5.869, -5.487, -5.358]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0696036397516
varray: [3.546, 3.647, 3.69, 3.697]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.60507265404
varray: [-6.701, -5.869, -5.487, -5.358]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 2.8118742232
actual_array: [ 1.93 -15.84]
varray: [3.546, 3.647, 3.69, 3.697]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 3.645
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -7.4429019667
actual_array: [ 18.1 -119. ]
varray: [-6.701, -5.869, -5.487, -5.358]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -5.85375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [94126.0, 94662.0, 95196.0, 95733.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9173
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -2.1938
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.119
varray: [2.146, 2.126, 2.095, 2.109]
inst: mean_all
operation: mean_all(residuals)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: residuals
SAMP_residuals_STD | 0.0220151463013
varray: [2.146, 2.126, 2.095, 2.109]
inst: std_all
operation: std_all(residuals)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97025
varray: [-0.988, -0.948, -1.036, -0.909]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0544204312123
varray: [-0.988, -0.948, -1.036, -0.909]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.2975
varray: [40.542, 40.295, 40.159, 40.194]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.172897850382
varray: [40.542, 40.295, 40.159, 40.194]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.11175
varray: [29.592, 28.934, 28.874, 29.047]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.328101991257
varray: [29.592, 28.934, 28.874, 29.047]
inst: std_all
operation: std_all(slope_shift)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: slope_shift
SAMP_d18O_STD | 0.338598459339
varray: [-8.028, -8.603, -8.746, -8.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.26215464328
varray: [-58.202, -63.513, -65.011, -65.145]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.47675
varray: [63.091, 62.853, 61.671, 62.292]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.633043639886
varray: [63.091, 62.853, 61.671, 62.292]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.119
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97025
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.2975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.11175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0424
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0424
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.96169149625
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.96565
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2587
varray: [-8.028, -8.603, -8.746, -8.73]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.7315
varray: [-58.202, -63.513, -65.011, -65.145]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.52675
varray: [-8.028, -8.603, -8.746, -8.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
dD_mean | -62.96775
varray: [-58.202, -63.513, -65.011, -65.145]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
d18O_stdev | 0.338598459339
varray: [-8.028, -8.603, -8.746, -8.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
dD_stdev | 3.26215464328
varray: [-58.202, -63.513, -65.011, -65.145]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.52675
varray: [-8.028, -8.603, -8.746, -8.73]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.96775
varray: [-58.202, -63.513, -65.011, -65.145]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.338598459339
varray: [-8.028, -8.603, -8.746, -8.73]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.26215464328
varray: [-58.202, -63.513, -65.011, -65.145]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.86851866074
actual_array: [ 1.93 -15.84]
varray: [-8.028, -8.603, -8.746, -8.73]
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.52675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.983271155
actual_array: [ 18.1 -119. ]
varray: [-58.202, -63.513, -65.011, -65.145]
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -62.96775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: [96318.0, 96854.0, 97389.0, 97925.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.9863
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.8078
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -10.263
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.6773244729
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -10.263
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -84.746462185
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -76.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.296
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.611
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6699146988
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.296
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.257692621
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -69.611
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.122
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.276
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4886434468
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.122
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7958550179
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.276
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.77
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4105092864
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2417802409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.77
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.344
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3417512253
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7753062193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.344
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.99
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.106
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3511273245
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.99
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5146939724
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.106
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.906
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2636170649
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.906
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4095730661
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.968
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.101
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3282079708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.968
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5092189252
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.101
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.895
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2521573881
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.282551971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.931
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.107
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2896617851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.931
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5157889819
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.107
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.061
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.559
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.16168075651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.061
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 7.2441859072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 7.559
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.672
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.176
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.7982137162
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.672
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.584872816
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 15.176
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.748
operation: col(d(18_16)Mean)
inst: col
dD_actual | 16.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.87738966536
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.748
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.5613648562
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 16.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.833
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.96594171376
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.833
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.5359232583
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 17.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.906
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.932
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04199229651
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.906
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.6977282747
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 18.932
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.27
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.09095637033
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0678414656
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 19.27
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.937
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.498
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.07428774945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.937
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3175036181
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 19.498
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.571
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06907880543
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.3974393072
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 19.571
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.962
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.767
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10033246957
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.962
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.6120611576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 19.767
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 20.104
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.7985
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.4729
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | 20.104
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.12012645686
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.9810793391
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: 20.104
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.315
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.399
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.856863905
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -13.315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.305777193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -93.399
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.148
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.413
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.7246739793
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.148
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.366211171
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -104.413
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.282
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8642736792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.282
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.738840248
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -107.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.393
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9799122365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.393
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.482095278
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -109.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.45
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.784
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0392941983
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.45
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.247506877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -109.784
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.428
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.069
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0163748446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.428
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.559584567
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -110.069
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.498
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.492
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.089300061
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.498
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.022773561
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -110.492
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.544
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.988
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.137222346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.544
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.565898243
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -110.988
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.548
operation: col(d(18_16)Mean)
inst: col
dD_actual | -111.294
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1413895012
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.548
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.900971132
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -111.294
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.981
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.298
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5506952489
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -13.981
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.810341729
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -110.298
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.101
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.573
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.3831882724
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.101
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1210728217
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.573
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.072
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.02
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4365540065
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.072
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6105420417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -69.02
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.984
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.72
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3448765917
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.984
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1870297689
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.72
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.007
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.462
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3688377342
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9045173333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.462
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.599
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3802974111
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0545336266
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.599
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.913
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.988
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2709095866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.913
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3854828584
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.988
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.926
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.826
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.284452841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.926
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2080913291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.826
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.988
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.367
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3490437469
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.988
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8004914364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.367
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.931
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.67
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2896617851
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.931
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.0372698563
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.67
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.935
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.874
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2938289403
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.935
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2606517822
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.874
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.722
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.281
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1137167296
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.722
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.7015188763
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -88.281
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.726
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.156
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1178838848
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.726
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -99.7546615774
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -90.156
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.712
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1032988415
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.712
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.06235923
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -90.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.765
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.159
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1585136482
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.765
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.852956046
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -91.159
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.97
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9772423962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.3609709403
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.97
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.592
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.668
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.978284185
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.592
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.0302780893
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.668
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.572
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.57
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.9574484089
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.572
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.9229671641
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.57
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.622
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.775
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.0095378491
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.622
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.1474440994
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.775
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.295
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.605
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.50171769106
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.5361508861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.605
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.09
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.909
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.28815098609
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.09
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.6790148749
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.032
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22772723542
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.032
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.0909948053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.996
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.268
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.19022283845
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.996
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.9771138235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.268
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.029
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22460186901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.029
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.0617554175
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.083
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.619
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.28085846446
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.083
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.6464904588
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.619
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.029
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.336
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.22460186901
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.029
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3366027871
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.336
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.046
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.661
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.24231227869
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.046
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.6924808553
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.661
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.991
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.632
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.10143628497
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.991
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.6156217353
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -62.632
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.802
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.90453820088
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.802
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.0234780088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -61.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.832
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.347
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.93579186502
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.832
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.2085346042
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -61.347
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.752
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.788
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.85244876064
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.752
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5964243269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -60.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.028
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.964
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3907152991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.9743631214
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -83.964
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.328
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.133
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7032519405
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.328
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.349438598
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -86.133
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.388
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7657592688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.1071851309
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -86.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.026
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.8366009075
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -96.3272820284
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.026
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -113.079
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5121490632
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -13.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -124.855562983
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -113.079
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.16
operation: col(d(18_16)Mean)
inst: col
dD_actual | -115.602
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.737175445
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.16
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -127.618271802
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -115.602
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.237
operation: col(d(18_16)Mean)
inst: col
dD_actual | -116.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8173931829
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.237
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -128.735181431
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -116.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.227
operation: col(d(18_16)Mean)
inst: col
dD_actual | -116.7
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8069752949
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -14.227
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -128.820592168
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -116.7
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.136
operation: col(d(18_16)Mean)
inst: col
dD_actual | -95.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.50322849
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.136
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -105.107067723
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -95.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.848
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.231
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2031933143
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.848
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.121815607
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -93.231
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.774
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1261009427
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.774
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -102.246903064
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -92.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.751
operation: col(d(18_16)Mean)
inst: col
dD_actual | -92.068
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1021398002
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.751
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.848319628
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -92.068
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.504
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2479902329
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.2355360514
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.504
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.371
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3125811388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.8998809147
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.371
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.95
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.711
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3094557723
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.95
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.177174684
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.711
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.943
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.376
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3021632507
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.943
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8103465214
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.376
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.531
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.83115745843
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.531
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.1670152966
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -80.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.415
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.6
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.71030995709
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.415
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.3857608039
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.6
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.419
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.036
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.7144771123
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.419
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.86318492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -82.036
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.388
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.68218165936
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.388
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.7679190987
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.272
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.093
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.56133415801
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.272
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.2106287798
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.093
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.254
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.123
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.54258195952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.254
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.243479063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.123
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.228
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.412
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.5154954506
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.228
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.5599367914
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.412
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.218
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.489
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.50507756255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.218
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.6442525183
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.489
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | -87.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99691913092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.3861561575
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -87.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.707
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0563010928
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.707
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.6095380833
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -88.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.701
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.335
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.05005036
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.701
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.7606493861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -88.335
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.717
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0667189808
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.717
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -97.9533710477
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -88.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.587
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.885
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.76413121732
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.587
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.93776297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.885
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.392
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.346
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.56098240039
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.392
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.252543441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.346
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.314
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.983
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.47972287363
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.314
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.855055014
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.983
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.303
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.816
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.46826319677
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.303
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.6721884375
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.816
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.096
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.04124703937
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.8837816403
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.096
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.848
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.741
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.03603809535
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.848
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.4950532889
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.741
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.794
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.05062313862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.5530887892
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.794
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.862
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.05062313862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.5936041385
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.099
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.21766110499
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.7421516336
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -80.099
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.073
operation: col(d(18_16)Mean)
inst: col
dD_actual | -80.776
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.3958069906
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.073
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.4834730249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -80.776
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.11
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.158
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.43435317637
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.11
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.9017666312
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.158
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.104
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.171
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.42810244354
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.104
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -89.9160017539
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.171
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.756
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.831
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.0655599395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.756
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.6387079847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.831
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.747
operation: col(d(18_16)Mean)
inst: col
dD_actual | -81.993
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.05618384026
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.747
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.8160995141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -81.993
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.722
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.092
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.03013912014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.722
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9245054487
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -82.092
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.721
operation: col(d(18_16)Mean)
inst: col
dD_actual | -82.117
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.02909733133
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.721
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -90.9518806847
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -82.117
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.176
operation: col(d(18_16)Mean)
inst: col
dD_actual | -90.646
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5449000422
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.176
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -100.291216203
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -90.646
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.192
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.328
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.5615686631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.192
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.038012642
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -91.328
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.233
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.624
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6042820041
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.233
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.362135436
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -91.624
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.24
operation: col(d(18_16)Mean)
inst: col
dD_actual | -91.462
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.6115745257
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -9.24
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -101.184743907
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -91.462
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3125811388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.1764055417
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -70.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.903
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.977
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2604916985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.903
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.4684471951
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.977
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.952
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.903
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.31153935
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.952
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3874164965
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.903
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.372
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3032050395
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8059664836
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -67.372
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.408
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.204
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.74480624018
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.408
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -92.1421559466
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -83.204
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -84.754
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.73438835213
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -93.8394205795
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -84.754
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.36
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.044
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.69480037755
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.36
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.1569733172
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.044
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.375
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.046
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.71042720962
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.1591633361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.046
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.704
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.28282478953
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.7395662612
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -73.704
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.892
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.473
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.04008799803
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.892
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.3916096399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.473
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.864
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.0109179115
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.864
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9174705522
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.808
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.905
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.95257773843
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.808
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.7696442777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.905
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.325
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.475
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.5329713552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.8688468616
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -77.475
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.404
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.193
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.61527267078
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.404
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.6550636399
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -78.193
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.447
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.415
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.66006958938
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.447
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.8981557357
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -78.415
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.429
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6413173909
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.429
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.846690292
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -78.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.778
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.362
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 3.99222093896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 4.778
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.4648842213
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -53.362
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.86
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.12
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.07764762095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 4.86
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.1048824961
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -52.12
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.928
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.442
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.14848925967
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 4.928
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3624660954
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.442
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 4.959
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.304
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.18078471261
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 4.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.2113547926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.304
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -77.628
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.0382389255
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -5.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.036383306
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -77.628
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.0
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.07
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.2361787984
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -6.0
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.6153869193
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -79.07
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.039
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.317
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.27680856178
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -6.039
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.8858542511
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -79.317
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.081
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.546
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.32056369158
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -6.081
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.136611413
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -79.546
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.869
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.2
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.01612685552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.869
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0926720626
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -72.2
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.779
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.692
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.9223658631
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.779
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.5364072668
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.692
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.753
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.543
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.89527935417
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.753
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.3732508602
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.761
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.447
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.90361366461
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.761
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.2681299539
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -71.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.457
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.103
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.57422912319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.457
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.3712956573
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -55.103
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.527
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.845
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.64715433952
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.527
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.9937737811
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -53.845
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.554
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.621
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.67528263724
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.554
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.7484916664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -53.621
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.51
operation: col(d(18_16)Mean)
inst: col
dD_actual | -53.804
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.62944392984
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 2.51
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -59.9488783941
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -53.804
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.831
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.165
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.93475007621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.831
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.7242145643
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -58.165
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.839
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.824
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.94308438665
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.839
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3508163451
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -57.824
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.89
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.835
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.99621561569
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.89
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.362861449
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -57.835
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.637
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.94100080904
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.1460495797
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -57.637
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.512
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.481
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.85315227587
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.512
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7353036315
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -65.481
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.794
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.368
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1469367188
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.794
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7065770053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.368
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.859
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.788
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2146529911
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.859
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1664809703
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.788
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.901
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2500738105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2902170371
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -66.901
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.087
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.511
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.24323681495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -4.087
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0531822364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.511
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.887
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.633
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.03487905401
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.887
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.1867733882
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.633
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.82
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.699
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.96507920409
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.82
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.2590440112
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.699
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.78
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.905
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.9234076519
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.78
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.484615956
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -68.905
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.367
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.641
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.77298976385
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 9.367
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5803729741
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.641
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.475
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.452
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.88550295476
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 9.475
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3734161898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.452
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.55
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.636
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.96363711512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 9.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5748979269
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.636
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 9.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.502
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 8.99176541285
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 9.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.4281666619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -51.502
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.725
operation: col(d(18_16)Mean)
inst: col
dD_actual | -83.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.03326448655
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -7.725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -91.9341041529
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -83.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.457
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.47810830617
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -94.6092122162
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -85.457
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.294
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.197
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.62604231644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.294
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.4195192022
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -86.197
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.33
operation: col(d(18_16)Mean)
inst: col
dD_actual | -86.295
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.66354671341
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.33
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.5268301274
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -86.295
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.701
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.383
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.7993175316
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.701
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6279927063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -65.383
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.4
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.543
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.48573910138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.4
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.6131753357
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -63.543
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.389
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.965
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.47427942452
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.389
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.980259879
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -62.965
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.295
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.531
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.37635127688
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -2.295
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5050257818
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -62.531
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.326
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.841
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.45043553455
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.326
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0344975895
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -64.841
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.359
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.61
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.48481456511
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.359
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7815504088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -64.61
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.35
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.779
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.47543846587
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.35
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.9666070042
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -64.779
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.351
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.656
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.47648025467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -3.351
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.831920843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -64.656
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -19.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.279113160425
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 0.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -22.8904738971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -19.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.907
operation: col(d(18_16)Mean)
inst: col
dD_actual | -15.051
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0405435241423
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 0.907
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -17.5139775439
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -15.051
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.902
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.768
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.0457524681659
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 0.902
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -16.1090804317
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -13.768
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.945
operation: col(d(18_16)Mean)
inst: col
dD_actual | -13.058
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.000955549562648
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 0.945
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -15.3316237289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -13.058
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.546
operation: col(d(18_16)Mean)
inst: col
dD_actual | -6.701
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.70873713153
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 3.546
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -8.37064871522
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -6.701
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.647
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.869
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.81395780081
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 3.647
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.45960086067
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -5.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.69
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.85875471941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 3.69
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -7.04130725438
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -5.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.697
operation: col(d(18_16)Mean)
inst: col
dD_actual | -5.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.86604724105
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: 3.697
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -6.90005103655
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -5.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:37:29
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.028
operation: col(d(18_16)Mean)
inst: col
dD_actual | -58.202
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.34892649438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.028
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.7647299136
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -58.202
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.603
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.513
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.9479550571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.603
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.5803250525
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -63.513
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.746
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.011
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0969308562
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.746
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.2206491944
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -65.011
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.7985
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.4729
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -13.981
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.298
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2587
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.7315
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.73
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.145
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0802622353
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.945917220034
measure_array: [ 2.7985 -14.2587]
inst: fit_type1
actual_mean: -8.73
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.95988745076
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.3673804595
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.94336214442
measure_array: [ 17.4729 -107.7315]
inst: fit_type1
actual_mean: -65.145
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.913234135667
tarray: []
source_col: d(D_H)Mean