## Summary
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
230622_22-304_23-082_HIDS2046
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.3581 | -75.6037 |  | 11.7066 |  | 0.29274 | 2.1958 |  | 0.51645
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25961 | 3.7384
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37648 | 5.4151
4 |  | EV | -10.4387 | -75.7944 |  |  |  | 0.12125 | 1.7949
5 |  | UNDE4-E-2 | -4.3166 | -66.4771 |  |  |  | 0.16818 | 0.4869
6 |  | UNDE4-010-1 | -7.134 | -57.6312 |  |  |  | 0.09704 | 0.35239
7 |  | UNDE4-1020-1 | -8.9597 | -69.3134 |  |  |  | 0.041508 | 0.50288
8 |  | UNDE4-D-5 | -9.1654 | -81.6101 | -1.5787 | 3.1796 |  | 0.017289 | 0.62266
9 |  | UNDE4-A-4 | -5.9497 | -76.9531 | -3.068 | 515.7761 | 14.8206 | 0.043699 | 0.25728 |  | -8.8284
10 |  | UNDE4-D-4 | -7.7719 | -73.0139 | -1.5633 | 4.8146 |  | 0.025729 | 0.43686
11 |  | UNDE4-C-1 | -6.632 | -73.5976 |  | 127.1108 |  | 0.029228 | 0.088024 |  | -3.8537
12 |  | UNDE4-2030-2 | -7.4286 | -61.2144 |  |  |  | 0.020662 | 0.77066
13 |  | EV | -10.2598 | -73.6605 |  |  |  | 0.086479 | 0.70494
14 |  | UNDE4-010-2 | -5.8224 | -48.4453 |  |  |  | 0.14865 | 1.4553
15 |  | UNDE4-3040-1 | -9.9487 | -75.4558 |  |  |  | 0.097806 | 1.4701
16 |  | UNDE4-C-4 | -4.8355 | -67.5318 |  | 85.5546 |  | 0.15647 | 0.38786 |  | -4.3509
17 |  | UNDE4-2030-1 | -13.2015 | -97.5674 |  |  |  | 0.21643 | 1.5518
18 |  | UNDE4-C-2 | 0.66147 | -52.2214 |  | 8.2551 |  | 0.37873 | 2.4548 |  | -1.4154
19 |  | UNDE4-010-3 | -4.4035 | -45.8823 |  |  |  | 0.11393 | 0.49781
20 |  | UNDE4-B-4 | -6.0233 | -70.8945 | -1.4693 | 227.0998 |  | 0.0667 | 1.3397 |  | -5.6834
21 |  | UNDE4-B-5 | -10.0526 | -86.575 | -1.596 | 360.1991 |  | 0.088055 | 0.7213 |  | -8.9762
22 |  | EV | -10.3335 | -74.9227 |  | 1.6371 |  | 0.024365 | 0.69792
23 |  | UNDE4-C-3 | -9.4267 | -79.2025 | -1.3995 | 242.6571 |  | 0.050675 | 0.21418 |  | -4.0922
24 |  | UNDE4-D-1 | -3.6265 | -66.3482 |  | 7.6481 |  | 0.15652 | 0.56732 |  | 0.9978
25 |  | UNDE4-C-5 | -6.4911 | -71.4518 |  | 1.4571 |  | 0.068485 | 0.25696
26 |  | UNDE4-B-1 | -10.0364 | -87.9511 | -1.7695 | 377.0581 |  | 0.060487 | 0.58344 |  | -6.6279
27 |  | UNDE4-A-2 | -5.8529 | -69.9084 | -2.0023 | 258.6371 |  | 0.089789 | 0.76285 |  | -4.0704
28 |  | UNDE4-D-3 | -6.2551 | -69.765 | -1.4752 | 2.9411 |  | 0.036427 | 0.22073
29 |  | UNDE4-deep-3 | -10.0568 | -79.6912 |  |  |  | 0.069496 | 0.37713
30 |  | UNDE4-E-5 | -4.8793 | -63.1606 | -1.63 | 51.8156 |  | 0.10376 | 0.73468 |  | 0.32505
31 |  | EV | -10.2254 | -73.8251 |  |  |  | 0.11021 | 0.35016
32 |  | UNDE4-B-2 | -8.9921 | -82.021 | -1.505 | 122.0628 |  | 0.051292 | 0.30936 |  | -2.4609
33 |  | UNDE4-1020-2 | -6.0609 | -53.6363 |  |  |  | 0.087774 | 1.3801
34 |  | UNDE4-A-3 | -1.0432 | -58.1196 |  | 1.1951 |  | 0.13476 | 0.20566
35 |  | UNDE4-3040-3 | -9.6386 | -74.7115 |  |  |  | 0.17913 | 0.52973
36 |  | UNDE4-E-4 | -3.8089 | -66.2878 |  |  |  | 0.14663 | 0.56346
37 |  | UNDE4-D-2 | 1.0212 | -55.0514 |  |  |  | 0.16176 | 0.7068
38 |  | UNDE4-E-3 | -6.0158 | -69.2283 | -2.2323 | 1.8223 |  | 0.29505 | 1.0465
39 |  | UNDE4-1020-3 | -3.8332 | -48.6835 |  |  |  | 0.12945 | 1.2818
40 |  | EV | -10.1168 | -72.6571 |  |  |  | 0.17706 | 1.1749
41 |  | UNDE4-A-5 | -5.7383 | -75.3564 |  | 1.8543 |  | 0.11071 | 0.11799
42 |  | UNDE4-B-3 | 4.6594 | -43.6104 |  |  |  | 0.25354 | 1.6308
43 |  | UNDE4-3040-2 | -9.5637 | -73.5023 |  |  |  | 0.33877 | 1.3825
44 |  | UNDE4-A-1 | -4.5316 | -63.104 |  |  |  | 0.21486 | 0.773
45 |  | UNDE4-E-1 | 0.3908 | -53.4789 |  |  |  | 0.14199 | 0.48253
46 |  | UNDE4-2030-3 | -7.0178 | -63.3993 |  |  |  | 0.18825 | 0.46909
47 |  | UNDE4-deep-1 | -9.9594 | -74.0259 |  |  |  | 0.061597 | 0.5249
48 |  | WOOD6-D-3 | -7.4545 | -75.1433 |  | 7.1111 |  | 0.070124 | 0.15879 |  | -1.7194
49 |  | EV | -10.2518 | -74.3829 |  |  |  | 0.05879 | 0.074668
50 |  | WOOD6-D-1 | -8.5278 | -81.8908 |  | 4.4581 |  | 0.045746 | 0.38143
51 |  | WOOD6-C-4 | -4.6885 | -54.8731 | -1.386 | 25.7208 |  | 0.12187 | 1.4054 |  | -4.2019
52 |  | WOOD6-C-2 | -6.7119 | -57.0961 | -1.5917 | 341.9301 |  | 0.055114 | 0.16751 |  | -15.2257
53 |  | EV | -10.1212 | -73.043 |  | 1.2221 |  | 0.08216 | 0.6867

## Detail
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
1 |  | EV | -9.755 | -73.851 | -11.1985 | -82.1498 | -1.148 | 111.837 | 2023/06/22 12:23:53
1 |  | EV | -9.058 | -68.957 | -10.4707 | -76.7727 | -1.195 | 55.478 | 2023/06/22 12:32:51
1 |  | EV | -8.837 | -67.42 | -10.24 | -75.0839 | -0.992 | 47.186 | 2023/06/22 12:41:47
1 |  | EV | -8.857 | -67.358 | -10.2609 | -75.0158 | -1.032 | 47.063 | 2023/06/22 12:50:44
1 |  | EV | -8.846 | -67.007 | -10.2494 | -74.6302 | -1.078 | 46.843 | 2023/06/22 12:59:41
1 |  | EV | -8.895 | -67.268 | -10.3006 | -74.9169 | -0.97 | 46.513 | 2023/06/22 13:08:38
1 |  | EV | -8.81 | -66.746 | -10.2118 | -74.3434 | -1.039 | 45.494 | 2023/06/22 13:17:34
1 |  | EV | -8.822 | -66.708 | -10.2243 | -74.3016 | -0.889 | 44.315 | 2023/06/22 13:26:30
1 |  | EV | -8.837 | -66.86 | -10.24 | -74.4686 | -1.106 | 43.768 | 2023/06/22 13:35:27
1 |  | EV | -8.784 | -66.756 | -10.1847 | -74.3544 | -1.186 | 43.588 | 2023/06/22 13:44:24
Sample 1 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
1 |  | EV | -10.3581 | -75.6037 |  | 11.7066 |  | 0.29274 | 2.1958 |  | 0.51645
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
2 |  | PZ | 2.133 | 7.583 | 1.2133 | 7.3238 | -0.981 | 42.144 | 2023/06/22 13:54:13
2 |  | PZ | 2.673 | 15.076 | 1.777 | 15.5566 | -1.134 | 41.883 | 2023/06/22 14:03:10
2 |  | PZ | 2.829 | 17.295 | 1.9399 | 17.9946 | -1.16 | 41.856 | 2023/06/22 14:12:07
2 |  | PZ | 2.861 | 18.004 | 1.9733 | 18.7736 | -1.12 | 41.65 | 2023/06/22 14:21:04
2 |  | PZ | 2.884 | 18.459 | 1.9973 | 19.2735 | -1.14 | 42.096 | 2023/06/22 14:30:00
2 |  | PZ | 2.925 | 19.137 | 2.0401 | 20.0185 | -1.142 | 41.526 | 2023/06/22 14:38:57
2 |  | PZ | 2.953 | 19.25 | 2.0694 | 20.1426 | -1.071 | 41.695 | 2023/06/22 14:47:54
2 |  | PZ | 2.94 | 19.371 | 2.0558 | 20.2756 | -1.005 | 41.594 | 2023/06/22 14:56:50
2 |  | PZ | 3.011 | 19.783 | 2.1299 | 20.7283 | -0.915 | 41.68 | 2023/06/22 15:05:47
2 |  | PZ | 2.986 | 19.951 | 2.1038 | 20.9128 | -1.111 | 41.289 | 2023/06/22 15:14:44
Sample 2 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
2 |  | PZ | 1.93 | 18.1 |  |  |  | 0.25961 | 3.7384
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
3 |  | UT2 | -13.221 | -93.149 | -14.8171 | -103.353 | -1.065 | 41.734 | 2023/06/22 15:24:33
3 |  | UT2 | -13.964 | -104.054 | -15.5929 | -115.3347 | -1.056 | 41.615 | 2023/06/22 15:33:30
3 |  | UT2 | -14.202 | -107.311 | -15.8414 | -118.9132 | -1.075 | 41.058 | 2023/06/22 15:42:26
3 |  | UT2 | -14.22 | -108.263 | -15.8602 | -119.9592 | -0.915 | 41.273 | 2023/06/22 15:51:24
3 |  | UT2 | -14.277 | -109.013 | -15.9197 | -120.7832 | -0.867 | 41.179 | 2023/06/22 16:00:20
3 |  | UT2 | -14.394 | -109.935 | -16.0418 | -121.7963 | -0.978 | 41.062 | 2023/06/22 16:09:17
3 |  | UT2 | -14.404 | -110.138 | -16.0523 | -122.0193 | -0.756 | 41.091 | 2023/06/22 16:18:13
3 |  | UT2 | -14.449 | -110.462 | -16.0992 | -122.3753 | -0.913 | 41.099 | 2023/06/22 16:27:10
3 |  | UT2 | -14.456 | -110.687 | -16.1065 | -122.6225 | -0.938 | 41.301 | 2023/06/22 16:36:05
3 |  | UT2 | -14.42 | -110.888 | -16.069 | -122.8433 | -0.77 | 41.213 | 2023/06/22 16:45:02
Sample 3 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
3 |  | UT2 | -15.84 | -119 |  |  |  | 0.37648 | 5.4151
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
4 |  | EV | -9.351 | -72.894 | -10.7767 | -81.0984 | -0.903 | 41.13 | 2023/06/22 16:54:52
4 |  | EV | -9.085 | -69.011 | -10.4989 | -76.832 | -0.878 | 41.174 | 2023/06/22 17:03:49
4 |  | EV | -9.023 | -67.995 | -10.4342 | -75.7157 | -0.955 | 41.058 | 2023/06/22 17:12:46
4 |  | EV | -8.944 | -67.433 | -10.3517 | -75.0982 | -0.986 | 41.286 | 2023/06/22 17:21:42
4 |  | EV | -9.004 | -67.396 | -10.4144 | -75.0576 | -0.881 | 41.311 | 2023/06/22 17:30:39
4 |  | EV | -8.992 | -67.375 | -10.4018 | -75.0345 | -0.935 | 40.715 | 2023/06/22 17:39:35
4 |  | EV | -8.997 | -67.196 | -10.4071 | -74.8378 | -1.009 | 41.004 | 2023/06/22 17:48:32
4 |  | EV | -8.96 | -67.127 | -10.3684 | -74.762 | -1.115 | 41.17 | 2023/06/22 17:57:29
4 |  | EV | -8.975 | -67.245 | -10.3841 | -74.8917 | -1.04 | 41.234 | 2023/06/22 18:06:27
4 |  | EV | -8.942 | -66.994 | -10.3496 | -74.6159 | -0.971 | 41.4 | 2023/06/22 18:15:24
Sample 4 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
4 |  | EV | -10.4387 | -75.7944 |  |  |  | 0.12125 | 1.7949
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
5 |  | UNDE4-E-2 | -3.409 | -60.25 | -4.5729 | -67.2061 | -1.392 | 42.449 | 2023/06/22 18:25:14
5 |  | UNDE4-E-2 | -3.136 | -59.653 | -4.2879 | -66.5501 | -1.213 | 42.041 | 2023/06/22 18:34:10
5 |  | UNDE4-E-2 | -3.06 | -59.229 | -4.2085 | -66.0843 | -1.266 | 42.002 | 2023/06/22 18:43:06
5 |  | UNDE4-E-2 | -3.049 | -59.214 | -4.197 | -66.0678 | -1.287 | 42.081 | 2023/06/22 18:52:03
Sample 5 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
5 |  | UNDE4-E-2 | -4.3166 | -66.4771 |  |  |  | 0.16818 | 0.4869
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
6 |  | UNDE4-010-1 | -5.718 | -52.055 | -6.9836 | -58.202 | -1.043 | 41.862 | 2023/06/22 19:01:54
6 |  | UNDE4-010-1 | -5.891 | -51.381 | -7.1642 | -57.4615 | -1.055 | 41.308 | 2023/06/22 19:10:51
6 |  | UNDE4-010-1 | -5.925 | -51.431 | -7.1997 | -57.5164 | -0.814 | 41.531 | 2023/06/22 19:19:47
6 |  | UNDE4-010-1 | -5.914 | -51.275 | -7.1882 | -57.345 | -1.015 | 41.094 | 2023/06/22 19:28:43
Sample 6 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
6 |  | UNDE4-010-1 | -7.134 | -57.6312 |  |  |  | 0.09704 | 0.35239
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
7 |  | UNDE4-1020-1 | -7.55 | -61.447 | -8.8963 | -68.5212 | -1.05 | 42.189 | 2023/06/22 19:38:34
7 |  | UNDE4-1020-1 | -7.643 | -62.229 | -8.9934 | -69.3804 | -0.944 | 41.009 | 2023/06/22 19:47:31
7 |  | UNDE4-1020-1 | -7.621 | -62.591 | -8.9704 | -69.7782 | -1.048 | 41.734 | 2023/06/22 19:56:27
7 |  | UNDE4-1020-1 | -7.629 | -62.405 | -8.9788 | -69.5738 | -0.983 | 41.42 | 2023/06/22 20:05:24
Sample 7 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
7 |  | UNDE4-1020-1 | -8.9597 | -69.3134 |  |  |  | 0.041508 | 0.50288
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
8 |  | UNDE4-D-5 | -7.804 | -72.461 | -9.1615 | -80.6226 | -1.513 | 43.555 | 2023/06/22 20:15:11
8 |  | UNDE4-D-5 | -7.786 | -73.45 | -9.1427 | -81.7092 | -1.569 | 43.78 | 2023/06/22 20:24:08
8 |  | UNDE4-D-5 | -7.814 | -73.663 | -9.1719 | -81.9433 | -1.607 | 45.221 | 2023/06/22 20:33:05
8 |  | UNDE4-D-5 | -7.827 | -73.865 | -9.1855 | -82.1652 | -1.626 | 46.17 | 2023/06/22 20:42:02
Sample 8 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
8 |  | UNDE4-D-5 | -9.1654 | -81.6101 | -1.5787 | 3.1796 |  | 0.017289 | 0.62266
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
9 |  | UNDE4-A-4 | -4.792 | -69.431 | -6.0168 | -77.2935 | -3.109 | 556.171 | 2023/06/22 20:51:50
9 |  | UNDE4-A-4 | -4.716 | -69.178 | -5.9375 | -77.0155 | -3.158 | 561.894 | 2023/06/22 21:00:47
9 |  | UNDE4-A-4 | -4.708 | -68.811 | -5.9291 | -76.6123 | -3.096 | 557.361 | 2023/06/22 21:09:43
9 |  | UNDE4-A-4 | -4.695 | -69.065 | -5.9155 | -76.8913 | -2.909 | 553.686 | 2023/06/22 21:18:38
Sample 9 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
9 |  | UNDE4-A-4 | -5.9497 | -76.9531 | -3.068 | 515.7761 | 14.8206 | 0.043699 | 0.25728 |  | -8.8284
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
10 |  | UNDE4-D-4 | -6.487 | -66.168 | -7.7865 | -73.7083 | -1.426 | 49.509 | 2023/06/22 21:28:27
10 |  | UNDE4-D-4 | -6.448 | -65.487 | -7.7458 | -72.9601 | -1.704 | 46.662 | 2023/06/22 21:37:24
10 |  | UNDE4-D-4 | -6.455 | -65.264 | -7.7531 | -72.7151 | -1.485 | 44.988 | 2023/06/22 21:46:19
10 |  | UNDE4-D-4 | -6.502 | -65.225 | -7.8021 | -72.6722 | -1.638 | 44.107 | 2023/06/22 21:55:16
Sample 10 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
10 |  | UNDE4-D-4 | -7.7719 | -73.0139 | -1.5633 | 4.8146 |  | 0.025729 | 0.43686
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
11 |  | UNDE4-C-1 | -5.424 | -66.081 | -6.6767 | -73.6127 | -1.342 | 170.908 | 2023/06/22 22:05:06
11 |  | UNDE4-C-1 | -5.373 | -66.077 | -6.6234 | -73.6083 | -1.275 | 166.666 | 2023/06/22 22:14:03
11 |  | UNDE4-C-1 | -5.37 | -66.162 | -6.6203 | -73.7017 | -1.247 | 167.803 | 2023/06/22 22:22:59
11 |  | UNDE4-C-1 | -5.358 | -65.949 | -6.6077 | -73.4677 | -1.489 | 169.074 | 2023/06/22 22:31:56
Sample 11 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
11 |  | UNDE4-C-1 | -6.632 | -73.5976 |  | 127.1108 |  | 0.029228 | 0.088024 |  | -3.8537
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
12 |  | UNDE4-2030-2 | -6.142 | -55.83 | -7.4263 | -62.3497 | -1.06 | 42.502 | 2023/06/22 22:41:45
12 |  | UNDE4-2030-2 | -6.166 | -54.896 | -7.4513 | -61.3235 | -0.9 | 40.784 | 2023/06/22 22:50:42
12 |  | UNDE4-2030-2 | -6.152 | -54.408 | -7.4367 | -60.7873 | -1.106 | 40.877 | 2023/06/22 22:59:40
12 |  | UNDE4-2030-2 | -6.117 | -54.053 | -7.4002 | -60.3973 | -0.982 | 40.315 | 2023/06/22 23:08:36
Sample 12 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
12 |  | UNDE4-2030-2 | -7.4286 | -61.2144 |  |  |  | 0.020662 | 0.77066
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
13 |  | EV | -8.735 | -65.085 | -10.1335 | -72.5184 | -0.89 | 41.585 | 2023/06/22 23:18:26
13 |  | EV | -8.853 | -66.291 | -10.2567 | -73.8435 | -0.893 | 41.101 | 2023/06/22 23:27:23
13 |  | EV | -8.911 | -66.534 | -10.3173 | -74.1105 | -0.947 | 40.778 | 2023/06/22 23:36:20
13 |  | EV | -8.925 | -66.588 | -10.3319 | -74.1698 | -0.83 | 40.588 | 2023/06/22 23:45:17
Sample 13 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
13 |  | EV | -10.2598 | -73.6605 |  |  |  | 0.086479 | 0.70494
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
14 |  | UNDE4-010-2 | -4.825 | -45.274 | -6.0513 | -50.7515 | -0.889 | 39.956 | 2023/06/22 23:55:06
14 |  | UNDE4-010-2 | -4.569 | -43.04 | -5.784 | -48.297 | -0.986 | 40.689 | 2023/06/23 00:04:02
14 |  | UNDE4-010-2 | -4.504 | -42.213 | -5.7161 | -47.3883 | -0.837 | 39.52 | 2023/06/23 00:12:58
14 |  | UNDE4-010-2 | -4.525 | -42.173 | -5.7381 | -47.3444 | -0.886 | 39.898 | 2023/06/23 00:21:55
Sample 14 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
14 |  | UNDE4-010-2 | -5.8224 | -48.4453 |  |  |  | 0.14865 | 1.4553
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
15 |  | UNDE4-3040-1 | -8.415 | -65.61 | -9.7994 | -73.0952 | -1.055 | 40.946 | 2023/06/23 00:31:41
15 |  | UNDE4-3040-1 | -8.61 | -68.034 | -10.003 | -75.7585 | -1.057 | 41.124 | 2023/06/23 00:40:38
15 |  | UNDE4-3040-1 | -8.577 | -68.558 | -9.9686 | -76.3343 | -1.087 | 40.143 | 2023/06/23 00:49:34
15 |  | UNDE4-3040-1 | -8.63 | -68.832 | -10.0239 | -76.6353 | -1.046 | 40.407 | 2023/06/23 00:58:31
Sample 15 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
15 |  | UNDE4-3040-1 | -9.9487 | -75.4558 |  |  |  | 0.097806 | 1.4701
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
16 |  | UNDE4-C-4 | -3.885 | -61.094 | -5.0699 | -68.1334 | -1.203 | 127.682 | 2023/06/23 01:08:18
16 |  | UNDE4-C-4 | -3.65 | -60.493 | -4.8245 | -67.4731 | -1.137 | 126.835 | 2023/06/23 01:17:14
16 |  | UNDE4-C-4 | -3.558 | -60.415 | -4.7285 | -67.3874 | -1.292 | 127.125 | 2023/06/23 01:26:11
16 |  | UNDE4-C-4 | -3.549 | -60.184 | -4.7191 | -67.1335 | -1.368 | 126.584 | 2023/06/23 01:35:07
Sample 16 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
16 |  | UNDE4-C-4 | -4.8355 | -67.5318 |  | 85.5546 |  | 0.15647 | 0.38786 |  | -4.3509
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
17 |  | UNDE4-2030-1 | -11.353 | -85.58 | -12.8668 | -95.0368 | -1.122 | 42.035 | 2023/06/23 01:44:54
17 |  | UNDE4-2030-1 | -11.759 | -88.369 | -13.2907 | -98.1011 | -0.885 | 39.903 | 2023/06/23 01:53:50
17 |  | UNDE4-2030-1 | -11.753 | -88.666 | -13.2845 | -98.4275 | -0.982 | 40.548 | 2023/06/23 02:02:46
17 |  | UNDE4-2030-1 | -11.829 | -88.918 | -13.3638 | -98.7043 | -0.957 | 39.695 | 2023/06/23 02:11:43
Sample 17 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
17 |  | UNDE4-2030-1 | -13.2015 | -97.5674 |  |  |  | 0.21643 | 1.5518
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
18 |  | UNDE4-C-2 | 1.047 | -50.222 | 0.079415 | -56.188 | -1.09 | 50.542 | 2023/06/23 02:21:31
18 |  | UNDE4-C-2 | 1.694 | -46.046 | 0.75492 | -51.5998 | -1.378 | 49.593 | 2023/06/23 02:30:28
18 |  | UNDE4-C-2 | 1.807 | -45.302 | 0.8729 | -50.7823 | -1.208 | 49.337 | 2023/06/23 02:39:24
18 |  | UNDE4-C-2 | 1.87 | -44.877 | 0.93867 | -50.3154 | -1.168 | 49.556 | 2023/06/23 02:48:21
Sample 18 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
18 |  | UNDE4-C-2 | 0.66147 | -52.2214 |  | 8.2551 |  | 0.37873 | 2.4548 |  | -1.4154
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
19 |  | UNDE4-010-3 | -3.084 | -41.581 | -4.2336 | -46.694 | -1.103 | 39.88 | 2023/06/23 02:58:09
19 |  | UNDE4-010-3 | -3.257 | -40.662 | -4.4142 | -45.6842 | -0.914 | 39.572 | 2023/06/23 03:07:05
19 |  | UNDE4-010-3 | -3.304 | -40.495 | -4.4633 | -45.5007 | -0.799 | 39.705 | 2023/06/23 03:16:02
19 |  | UNDE4-010-3 | -3.342 | -40.631 | -4.5029 | -45.6502 | -1.023 | 39.656 | 2023/06/23 03:24:59
Sample 19 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
19 |  | UNDE4-010-3 | -4.4035 | -45.8823 |  |  |  | 0.11393 | 0.49781
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
20 |  | UNDE4-B-4 | -4.702 | -61.642 | -5.9228 | -68.7355 | -1.734 | 272.56 | 2023/06/23 03:34:48
20 |  | UNDE4-B-4 | -4.839 | -63.869 | -6.0659 | -71.1824 | -1.34 | 265.883 | 2023/06/23 03:43:45
20 |  | UNDE4-B-4 | -4.847 | -64.506 | -6.0742 | -71.8822 | -1.48 | 270.474 | 2023/06/23 03:52:41
20 |  | UNDE4-B-4 | -4.805 | -64.411 | -6.0304 | -71.7779 | -1.323 | 265.49 | 2023/06/23 04:01:37
Sample 20 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
20 |  | UNDE4-B-4 | -6.0233 | -70.8945 | -1.4693 | 227.0998 |  | 0.0667 | 1.3397 |  | -5.6834
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
21 |  | UNDE4-B-5 | -8.53 | -76.829 | -9.9195 | -85.4218 | -1.682 | 421.791 | 2023/06/23 04:11:26
21 |  | UNDE4-B-5 | -8.693 | -78.004 | -10.0897 | -86.7128 | -1.523 | 398.317 | 2023/06/23 04:20:23
21 |  | UNDE4-B-5 | -8.676 | -78.249 | -10.0719 | -86.982 | -1.617 | 389.799 | 2023/06/23 04:29:19
21 |  | UNDE4-B-5 | -8.731 | -78.432 | -10.1293 | -87.1831 | -1.562 | 396.897 | 2023/06/23 04:38:16
Sample 21 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
21 |  | UNDE4-B-5 | -10.0526 | -86.575 | -1.596 | 360.1991 |  | 0.088055 | 0.7213 |  | -8.9762
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
22 |  | EV | -8.922 | -68.275 | -10.3288 | -76.0233 | -0.953 | 48.965 | 2023/06/23 04:48:01
22 |  | EV | -8.925 | -67.216 | -10.3319 | -74.8598 | -0.951 | 42.09 | 2023/06/23 04:56:57
22 |  | EV | -8.9 | -66.736 | -10.3058 | -74.3324 | -0.808 | 40.977 | 2023/06/23 05:05:53
22 |  | EV | -8.959 | -66.866 | -10.3674 | -74.4752 | -1.058 | 40.524 | 2023/06/23 05:14:51
Sample 22 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
22 |  | EV | -10.3335 | -74.9227 |  | 1.6371 |  | 0.024365 | 0.69792
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
23 |  | UNDE4-C-3 | -8.132 | -71.199 | -9.504 | -79.236 | -1.692 | 324.393 | 2023/06/23 05:24:36
23 |  | UNDE4-C-3 | -8.036 | -70.941 | -9.4037 | -78.9525 | -1.309 | 269.445 | 2023/06/23 05:33:32
23 |  | UNDE4-C-3 | -8.018 | -71.448 | -9.3849 | -79.5096 | -1.223 | 272.565 | 2023/06/23 05:42:28
23 |  | UNDE4-C-3 | -8.046 | -71.086 | -9.4142 | -79.1119 | -1.374 | 270.233 | 2023/06/23 05:51:25
Sample 23 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
23 |  | UNDE4-C-3 | -9.4267 | -79.2025 | -1.3995 | 242.6571 |  | 0.050675 | 0.21418 |  | -4.0922
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
24 |  | UNDE4-D-1 | -2.724 | -60.298 | -3.8577 | -67.2588 | -1.184 | 55.296 | 2023/06/23 06:01:12
24 |  | UNDE4-D-1 | -2.502 | -59.312 | -3.6259 | -66.1755 | -1.234 | 47.675 | 2023/06/23 06:10:08
24 |  | UNDE4-D-1 | -2.39 | -59.253 | -3.509 | -66.1106 | -1.224 | 47.057 | 2023/06/23 06:19:05
24 |  | UNDE4-D-1 | -2.394 | -59.014 | -3.5132 | -65.848 | -1.135 | 46.572 | 2023/06/23 06:28:00
Sample 24 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
24 |  | UNDE4-D-1 | -3.6265 | -66.3482 |  | 7.6481 |  | 0.15652 | 0.56732 |  | 0.9978
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
25 |  | UNDE4-C-5 | -5.156 | -63.753 | -6.3968 | -71.0549 | -1.189 | 47.222 | 2023/06/23 06:37:47
25 |  | UNDE4-C-5 | -5.315 | -64.18 | -6.5629 | -71.5241 | -1.239 | 41.733 | 2023/06/23 06:46:44
25 |  | UNDE4-C-5 | -5.235 | -64.163 | -6.4793 | -71.5054 | -1.19 | 41.768 | 2023/06/23 06:55:40
25 |  | UNDE4-C-5 | -5.279 | -64.361 | -6.5253 | -71.7229 | -1.138 | 41.113 | 2023/06/23 07:04:37
Sample 25 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
25 |  | UNDE4-C-5 | -6.4911 | -71.4518 |  | 1.4571 |  | 0.068485 | 0.25696
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
26 |  | UNDE4-B-1 | -8.562 | -78.302 | -9.9529 | -87.0403 | -2.031 | 483.103 | 2023/06/23 07:14:25
26 |  | UNDE4-B-1 | -8.63 | -79.146 | -10.0239 | -87.9676 | -1.675 | 400.646 | 2023/06/23 07:23:21
26 |  | UNDE4-B-1 | -8.698 | -79.573 | -10.0949 | -88.4367 | -1.614 | 397.715 | 2023/06/23 07:32:18
26 |  | UNDE4-B-1 | -8.678 | -79.503 | -10.074 | -88.3598 | -1.758 | 392.776 | 2023/06/23 07:41:14
Sample 26 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
26 |  | UNDE4-B-1 | -10.0364 | -87.9511 | -1.7695 | 377.0581 |  | 0.060487 | 0.58344 |  | -6.6279
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
27 |  | UNDE4-A-2 | -4.769 | -63.815 | -5.9928 | -71.123 | -2.315 | 346.909 | 2023/06/23 07:51:02
27 |  | UNDE4-A-2 | -4.602 | -62.614 | -5.8184 | -69.8035 | -1.84 | 288.991 | 2023/06/23 07:59:57
27 |  | UNDE4-A-2 | -4.589 | -62.252 | -5.8049 | -69.4057 | -1.895 | 280.568 | 2023/06/23 08:08:54
27 |  | UNDE4-A-2 | -4.58 | -62.157 | -5.7955 | -69.3013 | -1.959 | 284.088 | 2023/06/23 08:17:50
Sample 27 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
27 |  | UNDE4-A-2 | -5.8529 | -69.9084 | -2.0023 | 258.6371 |  | 0.089789 | 0.76285 |  | -4.0704
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
28 |  | UNDE4-D-3 | -5.054 | -62.839 | -6.2904 | -70.0507 | -1.604 | 51.321 | 2023/06/23 08:27:38
28 |  | UNDE4-D-3 | -4.98 | -62.637 | -6.2131 | -69.8287 | -1.427 | 42.974 | 2023/06/23 08:36:34
28 |  | UNDE4-D-3 | -5.048 | -62.531 | -6.2841 | -69.7123 | -1.336 | 41.83 | 2023/06/23 08:45:31
28 |  | UNDE4-D-3 | -4.999 | -62.309 | -6.2329 | -69.4683 | -1.534 | 41.647 | 2023/06/23 08:54:27
Sample 28 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
28 |  | UNDE4-D-3 | -6.2551 | -69.765 | -1.4752 | 2.9411 |  | 0.036427 | 0.22073
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
29 |  | UNDE4-deep-3 | -8.559 | -71.1 | -9.9498 | -79.1272 | -1.226 | 48.86 | 2023/06/23 09:04:13
29 |  | UNDE4-deep-3 | -8.678 | -71.573 | -10.074 | -79.6469 | -0.906 | 40.591 | 2023/06/23 09:13:08
29 |  | UNDE4-deep-3 | -8.707 | -71.955 | -10.1043 | -80.0666 | -0.948 | 40.314 | 2023/06/23 09:22:04
29 |  | UNDE4-deep-3 | -8.702 | -71.825 | -10.0991 | -79.9238 | -0.951 | 40.283 | 2023/06/23 09:30:59
Sample 29 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
29 |  | UNDE4-deep-3 | -10.0568 | -79.6912 |  |  |  | 0.069496 | 0.37713
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
30 |  | UNDE4-E-5 | -3.844 | -57.627 | -5.0271 | -64.3241 | -1.895 | 107.279 | 2023/06/23 09:40:45
30 |  | UNDE4-E-5 | -3.713 | -56.447 | -4.8903 | -63.0276 | -1.577 | 89.553 | 2023/06/23 09:49:41
30 |  | UNDE4-E-5 | -3.644 | -56.246 | -4.8182 | -62.8068 | -1.413 | 88.398 | 2023/06/23 09:58:37
30 |  | UNDE4-E-5 | -3.609 | -55.952 | -4.7817 | -62.4837 | -1.635 | 88.04 | 2023/06/23 10:07:33
Sample 30 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
30 |  | UNDE4-E-5 | -4.8793 | -63.1606 | -1.63 | 51.8156 |  | 0.10376 | 0.73468 |  | 0.32505
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
31 |  | EV | -8.666 | -65.926 | -10.0615 | -73.4424 | -1.039 | 47.098 | 2023/06/23 10:17:19
31 |  | EV | -8.848 | -66.072 | -10.2515 | -73.6028 | -0.849 | 40.268 | 2023/06/23 10:26:15
31 |  | EV | -8.924 | -66.714 | -10.3308 | -74.3082 | -1.086 | 40.017 | 2023/06/23 10:35:11
31 |  | EV | -8.854 | -66.385 | -10.2578 | -73.9467 | -0.938 | 39.677 | 2023/06/23 10:44:06
Sample 31 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
31 |  | EV | -10.2254 | -73.8251 |  |  |  | 0.11021 | 0.35016
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
32 |  | UNDE4-B-2 | -7.714 | -73.297 | -9.0675 | -81.5411 | -1.826 | 186.8 | 2023/06/23 10:53:52
32 |  | UNDE4-B-2 | -7.62 | -73.743 | -8.9694 | -82.0312 | -1.51 | 157.46 | 2023/06/23 11:02:48
32 |  | UNDE4-B-2 | -7.638 | -73.997 | -8.9882 | -82.3102 | -1.304 | 154.986 | 2023/06/23 11:11:45
32 |  | UNDE4-B-2 | -7.595 | -73.898 | -8.9433 | -82.2015 | -1.38 | 155.013 | 2023/06/23 11:20:41
Sample 32 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
32 |  | UNDE4-B-2 | -8.9921 | -82.021 | -1.505 | 122.0628 |  | 0.051292 | 0.30936 |  | -2.4609
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
33 |  | UNDE4-1020-2 | -4.95 | -49.862 | -6.1818 | -55.7925 | -1.141 | 48.462 | 2023/06/23 11:30:28
33 |  | UNDE4-1020-2 | -4.854 | -47.844 | -6.0815 | -53.5753 | -0.946 | 40.086 | 2023/06/23 11:39:23
33 |  | UNDE4-1020-2 | -4.775 | -47.09 | -5.9991 | -52.7468 | -1.163 | 40.774 | 2023/06/23 11:48:19
33 |  | UNDE4-1020-2 | -4.758 | -46.802 | -5.9813 | -52.4304 | -0.971 | 40.683 | 2023/06/23 11:57:15
Sample 33 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
33 |  | UNDE4-1020-2 | -6.0609 | -53.6363 |  |  |  | 0.087774 | 1.3801
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
34 |  | UNDE4-A-3 | -0.222 | -52.01 | -1.2455 | -58.1526 | -1.113 | 49.513 | 2023/06/23 12:07:02
34 |  | UNDE4-A-3 | 0.002 | -51.714 | -1.0116 | -57.8273 | -0.97 | 40.309 | 2023/06/23 12:15:58
34 |  | UNDE4-A-3 | 0.017 | -52.215 | -0.99596 | -58.3778 | -1.072 | 40.589 | 2023/06/23 12:24:53
34 |  | UNDE4-A-3 | 0.09 | -51.981 | -0.91974 | -58.1207 | -1.038 | 40.377 | 2023/06/23 12:33:49
Sample 34 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
34 |  | UNDE4-A-3 | -1.0432 | -58.1196 |  | 1.1951 |  | 0.13476 | 0.20566
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
35 |  | UNDE4-3040-3 | -7.994 | -66.314 | -9.3599 | -73.8687 | -1.064 | 45.483 | 2023/06/23 12:43:36
35 |  | UNDE4-3040-3 | -8.327 | -67.141 | -9.7075 | -74.7774 | -1.026 | 39.143 | 2023/06/23 12:52:32
35 |  | UNDE4-3040-3 | -8.347 | -67.432 | -9.7284 | -75.0971 | -1.007 | 39.581 | 2023/06/23 13:01:29
35 |  | UNDE4-3040-3 | -8.376 | -67.437 | -9.7587 | -75.1026 | -1.132 | 39.765 | 2023/06/23 13:10:25
Sample 35 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
35 |  | UNDE4-3040-3 | -9.6386 | -74.7115 |  |  |  | 0.17913 | 0.52973
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
36 |  | UNDE4-E-4 | -2.897 | -60.253 | -4.0383 | -67.2094 | -1.503 | 43.023 | 2023/06/23 13:20:09
36 |  | UNDE4-E-4 | -2.605 | -59.111 | -3.7335 | -65.9546 | -1.004 | 37.472 | 2023/06/23 13:29:06
36 |  | UNDE4-E-4 | -2.611 | -59.229 | -3.7397 | -66.0843 | -1.163 | 39.348 | 2023/06/23 13:38:01
36 |  | UNDE4-E-4 | -2.596 | -59.064 | -3.7241 | -65.903 | -1.256 | 38.941 | 2023/06/23 13:46:57
Sample 36 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
36 |  | UNDE4-E-4 | -3.8089 | -66.2878 |  |  |  | 0.14663 | 0.56346
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
37 |  | UNDE4-D-2 | 1.71 | -50.196 | 0.77162 | -56.1595 | -1.317 | 40.018 | 2023/06/23 13:56:40
37 |  | UNDE4-D-2 | 2.006 | -49.055 | 1.0807 | -54.9058 | -1.124 | 40.714 | 2023/06/23 14:05:36
37 |  | UNDE4-D-2 | 2.012 | -48.95 | 1.0869 | -54.7905 | -1.078 | 40.467 | 2023/06/23 14:14:31
37 |  | UNDE4-D-2 | 2.068 | -48.549 | 1.1454 | -54.3499 | -1.346 | 40.924 | 2023/06/23 14:23:26
Sample 37 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
37 |  | UNDE4-D-2 | 1.0212 | -55.0514 |  |  |  | 0.16176 | 0.7068
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
38 |  | UNDE4-E-3 | -4.359 | -60.585 | -5.5647 | -67.5741 | -2.183 | 42.313 | 2023/06/23 14:33:11
38 |  | UNDE4-E-3 | -4.858 | -62.235 | -6.0857 | -69.387 | -2.294 | 43.833 | 2023/06/23 14:42:07
38 |  | UNDE4-E-3 | -4.932 | -62.581 | -6.163 | -69.7672 | -2.225 | 43.587 | 2023/06/23 14:51:03
38 |  | UNDE4-E-3 | -5.015 | -62.961 | -6.2496 | -70.1847 | -2.227 | 43.564 | 2023/06/23 14:59:59
Sample 38 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
38 |  | UNDE4-E-3 | -6.0158 | -69.2283 | -2.2323 | 1.8223 |  | 0.29505 | 1.0465
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
39 |  | UNDE4-1020-3 | -2.888 | -45.261 | -4.0289 | -50.7373 | -1.078 | 40.047 | 2023/06/23 15:09:44
39 |  | UNDE4-1020-3 | -2.655 | -43.142 | -3.7857 | -48.4091 | -1.052 | 39.877 | 2023/06/23 15:18:39
39 |  | UNDE4-1020-3 | -2.668 | -42.757 | -3.7992 | -47.9861 | -1.131 | 40.038 | 2023/06/23 15:27:35
39 |  | UNDE4-1020-3 | -2.591 | -42.407 | -3.7189 | -47.6015 | -1.104 | 39.924 | 2023/06/23 15:36:32
Sample 39 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
39 |  | UNDE4-1020-3 | -3.8332 | -48.6835 |  |  |  | 0.12945 | 1.2818
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
40 |  | EV | -8.458 | -63.493 | -9.8443 | -70.7692 | -0.963 | 42.131 | 2023/06/23 15:46:17
40 |  | EV | -8.764 | -65.418 | -10.1638 | -72.8843 | -0.858 | 40.384 | 2023/06/23 15:55:12
40 |  | EV | -8.81 | -65.924 | -10.2118 | -73.4402 | -1.077 | 40.032 | 2023/06/23 16:04:08
40 |  | EV | -8.844 | -66.01 | -10.2473 | -73.5347 | -0.999 | 39.863 | 2023/06/23 16:13:04
Sample 40 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
40 |  | EV | -10.1168 | -72.6571 |  |  |  | 0.17706 | 1.1749
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
41 |  | UNDE4-A-5 | -4.685 | -67.83 | -5.9051 | -75.5344 | -1.485 | 53.102 | 2023/06/23 16:22:51
41 |  | UNDE4-A-5 | -4.503 | -67.681 | -5.7151 | -75.3707 | -1.083 | 40.452 | 2023/06/23 16:31:47
41 |  | UNDE4-A-5 | -4.482 | -67.576 | -5.6932 | -75.2553 | -1.09 | 40.264 | 2023/06/23 16:40:42
41 |  | UNDE4-A-5 | -4.431 | -67.585 | -5.6399 | -75.2652 | -1.095 | 39.607 | 2023/06/23 16:49:36
Sample 41 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
41 |  | UNDE4-A-5 | -5.7383 | -75.3564 |  | 1.8543 |  | 0.11071 | 0.11799
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
42 |  | UNDE4-B-3 | 5.069 | -41.145 | 4.2786 | -46.2149 | -1.304 | 46.187 | 2023/06/23 16:59:22
42 |  | UNDE4-B-3 | 5.463 | -38.515 | 4.69 | -43.3253 | -1.25 | 40.663 | 2023/06/23 17:08:18
42 |  | UNDE4-B-3 | 5.565 | -37.897 | 4.7964 | -42.6462 | -1.245 | 40.864 | 2023/06/23 17:17:14
42 |  | UNDE4-B-3 | 5.638 | -37.541 | 4.8727 | -42.2551 | -1.194 | 41.193 | 2023/06/23 17:26:10
Sample 42 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
42 |  | UNDE4-B-3 | 4.6594 | -43.6104 |  |  |  | 0.25354 | 1.6308
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
43 |  | UNDE4-3040-2 | -7.693 | -63.942 | -9.0456 | -71.2626 | -1.006 | 44.931 | 2023/06/23 17:35:54
43 |  | UNDE4-3040-2 | -8.252 | -66.324 | -9.6292 | -73.8797 | -0.955 | 40.176 | 2023/06/23 17:44:50
43 |  | UNDE4-3040-2 | -8.414 | -66.937 | -9.7984 | -74.5532 | -1.019 | 39.296 | 2023/06/23 17:53:46
43 |  | UNDE4-3040-2 | -8.398 | -66.719 | -9.7817 | -74.3137 | -0.892 | 39.32 | 2023/06/23 18:02:41
Sample 43 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
43 |  | UNDE4-3040-2 | -9.5637 | -73.5023 |  |  |  | 0.33877 | 1.3825
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
44 |  | UNDE4-A-1 | -3.683 | -57.614 | -4.859 | -64.3098 | -0.985 | 41.526 | 2023/06/23 18:12:25
44 |  | UNDE4-A-1 | -3.333 | -56.503 | -4.4935 | -63.0891 | -1.102 | 39.93 | 2023/06/23 18:21:21
44 |  | UNDE4-A-1 | -3.247 | -55.961 | -4.4038 | -62.4936 | -0.874 | 39.356 | 2023/06/23 18:30:18
44 |  | UNDE4-A-1 | -3.215 | -55.988 | -4.3703 | -62.5233 | -1.195 | 39.598 | 2023/06/23 18:39:13
Sample 44 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
44 |  | UNDE4-A-1 | -4.5316 | -63.104 |  |  |  | 0.21486 | 0.773
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
45 |  | UNDE4-E-1 | 1.149 | -48.454 | 0.18591 | -54.2455 | -0.987 | 40.61 | 2023/06/23 18:48:58
45 |  | UNDE4-E-1 | 1.335 | -47.705 | 0.3801 | -53.4225 | -1.039 | 39.921 | 2023/06/23 18:57:54
45 |  | UNDE4-E-1 | 1.432 | -47.438 | 0.48138 | -53.1292 | -1.2 | 40.183 | 2023/06/23 19:06:49
45 |  | UNDE4-E-1 | 1.465 | -47.428 | 0.51583 | -53.1182 | -1.101 | 39.941 | 2023/06/23 19:15:45
Sample 45 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
45 |  | UNDE4-E-1 | 0.3908 | -53.4789 |  |  |  | 0.14199 | 0.48253
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
46 |  | UNDE4-2030-3 | -5.478 | -56.1 | -6.733 | -62.6464 | -0.959 | 39.746 | 2023/06/23 19:25:31
46 |  | UNDE4-2030-3 | -5.785 | -56.863 | -7.0536 | -63.4847 | -0.958 | 39.67 | 2023/06/23 19:34:26
46 |  | UNDE4-2030-3 | -5.836 | -57.09 | -7.1068 | -63.7341 | -1.168 | 39.705 | 2023/06/23 19:43:21
46 |  | UNDE4-2030-3 | -5.904 | -57.088 | -7.1778 | -63.7319 | -0.804 | 39.438 | 2023/06/23 19:52:16
Sample 46 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
46 |  | UNDE4-2030-3 | -7.0178 | -63.3993 |  |  |  | 0.18825 | 0.46909
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
47 |  | UNDE4-deep-1 | -8.477 | -65.688 | -9.8642 | -73.1809 | -0.86 | 39.942 | 2023/06/23 20:02:01
47 |  | UNDE4-deep-1 | -8.585 | -66.606 | -9.9769 | -74.1896 | -1.077 | 40.147 | 2023/06/23 20:10:57
47 |  | UNDE4-deep-1 | -8.605 | -66.665 | -9.9978 | -74.2544 | -0.989 | 39.625 | 2023/06/23 20:19:53
47 |  | UNDE4-deep-1 | -8.606 | -66.869 | -9.9988 | -74.4785 | -0.956 | 39.587 | 2023/06/23 20:28:49
Sample 47 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
47 |  | UNDE4-deep-1 | -9.9594 | -74.0259 |  |  |  | 0.061597 | 0.5249
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
48 |  | WOOD6-D-3 | -6.273 | -67.278 | -7.5631 | -74.9279 | -1.273 | 49.761 | 2023/06/23 20:38:35
48 |  | WOOD6-D-3 | -6.125 | -67.429 | -7.4085 | -75.0938 | -1.383 | 48.482 | 2023/06/23 20:47:31
48 |  | WOOD6-D-3 | -6.149 | -67.539 | -7.4336 | -75.2147 | -1.308 | 48.06 | 2023/06/23 20:56:26
48 |  | WOOD6-D-3 | -6.129 | -67.65 | -7.4127 | -75.3366 | -1.254 | 48.149 | 2023/06/23 21:05:22
Sample 48 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
48 |  | WOOD6-D-3 | -7.4545 | -75.1433 |  | 7.1111 |  | 0.070124 | 0.15879 |  | -1.7194
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
49 |  | EV | -8.77 | -66.871 | -10.1701 | -74.4807 | -0.891 | 42.607 | 2023/06/23 21:15:08
49 |  | EV | -8.836 | -66.727 | -10.239 | -74.3225 | -0.868 | 39.793 | 2023/06/23 21:24:04
49 |  | EV | -8.893 | -66.816 | -10.2985 | -74.4203 | -0.883 | 39.839 | 2023/06/23 21:33:00
49 |  | EV | -8.894 | -66.714 | -10.2995 | -74.3082 | -0.943 | 39.464 | 2023/06/23 21:41:56
Sample 49 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
49 |  | EV | -10.2518 | -74.3829 |  |  |  | 0.05879 | 0.074668
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
50 |  | WOOD6-D-1 | -7.262 | -73.104 | -8.5956 | -81.3291 | -1.239 | 48.685 | 2023/06/23 21:51:40
50 |  | WOOD6-D-1 | -7.155 | -73.558 | -8.4839 | -81.8279 | -0.984 | 45.299 | 2023/06/23 22:00:35
50 |  | WOOD6-D-1 | -7.182 | -73.828 | -8.5121 | -82.1246 | -0.997 | 44.955 | 2023/06/23 22:09:30
50 |  | WOOD6-D-1 | -7.189 | -73.971 | -8.5194 | -82.2817 | -1.135 | 44.901 | 2023/06/23 22:18:26
Sample 50 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
50 |  | WOOD6-D-1 | -8.5278 | -81.8908 |  | 4.4581 |  | 0.045746 | 0.38143
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
51 |  | WOOD6-C-4 | -3.695 | -51.049 | -4.8715 | -57.0967 | -1.6 | 80.334 | 2023/06/23 22:28:12
51 |  | WOOD6-C-4 | -3.508 | -48.887 | -4.6762 | -54.7212 | -1.484 | 64.152 | 2023/06/23 22:37:07
51 |  | WOOD6-C-4 | -3.451 | -48.211 | -4.6167 | -53.9785 | -1.208 | 62.189 | 2023/06/23 22:46:02
51 |  | WOOD6-C-4 | -3.425 | -47.954 | -4.5896 | -53.6961 | -1.252 | 62.216 | 2023/06/23 22:54:57
Sample 51 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
51 |  | WOOD6-C-4 | -4.6885 | -54.8731 | -1.386 | 25.7208 |  | 0.12187 | 1.4054 |  | -4.2019
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
52 |  | WOOD6-C-2 | -5.377 | -50.83 | -6.6276 | -56.8561 | -1.508 | 417.821 | 2023/06/23 23:04:43
52 |  | WOOD6-C-2 | -5.468 | -51.212 | -6.7226 | -57.2758 | -1.709 | 407.499 | 2023/06/23 23:13:37
52 |  | WOOD6-C-2 | -5.492 | -51.14 | -6.7477 | -57.1967 | -1.652 | 364.269 | 2023/06/23 23:22:32
52 |  | WOOD6-C-2 | -5.494 | -51.012 | -6.7497 | -57.056 | -1.498 | 344.139 | 2023/06/23 23:31:28
Sample 52 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
52 |  | WOOD6-C-2 | -6.7119 | -57.0961 | -1.5917 | 341.9301 |  | 0.055114 | 0.16751 |  | -15.2257
Sample |  | Name | Measured d(18_16) Mean | Measured d(D_H) Mean | Calibrated d(18_16) Mean | Calibrated d(D_H) Mean | CH<sub>4</sub> | Other | Time Code
53 |  | EV | -8.628 | -64.622 | -10.0218 | -72.0097 | -1.154 | 48.555 | 2023/06/23 23:41:15
53 |  | EV | -8.686 | -65.48 | -10.0824 | -72.9524 | -0.978 | 41.556 | 2023/06/23 23:50:11
53 |  | EV | -8.767 | -66.068 | -10.1669 | -73.5985 | -1.061 | 40.803 | 2023/06/23 23:59:06
53 |  | EV | -8.812 | -66.08 | -10.2139 | -73.6116 | -0.84 | 39.982 | 2023/06/24 00:08:01
Sample 53 Summary
Sample | Data Error | Name | Calibrated d<sup>18</sup>O Mean | Calibrated d<sup>2</sup>H Mean | CH<sub>4</sub> | C<sub>2+</sub> alcohols | relative %deviation | Uncalibrated d<sup>18</sup>O precision | Uncalibrated d<sup>2</sup>H precision | slope | curvature
53 |  | EV | -10.1212 | -73.043 |  | 1.2221 |  | 0.08216 | 0.6867

## Source
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | Timestamp Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | baseline_shift | slope_shift | residuals | baseline_curvature | interval | ch4_ppm | h16od_adjust | h16od_shift | n2_flag | DAS Temp | Tray | Sample | Job | Method | Error Code
1 | P-24803 | 2023/06/22 12:23:53 | 1-01 | 1 | -9.755 | -73.851 | 15121 | -1 | 1 | EV | standard | H2O | 1687458539.1 | 0.218 | 0.709 | 85.203 | 0.001 | 0.002 | 2.067 | 111.837 | 34.293 | 2.547 | 72.062 | 0.857 | -1.148 | 0 | 0 | 1 | 44 | 1 | 1 | 1 | GJB_Rinse | 0
2 | P-24803 | 2023/06/22 12:32:51 | 1-01 | 2 | -9.058 | -68.957 | 12900 | -1 | 1 | EV | standard | H2O | 1687459075.68 | 0.22 | 0.728 | 74.875 | 0.001 | -0.001 | 1.849 | 55.478 | 29.983 | 2.197 | 64.164 | 0.857 | -1.195 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
3 | P-24803 | 2023/06/22 12:41:47 | 1-01 | 3 | -8.837 | -67.42 | 12993 | -1 | 1 | EV | standard | H2O | 1687459612.95 | 0.211 | 0.81 | 76.456 | 0.001 | 0.001 | 1.886 | 47.186 | 30.621 | 2.201 | 64.619 | 0.856 | -0.992 | 0 | 0 | 1 | 42.875 | 1 | 1 | 1 | GJB_Rinse | 0
4 | P-24803 | 2023/06/22 12:50:44 | 1-01 | 4 | -8.857 | -67.358 | 12950 | 0 | 1 | EV | standard | H2O | 1687460149.89 | 0.184 | 0.84 | 74.328 | 0.002 | 0 | 1.839 | 47.063 | 30.63 | 2.192 | 63.386 | 0.857 | -1.032 | 0 | 0 | 1 | 42.75 | 1 | 1 | 1 | GJB_Rinse | 0
5 | P-24803 | 2023/06/22 12:59:41 | 1-01 | 5 | -8.846 | -67.007 | 13099 | 0 | 1 | EV | standard | H2O | 1687460686.52 | 0.218 | 0.796 | 66.124 | 0.001 | 0 | 1.623 | 46.843 | 30.988 | 2.214 | 64.464 | 0.857 | -1.078 | 0 | 0 | 1 | 42.812 | 1 | 1 | 1 | GJB_Rinse | 0
6 | P-24803 | 2023/06/22 13:08:38 | 1-01 | 6 | -8.895 | -67.268 | 13111 | 0 | 1 | EV | standard | H2O | 1687461223.28 | 0.197 | 0.756 | 70.773 | 0.001 | 0.001 | 1.74 | 46.513 | 30.463 | 2.232 | 65.887 | 0.856 | -0.97 | 0 | 0 | 1 | 42.938 | 1 | 1 | 1 | GJB_Rinse | 0
7 | P-24803 | 2023/06/22 13:17:34 | 1-01 | 7 | -8.81 | -66.746 | 13175 | 0 | 1 | EV | standard | H2O | 1687461760.14 | 0.182 | 0.731 | 69.317 | 0 | 0.002 | 1.7 | 45.494 | 31.008 | 2.236 | 65.141 | 0.856 | -1.039 | 0 | 0 | 1 | 43.188 | 1 | 1 | 1 | GJB_Rinse | 0
8 | P-24803 | 2023/06/22 13:26:30 | 1-01 | 8 | -8.822 | -66.708 | 13019 | 0 | 1 | EV | standard | H2O | 1687462296.18 | 0.213 | 0.846 | 71.31 | 0.001 | -0.002 | 1.753 | 44.315 | 30.518 | 2.209 | 65.168 | 0.857 | -0.889 | 0 | 0 | 1 | 43.312 | 1 | 1 | 1 | GJB_Rinse | 0
9 | P-24803 | 2023/06/22 13:35:27 | 1-01 | 9 | -8.837 | -66.86 | 13005 | 0 | 1 | EV | standard | H2O | 1687462832.28 | 0.216 | 0.849 | 71.337 | 0.001 | 0.003 | 1.767 | 43.768 | 30.581 | 2.195 | 64.205 | 0.856 | -1.106 | 0 | 0 | 1 | 43.438 | 1 | 1 | 1 | GJB_Rinse | 0
10 | P-24803 | 2023/06/22 13:44:24 | 1-01 | 10 | -8.784 | -66.756 | 13053 | 0 | 1 | EV | standard | H2O | 1687463369.5 | 0.209 | 0.846 | 67.583 | 0.001 | -0.002 | 1.66 | 43.588 | 30.836 | 2.203 | 63.973 | 0.856 | -1.186 | 0 | 0 | 1 | 43.562 | 1 | 1 | 1 | GJB_Rinse | 0
11 | P-24804 | 2023/06/22 13:54:13 | 1-02 | 1 | 2.133 | 7.583 | 13185 | -1 | 1 | PZ | standard | H2O | 1687463958.46 | 0.201 | 0.706 | 66.634 | 0 | 0.002 | 1.639 | 42.144 | 31.57 | 2.359 | 67.948 | 0.856 | -0.981 | 0 | 0 | 1 | 43.562 | 1 | 2 | 1 | GJB_Rinse | 0
12 | P-24804 | 2023/06/22 14:03:10 | 1-02 | 2 | 2.673 | 15.076 | 12875 | -1 | 1 | PZ | standard | H2O | 1687464495.68 | 0.19 | 0.801 | 67.692 | 0 | -0.003 | 1.665 | 41.883 | 30.852 | 2.316 | 65.37 | 0.856 | -1.134 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
13 | P-24804 | 2023/06/22 14:12:07 | 1-02 | 3 | 2.829 | 17.295 | 12909 | -1 | 1 | PZ | standard | H2O | 1687465032.54 | 0.196 | 0.857 | 68.073 | 0.001 | -0.006 | 1.675 | 41.856 | 30.793 | 2.327 | 66.304 | 0.856 | -1.16 | 0 | 0 | 1 | 43.188 | 1 | 2 | 1 | GJB_Rinse | 0
14 | P-24804 | 2023/06/22 14:21:04 | 1-02 | 4 | 2.861 | 18.004 | 12817 | 0 | 1 | PZ | standard | H2O | 1687465569.38 | 0.199 | 0.773 | 70.852 | 0 | -0.001 | 1.742 | 41.65 | 30.317 | 2.302 | 65.947 | 0.856 | -1.12 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
15 | P-24804 | 2023/06/22 14:30:00 | 1-02 | 5 | 2.884 | 18.459 | 12975 | 0 | 1 | PZ | standard | H2O | 1687466105.45 | 0.207 | 0.818 | 69.41 | 0.001 | -0.003 | 1.715 | 42.096 | 30.803 | 2.338 | 66.619 | 0.857 | -1.14 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
16 | P-24804 | 2023/06/22 14:38:57 | 1-02 | 6 | 2.925 | 19.137 | 12880 | 0 | 1 | PZ | standard | H2O | 1687466642.24 | 0.197 | 0.832 | 73.463 | 0.001 | -0.004 | 1.804 | 41.526 | 30.461 | 2.334 | 66.623 | 0.856 | -1.142 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
17 | P-24804 | 2023/06/22 14:47:54 | 1-02 | 7 | 2.953 | 19.25 | 12877 | 0 | 1 | PZ | standard | H2O | 1687467179.62 | 0.196 | 0.915 | 70.137 | 0.001 | -0.003 | 1.737 | 41.695 | 30.315 | 2.326 | 66.394 | 0.856 | -1.071 | 0 | 0 | 1 | 43.25 | 1 | 2 | 1 | GJB_Rinse | 0
18 | P-24804 | 2023/06/22 14:56:50 | 1-02 | 8 | 2.94 | 19.371 | 12848 | 0 | 1 | PZ | standard | H2O | 1687467716.12 | 0.213 | 0.729 | 72.477 | 0 | -0.001 | 1.793 | 41.594 | 30.37 | 2.324 | 66.775 | 0.856 | -1.005 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
19 | P-24804 | 2023/06/22 15:05:47 | 1-02 | 9 | 3.011 | 19.783 | 12987 | 0 | 1 | PZ | standard | H2O | 1687468252.67 | 0.209 | 0.862 | 68.336 | 0.001 | -0.004 | 1.674 | 41.68 | 30.555 | 2.345 | 68.148 | 0.857 | -0.915 | 0 | 0 | 1 | 43.312 | 1 | 2 | 1 | GJB_Rinse | 0
20 | P-24804 | 2023/06/22 15:14:44 | 1-02 | 10 | 2.986 | 19.951 | 12768 | 0 | 1 | PZ | standard | H2O | 1687468790.03 | 0.181 | 0.957 | 64.762 | 0.001 | -0.005 | 1.591 | 41.289 | 30.124 | 2.302 | 65.874 | 0.857 | -1.111 | 0 | 0 | 1 | 43.375 | 1 | 2 | 1 | GJB_Rinse | 0
21 | P-24805 | 2023/06/22 15:24:33 | 1-03 | 1 | -13.221 | -93.149 | 13143 | -1 | 1 | UT2 | standard | H2O | 1687469378.98 | 0.192 | 0.777 | 69.665 | 0.002 | 0 | 1.722 | 41.734 | 31.43 | 2.179 | 63.967 | 0.856 | -1.065 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
22 | P-24805 | 2023/06/22 15:33:30 | 1-03 | 2 | -13.964 | -104.054 | 13005 | -1 | 1 | UT2 | standard | H2O | 1687469915.64 | 0.18 | 0.705 | 70.673 | 0.002 | 0.006 | 1.735 | 41.615 | 30.977 | 2.147 | 62.721 | 0.857 | -1.056 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
23 | P-24805 | 2023/06/22 15:42:26 | 1-03 | 3 | -14.202 | -107.311 | 12768 | -1 | 1 | UT2 | standard | H2O | 1687470452.1 | 0.209 | 0.796 | 65.175 | 0.002 | 0.005 | 1.61 | 41.058 | 30.386 | 2.106 | 61.556 | 0.856 | -1.075 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
24 | P-24805 | 2023/06/22 15:51:24 | 1-03 | 4 | -14.22 | -108.263 | 12809 | 0 | 1 | UT2 | standard | H2O | 1687470989.37 | 0.219 | 0.797 | 67.736 | 0.002 | 0.004 | 1.669 | 41.273 | 30.001 | 2.115 | 63.043 | 0.856 | -0.915 | 0 | 0 | 1 | 43.312 | 1 | 3 | 1 | GJB_Rinse | 0
25 | P-24805 | 2023/06/22 16:00:20 | 1-03 | 5 | -14.277 | -109.013 | 12783 | 0 | 1 | UT2 | standard | H2O | 1687471525.44 | 0.189 | 0.77 | 72.096 | 0.001 | 0.004 | 1.772 | 41.179 | 29.859 | 2.109 | 63.154 | 0.856 | -0.867 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
26 | P-24805 | 2023/06/22 16:09:17 | 1-03 | 6 | -14.394 | -109.935 | 12691 | 0 | 1 | UT2 | standard | H2O | 1687472061.89 | 0.21 | 0.732 | 63.427 | 0.001 | 0.004 | 1.565 | 41.062 | 29.699 | 2.088 | 62.342 | 0.856 | -0.978 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
27 | P-24805 | 2023/06/22 16:18:13 | 1-03 | 7 | -14.404 | -110.138 | 12811 | 0 | 1 | UT2 | standard | H2O | 1687472598.38 | 0.186 | 0.785 | 65.445 | 0.001 | 0.006 | 1.608 | 41.091 | 29.878 | 2.121 | 63.531 | 0.856 | -0.756 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
28 | P-24805 | 2023/06/22 16:27:10 | 1-03 | 8 | -14.449 | -110.462 | 12813 | 0 | 1 | UT2 | standard | H2O | 1687473135.66 | 0.184 | 0.788 | 67.203 | 0.001 | 0.004 | 1.662 | 41.099 | 30.069 | 2.111 | 63.348 | 0.856 | -0.913 | 0 | 0 | 1 | 43.5 | 1 | 3 | 1 | GJB_Rinse | 0
29 | P-24805 | 2023/06/22 16:36:05 | 1-03 | 9 | -14.456 | -110.687 | 12848 | 0 | 1 | UT2 | standard | H2O | 1687473670.71 | 0.208 | 0.822 | 69.511 | 0.001 | 0.005 | 1.719 | 41.301 | 30.11 | 2.099 | 62.974 | 0.856 | -0.938 | 0 | 0 | 1 | 43.375 | 1 | 3 | 1 | GJB_Rinse | 0
30 | P-24805 | 2023/06/22 16:45:02 | 1-03 | 10 | -14.42 | -110.888 | 12843 | 0 | 1 | UT2 | standard | H2O | 1687474207.91 | 0.206 | 0.799 | 68.015 | 0.001 | 0.005 | 1.681 | 41.213 | 30.213 | 2.117 | 63.171 | 0.856 | -0.77 | 0 | 0 | 1 | 43.438 | 1 | 3 | 1 | GJB_Rinse | 0
31 | P-24806 | 2023/06/22 16:54:52 | 1-04 | 1 | -9.351 | -72.894 | 12887 | -1 | 1 | EV | standard | H2O | 1687474797.6 | 0.197 | 0.813 | 69.64 | 0.001 | 0.005 | 1.71 | 41.13 | 30.233 | 2.189 | 65.112 | 0.856 | -0.903 | 0 | 0 | 1 | 43.375 | 1 | 4 | 1 | GJB_Rinse | 0
32 | P-24806 | 2023/06/22 17:03:49 | 1-04 | 2 | -9.085 | -69.011 | 12902 | -1 | 1 | EV | standard | H2O | 1687475334.22 | 0.217 | 0.73 | 68.233 | 0.001 | 0.001 | 1.679 | 41.174 | 30.203 | 2.184 | 65.104 | 0.856 | -0.878 | 0 | 0 | 1 | 43.5 | 1 | 4 | 1 | GJB_Rinse | 0
33 | P-24806 | 2023/06/22 17:12:46 | 1-04 | 3 | -9.023 | -67.995 | 12736 | -1 | 1 | EV | standard | H2O | 1687475870.94 | 0.182 | 0.777 | 64.888 | 0.001 | -0.001 | 1.601 | 41.058 | 29.565 | 2.163 | 63.859 | 0.856 | -0.955 | 0 | 0 | 1 | 43.562 | 1 | 4 | 1 | GJB_Rinse | 0
34 | P-24806 | 2023/06/22 17:21:42 | 1-04 | 4 | -8.944 | -67.433 | 12884 | 0 | 1 | EV | standard | H2O | 1687476408.01 | 0.196 | 0.821 | 65.248 | 0.001 | -0.001 | 1.615 | 41.286 | 30.083 | 2.179 | 65.043 | 0.856 | -0.986 | 0 | 0 | 1 | 43.625 | 1 | 4 | 1 | GJB_Rinse | 0
35 | P-24806 | 2023/06/22 17:30:39 | 1-04 | 5 | -9.004 | -67.396 | 12923 | 0 | 1 | EV | standard | H2O | 1687476944.13 | 0.182 | 0.702 | 63.959 | 0.001 | 0 | 1.572 | 41.311 | 30.197 | 2.203 | 65.425 | 0.856 | -0.881 | 0 | 0 | 1 | 43.625 | 1 | 4 | 1 | GJB_Rinse | 0
36 | P-24806 | 2023/06/22 17:39:35 | 1-04 | 6 | -8.992 | -67.375 | 12694 | 0 | 1 | EV | standard | H2O | 1687477481.25 | 0.177 | 0.824 | 65.984 | 0 | 0.002 | 1.621 | 40.715 | 29.815 | 2.159 | 64.221 | 0.857 | -0.935 | 0 | 0 | 1 | 43.688 | 1 | 4 | 1 | GJB_Rinse | 0
37 | P-24806 | 2023/06/22 17:48:32 | 1-04 | 7 | -8.997 | -67.196 | 12804 | 0 | 1 | EV | standard | H2O | 1687478017.99 | 0.184 | 0.776 | 64.593 | 0.001 | 0.002 | 1.581 | 41.004 | 30.139 | 2.17 | 64.15 | 0.856 | -1.009 | 0 | 0 | 1 | 43.688 | 1 | 4 | 1 | GJB_Rinse | 0
38 | P-24806 | 2023/06/22 17:57:29 | 1-04 | 8 | -8.96 | -67.127 | 12826 | 0 | 1 | EV | standard | H2O | 1687478555.14 | 0.206 | 0.73 | 66.436 | 0.001 | 0.005 | 1.641 | 41.17 | 30.192 | 2.182 | 63.659 | 0.856 | -1.115 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
39 | P-24806 | 2023/06/22 18:06:27 | 1-04 | 9 | -8.975 | -67.245 | 12859 | 0 | 1 | EV | standard | H2O | 1687479092.59 | 0.188 | 0.85 | 64.921 | 0.001 | -0.001 | 1.606 | 41.234 | 30.106 | 2.174 | 64.08 | 0.856 | -1.04 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
40 | P-24806 | 2023/06/22 18:15:24 | 1-04 | 10 | -8.942 | -66.994 | 12983 | 0 | 1 | EV | standard | H2O | 1687479629.18 | 0.201 | 0.764 | 64.691 | 0.001 | 0.003 | 1.597 | 41.4 | 30.681 | 2.205 | 64.441 | 0.856 | -0.971 | 0 | 0 | 1 | 43.75 | 1 | 4 | 1 | GJB_Rinse | 0
41 | P-24807 | 2023/06/22 18:25:14 | 1-05 | 1 | -3.409 | -60.25 | 13011 | -1 | 1 | UNDE4-E-2 | 22-304 | H2O | 1687480219.5 | 0.213 | 0.857 | 65.982 | 0 | 0.004 | 1.627 | 42.449 | 30.107 | 2.254 | 64.738 | 0.856 | -1.392 | 0 | 0 | 1 | 43.688 | 1 | 5 | 1 | GJB_Rinse | 0
42 | P-24807 | 2023/06/22 18:34:10 | 1-05 | 2 | -3.136 | -59.653 | 12820 | -1 | 1 | UNDE4-E-2 | 22-304 | H2O | 1687480755.51 | 0.197 | 0.752 | 69.724 | 0 | 0.003 | 1.724 | 42.041 | 29.436 | 2.237 | 64.712 | 0.856 | -1.213 | 0 | 0 | 1 | 43.75 | 1 | 5 | 1 | GJB_Rinse | 0
43 | P-24807 | 2023/06/22 18:43:06 | 1-05 | 3 | -3.06 | -59.229 | 12809 | -1 | 1 | UNDE4-E-2 | 22-304 | H2O | 1687481291.79 | 0.188 | 0.821 | 67.633 | 0.001 | 0.001 | 1.672 | 42.002 | 29.461 | 2.234 | 65.415 | 0.856 | -1.266 | 0 | 0 | 1 | 43.625 | 1 | 5 | 1 | GJB_Rinse | 0
44 | P-24807 | 2023/06/22 18:52:03 | 1-05 | 4 | -3.049 | -59.214 | 12872 | 0 | 1 | UNDE4-E-2 | 22-304 | H2O | 1687481828.92 | 0.178 | 0.762 | 64.314 | 0 | -0.001 | 1.575 | 42.081 | 29.731 | 2.256 | 64.846 | 0.856 | -1.287 | 0 | 0 | 1 | 43.688 | 1 | 5 | 1 | GJB_Rinse | 0
45 | P-24808 | 2023/06/22 19:01:54 | 1-06 | 1 | -5.718 | -52.055 | 13212 | -1 | 1 | UNDE4-010-1 | 22-304 | H2O | 1687482419.71 | 0.212 | 0.794 | 69.379 | 0.001 | 0 | 1.702 | 41.862 | 31.348 | 2.272 | 66.268 | 0.856 | -1.043 | 0 | 0 | 1 | 43.5 | 1 | 6 | 1 | GJB_Rinse | 0
46 | P-24808 | 2023/06/22 19:10:51 | 1-06 | 2 | -5.891 | -51.381 | 12956 | -1 | 1 | UNDE4-010-1 | 22-304 | H2O | 1687482956.64 | 0.195 | 0.8 | 71.66 | 0.001 | 0.001 | 1.763 | 41.308 | 30.446 | 2.236 | 65.511 | 0.856 | -1.055 | 0 | 0 | 1 | 43.375 | 1 | 6 | 1 | GJB_Rinse | 0
47 | P-24808 | 2023/06/22 19:19:47 | 1-06 | 3 | -5.925 | -51.431 | 13096 | -1 | 1 | UNDE4-010-1 | 22-304 | H2O | 1687483492.48 | 0.194 | 0.751 | 66.817 | 0 | 0.001 | 1.634 | 41.531 | 30.671 | 2.258 | 66.501 | 0.856 | -0.814 | 0 | 0 | 1 | 43.438 | 1 | 6 | 1 | GJB_Rinse | 0
48 | P-24808 | 2023/06/22 19:28:43 | 1-06 | 4 | -5.914 | -51.275 | 12852 | 0 | 1 | UNDE4-010-1 | 22-304 | H2O | 1687484029.35 | 0.192 | 0.8 | 63.807 | 0.001 | -0.001 | 1.568 | 41.094 | 30.36 | 2.205 | 64.601 | 0.856 | -1.015 | 0 | 0 | 1 | 43.5 | 1 | 6 | 1 | GJB_Rinse | 0
49 | P-24809 | 2023/06/22 19:38:34 | 1-07 | 1 | -7.55 | -61.447 | 13407 | -1 | 1 | UNDE4-1020-1 | 22-304 | H2O | 1687484619.49 | 0.194 | 0.744 | 68.715 | 0.001 | 0 | 1.677 | 42.189 | 31.948 | 2.281 | 66.297 | 0.857 | -1.05 | 0 | 0 | 1 | 43.5 | 1 | 7 | 1 | GJB_Rinse | 0
50 | P-24809 | 2023/06/22 19:47:31 | 1-07 | 2 | -7.643 | -62.229 | 12856 | -1 | 1 | UNDE4-1020-1 | 22-304 | H2O | 1687485156.54 | 0.184 | 0.852 | 69.874 | 0.001 | -0.001 | 1.718 | 41.009 | 30.212 | 2.194 | 64.884 | 0.856 | -0.944 | 0 | 0 | 1 | 43.438 | 1 | 7 | 1 | GJB_Rinse | 0
51 | P-24809 | 2023/06/22 19:56:27 | 1-07 | 3 | -7.621 | -62.591 | 13209 | -1 | 1 | UNDE4-1020-1 | 22-304 | H2O | 1687485692.77 | 0.195 | 0.904 | 66.074 | 0 | 0.001 | 1.622 | 41.734 | 31.629 | 2.242 | 65.037 | 0.857 | -1.048 | 0 | 0 | 1 | 43.25 | 1 | 7 | 1 | GJB_Rinse | 0
52 | P-24809 | 2023/06/22 20:05:24 | 1-07 | 4 | -7.629 | -62.405 | 12956 | 0 | 1 | UNDE4-1020-1 | 22-304 | H2O | 1687486229.22 | 0.203 | 0.817 | 69.706 | 0.001 | 0 | 1.712 | 41.42 | 30.155 | 2.203 | 65.131 | 0.856 | -0.983 | 0 | 0 | 1 | 43.188 | 1 | 7 | 1 | GJB_Rinse | 0
53 | P-24810 | 2023/06/22 20:15:11 | 1-08 | 1 | -7.804 | -72.461 | 12994 | -1 | 1 | UNDE4-D-5 | 22-304 | H2O | 1687486816.9 | 0.162 | 0.849 | 65.088 | 0.001 | -0.001 | 1.603 | 43.555 | 29.061 | 2.306 | 66.164 | 0.856 | -1.513 | 0 | 0 | 1 | 43.188 | 1 | 8 | 1 | GJB_Rinse | 0
54 | P-24810 | 2023/06/22 20:24:08 | 1-08 | 2 | -7.786 | -73.45 | 13119 | -1 | 1 | UNDE4-D-5 | 22-304 | H2O | 1687487354.27 | 0.208 | 0.763 | 68.229 | 0 | 0.002 | 1.674 | 43.78 | 29.613 | 2.332 | 66.312 | 0.856 | -1.569 | 0 | 0 | 1 | 43.188 | 1 | 8 | 1 | GJB_Rinse | 0
55 | P-24810 | 2023/06/22 20:33:05 | 1-08 | 3 | -7.814 | -73.663 | 13723 | -1 | 1 | UNDE4-D-5 | 22-304 | H2O | 1687487890.72 | 0.189 | 0.712 | 67.085 | 0 | 0.001 | 1.634 | 45.221 | 31.387 | 2.422 | 67.585 | 0.857 | -1.607 | 0 | 0 | 1 | 43.25 | 1 | 8 | 1 | GJB_Rinse | 0
56 | P-24810 | 2023/06/22 20:42:02 | 1-08 | 4 | -7.827 | -73.865 | 14187 | 0 | 1 | UNDE4-D-5 | 22-304 | H2O | 1687488426.86 | 0.181 | 0.686 | 69.018 | 0.002 | 0.002 | 1.684 | 46.17 | 32.814 | 2.482 | 70.261 | 0.856 | -1.626 | 0 | 0 | 1 | 43.25 | 1 | 8 | 1 | GJB_Rinse | 0
57 | P-24811 | 2023/06/22 20:51:50 | 1-09 | 1 | -4.792 | -69.431 | 12515 | -1 | 1 | UNDE4-A-4 | 22-304 | H2O | 1687489015.13 | 0.223 | 0.991 | 65.754 | 0 | -0.002 | 1.628 | 556.171 | 9.144 | 2.542 | 55.808 | 0.859 | -3.109 | 0 | 0 | 1 | 43.125 | 1 | 9 | 1 | GJB_Rinse | 0
58 | P-24811 | 2023/06/22 21:00:47 | 1-09 | 2 | -4.716 | -69.178 | 12569 | -1 | 1 | UNDE4-A-4 | 22-304 | H2O | 1687489552.37 | 0.272 | 0.982 | 70.092 | 0 | 0 | 1.723 | 561.894 | 9.25 | 2.571 | 56.284 | 0.858 | -3.158 | 0 | 0 | 1 | 43.062 | 1 | 9 | 1 | GJB_Rinse | 0
59 | P-24811 | 2023/06/22 21:09:43 | 1-09 | 3 | -4.708 | -68.811 | 12428 | -1 | 1 | UNDE4-A-4 | 22-304 | H2O | 1687490088.41 | 0.241 | 0.827 | 67.152 | 0.001 | 0.001 | 1.664 | 557.361 | 9.093 | 2.567 | 55.429 | 0.858 | -3.096 | 0 | 0 | 1 | 43.188 | 1 | 9 | 1 | GJB_Rinse | 0
60 | P-24811 | 2023/06/22 21:18:38 | 1-09 | 4 | -4.695 | -69.065 | 12299 | 0 | 1 | UNDE4-A-4 | 22-304 | H2O | 1687490624.22 | 0.271 | 0.94 | 67.389 | 0.001 | -0.002 | 1.673 | 553.686 | 8.692 | 2.531 | 56.327 | 0.858 | -2.909 | 0 | 0 | 1 | 43.25 | 1 | 9 | 1 | GJB_Rinse | 0
61 | P-24812 | 2023/06/22 21:28:27 | 1-10 | 1 | -6.487 | -66.168 | 12797 | -1 | 1 | UNDE4-D-4 | 22-304 | H2O | 1687491212.62 | 0.208 | 0.857 | 67.595 | 0 | 0.006 | 1.67 | 49.509 | 28.646 | 2.283 | 65.385 | 0.856 | -1.426 | 0 | 0 | 1 | 43.25 | 1 | 10 | 1 | GJB_Rinse | 0
62 | P-24812 | 2023/06/22 21:37:24 | 1-10 | 2 | -6.448 | -65.487 | 12804 | -1 | 1 | UNDE4-D-4 | 22-304 | H2O | 1687491749.48 | 0.197 | 0.803 | 65.967 | 0.002 | -0.002 | 1.632 | 46.662 | 28.772 | 2.278 | 64.321 | 0.857 | -1.704 | 0 | 0 | 1 | 43.312 | 1 | 10 | 1 | GJB_Rinse | 0
63 | P-24812 | 2023/06/22 21:46:19 | 1-10 | 3 | -6.455 | -65.264 | 12756 | -1 | 1 | UNDE4-D-4 | 22-304 | H2O | 1687492284.55 | 0.204 | 0.647 | 67.654 | 0.001 | -0.002 | 1.665 | 44.988 | 28.678 | 2.261 | 65.236 | 0.856 | -1.485 | 0 | 0 | 1 | 43.312 | 1 | 10 | 1 | GJB_Rinse | 0
64 | P-24812 | 2023/06/22 21:55:16 | 1-10 | 4 | -6.502 | -65.225 | 12617 | 0 | 1 | UNDE4-D-4 | 22-304 | H2O | 1687492822.49 | 0.188 | 0.788 | 64.564 | 0 | -0.001 | 1.597 | 44.107 | 28.47 | 2.249 | 63.459 | 0.857 | -1.638 | 0 | 0 | 1 | 43.375 | 1 | 10 | 1 | GJB_Rinse | 0
65 | P-24813 | 2023/06/22 22:05:06 | 1-11 | 1 | -5.424 | -66.081 | 12798 | -1 | 1 | UNDE4-C-1 | 22-304 | H2O | 1687493411.63 | 0.225 | 0.904 | 66.039 | 0.001 | 0.002 | 1.631 | 170.908 | 25.469 | 2.24 | 61.707 | 0.857 | -1.342 | 0 | 0 | 1 | 43.375 | 1 | 11 | 1 | GJB_Rinse | 0
66 | P-24813 | 2023/06/22 22:14:03 | 1-11 | 2 | -5.373 | -66.077 | 12339 | -1 | 1 | UNDE4-C-1 | 22-304 | H2O | 1687493948.48 | 0.203 | 0.917 | 64.795 | 0.001 | -0.001 | 1.6 | 166.666 | 24.096 | 2.18 | 60.371 | 0.856 | -1.275 | 0 | 0 | 1 | 43.375 | 1 | 11 | 1 | GJB_Rinse | 0
67 | P-24813 | 2023/06/22 22:22:59 | 1-11 | 3 | -5.37 | -66.162 | 12425 | -1 | 1 | UNDE4-C-1 | 22-304 | H2O | 1687494485.04 | 0.206 | 0.861 | 66.554 | 0.001 | -0.001 | 1.643 | 167.803 | 24.46 | 2.193 | 60.904 | 0.856 | -1.247 | 0 | 0 | 1 | 43.375 | 1 | 11 | 1 | GJB_Rinse | 0
68 | P-24813 | 2023/06/22 22:31:56 | 1-11 | 4 | -5.358 | -65.949 | 12506 | 0 | 1 | UNDE4-C-1 | 22-304 | H2O | 1687495021.59 | 0.202 | 0.887 | 63.461 | 0.001 | 0.002 | 1.578 | 169.074 | 24.531 | 2.191 | 60.765 | 0.856 | -1.489 | 0 | 0 | 1 | 43.438 | 1 | 11 | 1 | GJB_Rinse | 0
69 | P-24814 | 2023/06/22 22:41:45 | 1-12 | 1 | -6.142 | -55.83 | 12481 | -1 | 1 | UNDE4-2030-2 | 22-304 | H2O | 1687495611 | 0.208 | 0.758 | 67.16 | 0.001 | 0.001 | 1.66 | 42.502 | 28.915 | 2.13 | 62.989 | 0.856 | -1.06 | 0 | 0 | 1 | 43.438 | 1 | 12 | 1 | GJB_Rinse | 0
70 | P-24814 | 2023/06/22 22:50:42 | 1-12 | 2 | -6.166 | -54.896 | 12209 | -1 | 1 | UNDE4-2030-2 | 22-304 | H2O | 1687496147.94 | 0.218 | 0.871 | 61.123 | 0.001 | -0.001 | 1.518 | 40.784 | 27.817 | 2.103 | 63.668 | 0.856 | -0.9 | 0 | 0 | 1 | 43.5 | 1 | 12 | 1 | GJB_Rinse | 0
71 | P-24814 | 2023/06/22 22:59:40 | 1-12 | 3 | -6.152 | -54.408 | 12379 | -1 | 1 | UNDE4-2030-2 | 22-304 | H2O | 1687496685.4 | 0.19 | 0.744 | 64.618 | 0.001 | 0 | 1.608 | 40.877 | 28.981 | 2.119 | 62.027 | 0.856 | -1.106 | 0 | 0 | 1 | 43.375 | 1 | 12 | 1 | GJB_Rinse | 0
72 | P-24814 | 2023/06/22 23:08:36 | 1-12 | 4 | -6.117 | -54.053 | 12069 | 0 | 1 | UNDE4-2030-2 | 22-304 | H2O | 1687497222.14 | 0.23 | 0.754 | 65.683 | 0.001 | 0.003 | 1.634 | 40.315 | 27.723 | 2.076 | 61.336 | 0.856 | -0.982 | 0 | 0 | 1 | 43.438 | 1 | 12 | 1 | GJB_Rinse | 0
73 | P-24815 | 2023/06/22 23:18:26 | 1-13 | 1 | -8.735 | -65.085 | 12968 | -1 | 1 | EV | standard | H2O | 1687497811.78 | 0.19 | 0.827 | 67.6 | 0.001 | 0.003 | 1.66 | 41.585 | 30.241 | 2.212 | 65.587 | 0.856 | -0.89 | 0 | 0 | 1 | 43.5 | 1 | 13 | 1 | GJB_Rinse | 0
74 | P-24815 | 2023/06/22 23:27:23 | 1-13 | 2 | -8.853 | -66.291 | 12646 | -1 | 1 | EV | standard | H2O | 1687498349.09 | 0.18 | 0.896 | 62.764 | 0 | 0.001 | 1.55 | 41.101 | 29.289 | 2.154 | 63.83 | 0.856 | -0.893 | 0 | 0 | 1 | 43.5 | 1 | 13 | 1 | GJB_Rinse | 0
75 | P-24815 | 2023/06/22 23:36:20 | 1-13 | 3 | -8.911 | -66.534 | 12523 | -1 | 1 | EV | standard | H2O | 1687498885.75 | 0.206 | 0.886 | 66.54 | 0.001 | 0.005 | 1.645 | 40.778 | 28.932 | 2.134 | 63.902 | 0.856 | -0.947 | 0 | 0 | 1 | 43.5 | 1 | 13 | 1 | GJB_Rinse | 0
76 | P-24815 | 2023/06/22 23:45:17 | 1-13 | 4 | -8.925 | -66.588 | 12505 | 0 | 1 | EV | standard | H2O | 1687499422.68 | 0.212 | 0.725 | 66.624 | 0.001 | 0.002 | 1.647 | 40.588 | 29.188 | 2.132 | 63.167 | 0.855 | -0.83 | 0 | 0 | 1 | 43.562 | 1 | 13 | 1 | GJB_Rinse | 0
77 | P-24816 | 2023/06/22 23:55:06 | 1-14 | 1 | -4.825 | -45.274 | 12213 | -1 | 1 | UNDE4-010-2 | 22-304 | H2O | 1687500012.2 | 0.193 | 0.788 | 69.362 | 0.001 | 0.003 | 1.716 | 39.956 | 28.299 | 2.115 | 62.144 | 0.856 | -0.889 | 0 | 0 | 1 | 43.562 | 1 | 14 | 1 | GJB_Rinse | 0
78 | P-24816 | 2023/06/23 00:04:02 | 1-14 | 2 | -4.569 | -43.04 | 12597 | -1 | 1 | UNDE4-010-2 | 22-304 | H2O | 1687500546.81 | 0.221 | 0.748 | 62.231 | 0.001 | -0.003 | 1.535 | 40.689 | 29.424 | 2.185 | 64.25 | 0.856 | -0.986 | 0 | 0 | 1 | 43.562 | 1 | 14 | 1 | GJB_Rinse | 0
79 | P-24816 | 2023/06/23 00:12:58 | 1-14 | 3 | -4.504 | -42.213 | 12073 | -1 | 1 | UNDE4-010-2 | 22-304 | H2O | 1687501083.85 | 0.186 | 0.739 | 66.281 | 0.001 | 0.001 | 1.639 | 39.52 | 27.84 | 2.109 | 63.02 | 0.856 | -0.837 | 0 | 0 | 1 | 43.562 | 1 | 14 | 1 | GJB_Rinse | 0
80 | P-24816 | 2023/06/23 00:21:55 | 1-14 | 4 | -4.525 | -42.173 | 12207 | 0 | 1 | UNDE4-010-2 | 22-304 | H2O | 1687501620.16 | 0.2 | 0.857 | 63.354 | 0.001 | -0.002 | 1.577 | 39.898 | 28.3 | 2.117 | 62.906 | 0.856 | -0.886 | 0 | 0 | 1 | 43.562 | 1 | 14 | 1 | GJB_Rinse | 0
81 | P-24817 | 2023/06/23 00:31:41 | 1-15 | 1 | -8.415 | -65.61 | 12772 | -1 | 1 | UNDE4-3040-1 | 22-304 | H2O | 1687502207.08 | 0.191 | 0.799 | 64.972 | 0.001 | 0.002 | 1.605 | 40.946 | 29.98 | 2.165 | 63.959 | 0.856 | -1.055 | 0 | 0 | 1 | 43.562 | 1 | 15 | 1 | GJB_Rinse | 0
82 | P-24817 | 2023/06/23 00:40:38 | 1-15 | 2 | -8.61 | -68.034 | 12816 | -1 | 1 | UNDE4-3040-1 | 22-304 | H2O | 1687502743.66 | 0.183 | 0.754 | 64.412 | 0 | 0.002 | 1.594 | 41.124 | 30.049 | 2.168 | 63.798 | 0.856 | -1.057 | 0 | 0 | 1 | 43.625 | 1 | 15 | 1 | GJB_Rinse | 0
83 | P-24817 | 2023/06/23 00:49:34 | 1-15 | 3 | -8.577 | -68.558 | 12334 | -1 | 1 | UNDE4-3040-1 | 22-304 | H2O | 1687503280.15 | 0.186 | 0.853 | 64.319 | 0 | 0.004 | 1.6 | 40.143 | 28.488 | 2.092 | 62.791 | 0.855 | -1.087 | 0 | 0 | 1 | 43.5 | 1 | 15 | 1 | GJB_Rinse | 0
84 | P-24817 | 2023/06/23 00:58:31 | 1-15 | 4 | -8.63 | -68.832 | 12474 | 0 | 1 | UNDE4-3040-1 | 22-304 | H2O | 1687503816.21 | 0.188 | 0.853 | 60.845 | 0.001 | 0.002 | 1.502 | 40.407 | 29.121 | 2.12 | 62.22 | 0.856 | -1.046 | 0 | 0 | 1 | 43.562 | 1 | 15 | 1 | GJB_Rinse | 0
85 | P-24818 | 2023/06/23 01:08:18 | 1-16 | 1 | -3.885 | -61.094 | 12309 | -1 | 1 | UNDE4-C-4 | 22-304 | H2O | 1687504403.7 | 0.196 | 0.854 | 64.027 | 0.001 | -0.002 | 1.593 | 127.682 | 25.217 | 2.173 | 61.422 | 0.855 | -1.203 | 0 | 0 | 1 | 43.625 | 1 | 16 | 1 | GJB_Rinse | 0
86 | P-24818 | 2023/06/23 01:17:14 | 1-16 | 2 | -3.65 | -60.493 | 12137 | -1 | 1 | UNDE4-C-4 | 22-304 | H2O | 1687504939.54 | 0.196 | 0.792 | 63.311 | 0 | 0.002 | 1.573 | 126.835 | 25.03 | 2.138 | 60.577 | 0.855 | -1.137 | 0 | 0 | 1 | 43.625 | 1 | 16 | 1 | GJB_Rinse | 0
87 | P-24818 | 2023/06/23 01:26:11 | 1-16 | 3 | -3.558 | -60.415 | 12104 | -1 | 1 | UNDE4-C-4 | 22-304 | H2O | 1687505476.34 | 0.177 | 0.868 | 66.8 | 0 | -0.002 | 1.66 | 127.125 | 24.834 | 2.115 | 59.988 | 0.856 | -1.292 | 0 | 0 | 1 | 43.625 | 1 | 16 | 1 | GJB_Rinse | 0
88 | P-24818 | 2023/06/23 01:35:07 | 1-16 | 4 | -3.549 | -60.184 | 12073 | 0 | 1 | UNDE4-C-4 | 22-304 | H2O | 1687506012.42 | 0.204 | 0.81 | 64.913 | 0.001 | -0.001 | 1.604 | 126.584 | 25.094 | 2.126 | 59.771 | 0.856 | -1.368 | 0 | 0 | 1 | 43.688 | 1 | 16 | 1 | GJB_Rinse | 0
89 | P-24819 | 2023/06/23 01:44:54 | 1-17 | 1 | -11.353 | -85.58 | 12604 | -1 | 1 | UNDE4-2030-1 | 22-304 | H2O | 1687506599.49 | 0.18 | 0.867 | 68.139 | 0.001 | 0.003 | 1.685 | 42.035 | 29.581 | 2.103 | 62.247 | 0.856 | -1.122 | 0 | 0 | 1 | 43.75 | 1 | 17 | 1 | GJB_Rinse | 0
90 | P-24819 | 2023/06/23 01:53:50 | 1-17 | 2 | -11.759 | -88.369 | 12053 | -1 | 1 | UNDE4-2030-1 | 22-304 | H2O | 1687507135.71 | 0.198 | 0.778 | 61.41 | 0.001 | 0.004 | 1.522 | 39.903 | 27.854 | 2.039 | 60.875 | 0.856 | -0.885 | 0 | 0 | 1 | 43.625 | 1 | 17 | 1 | GJB_Rinse | 0
91 | P-24819 | 2023/06/23 02:02:46 | 1-17 | 3 | -11.753 | -88.666 | 12400 | -1 | 1 | UNDE4-2030-1 | 22-304 | H2O | 1687507671.86 | 0.191 | 0.759 | 67.463 | 0.001 | 0.004 | 1.665 | 40.548 | 28.636 | 2.088 | 62.842 | 0.855 | -0.982 | 0 | 0 | 1 | 43.562 | 1 | 17 | 1 | GJB_Rinse | 0
92 | P-24819 | 2023/06/23 02:11:43 | 1-17 | 4 | -11.829 | -88.918 | 12039 | 0 | 1 | UNDE4-2030-1 | 22-304 | H2O | 1687508208.29 | 0.191 | 0.787 | 62.273 | 0.001 | 0.002 | 1.563 | 39.695 | 27.892 | 2.033 | 60.125 | 0.856 | -0.957 | 0 | 0 | 1 | 43.438 | 1 | 17 | 1 | GJB_Rinse | 0
93 | P-24820 | 2023/06/23 02:21:31 | 1-18 | 1 | 1.047 | -50.222 | 12831 | -1 | 1 | UNDE4-C-2 | 22-304 | H2O | 1687508797.03 | 0.207 | 0.919 | 70.095 | 0 | 0.003 | 1.732 | 50.542 | 29.338 | 2.234 | 65.25 | 0.856 | -1.09 | 0 | 0 | 1 | 43.5 | 1 | 18 | 1 | GJB_Rinse | 0
94 | P-24820 | 2023/06/23 02:30:28 | 1-18 | 2 | 1.694 | -46.046 | 12372 | -1 | 1 | UNDE4-C-2 | 22-304 | H2O | 1687509333.25 | 0.197 | 0.856 | 65.331 | 0 | 0.001 | 1.627 | 49.593 | 28.294 | 2.161 | 61.325 | 0.856 | -1.378 | 0 | 0 | 1 | 43.625 | 1 | 18 | 1 | GJB_Rinse | 0
95 | P-24820 | 2023/06/23 02:39:24 | 1-18 | 3 | 1.807 | -45.302 | 12348 | -1 | 1 | UNDE4-C-2 | 22-304 | H2O | 1687509869.86 | 0.203 | 0.821 | 68.546 | 0.001 | -0.001 | 1.698 | 49.337 | 27.914 | 2.166 | 63.249 | 0.855 | -1.208 | 0 | 0 | 1 | 43.688 | 1 | 18 | 1 | GJB_Rinse | 0
96 | P-24820 | 2023/06/23 02:48:21 | 1-18 | 4 | 1.87 | -44.877 | 12460 | 0 | 1 | UNDE4-C-2 | 22-304 | H2O | 1687510406.8 | 0.212 | 0.77 | 68.177 | 0.001 | 0.001 | 1.675 | 49.556 | 28.421 | 2.189 | 63.676 | 0.856 | -1.168 | 0 | 0 | 1 | 43.688 | 1 | 18 | 1 | GJB_Rinse | 0
97 | P-24821 | 2023/06/23 02:58:09 | 1-19 | 1 | -3.084 | -41.581 | 12138 | -1 | 1 | UNDE4-010-3 | 22-304 | H2O | 1687510994.79 | 0.192 | 0.858 | 65.226 | 0.001 | -0.004 | 1.611 | 39.88 | 28.241 | 2.1 | 61.811 | 0.856 | -1.103 | 0 | 0 | 1 | 43.75 | 1 | 19 | 1 | GJB_Rinse | 0
98 | P-24821 | 2023/06/23 03:07:05 | 1-19 | 2 | -3.257 | -40.662 | 12123 | -1 | 1 | UNDE4-010-3 | 22-304 | H2O | 1687511531.43 | 0.202 | 0.73 | 63.063 | 0.001 | -0.002 | 1.559 | 39.572 | 28.174 | 2.115 | 62.851 | 0.856 | -0.914 | 0 | 0 | 1 | 43.812 | 1 | 19 | 1 | GJB_Rinse | 0
99 | P-24821 | 2023/06/23 03:16:02 | 1-19 | 3 | -3.304 | -40.495 | 12186 | -1 | 1 | UNDE4-010-3 | 22-304 | H2O | 1687512067.68 | 0.205 | 0.788 | 60.458 | 0.001 | -0.001 | 1.506 | 39.705 | 28.271 | 2.112 | 62.682 | 0.856 | -0.799 | 0 | 0 | 1 | 43.75 | 1 | 19 | 1 | GJB_Rinse | 0
100 | P-24821 | 2023/06/23 03:24:59 | 1-19 | 4 | -3.342 | -40.631 | 12159 | 0 | 1 | UNDE4-010-3 | 22-304 | H2O | 1687512604.48 | 0.216 | 0.852 | 60.965 | 0.001 | -0.003 | 1.519 | 39.656 | 28.295 | 2.119 | 62.057 | 0.855 | -1.023 | 0 | 0 | 1 | 43.75 | 1 | 19 | 1 | GJB_Rinse | 0
101 | P-24822 | 2023/06/23 03:34:48 | 1-20 | 1 | -4.702 | -61.642 | 12653 | -1 | 1 | UNDE4-B-4 | 22-304 | H2O | 1687513193.41 | 0.213 | 0.837 | 61.988 | 0 | 0 | 1.526 | 272.56 | 21.758 | 2.244 | 59.004 | 0.857 | -1.734 | 0 | 0 | 1 | 43.812 | 1 | 20 | 1 | GJB_Rinse | 0
102 | P-24822 | 2023/06/23 03:43:45 | 1-20 | 2 | -4.839 | -63.869 | 12229 | -1 | 1 | UNDE4-B-4 | 22-304 | H2O | 1687513730.37 | 0.21 | 0.803 | 67.866 | 0 | 0.004 | 1.688 | 265.883 | 20.377 | 2.195 | 59.046 | 0.856 | -1.34 | 0 | 0 | 1 | 43.75 | 1 | 20 | 1 | GJB_Rinse | 0
103 | P-24822 | 2023/06/23 03:52:41 | 1-20 | 3 | -4.847 | -64.506 | 12491 | -1 | 1 | UNDE4-B-4 | 22-304 | H2O | 1687514267.28 | 0.229 | 0.788 | 64.987 | 0.001 | 0 | 1.603 | 270.474 | 21.095 | 2.217 | 59.312 | 0.857 | -1.48 | 0 | 0 | 1 | 43.688 | 1 | 20 | 1 | GJB_Rinse | 0
104 | P-24822 | 2023/06/23 04:01:37 | 1-20 | 4 | -4.805 | -64.411 | 12206 | 0 | 1 | UNDE4-B-4 | 22-304 | H2O | 1687514803.34 | 0.226 | 0.993 | 64.576 | 0 | 0.004 | 1.607 | 265.49 | 20.276 | 2.191 | 59.066 | 0.856 | -1.323 | 0 | 0 | 1 | 43.625 | 1 | 20 | 1 | GJB_Rinse | 0
105 | P-24823 | 2023/06/23 04:11:26 | 1-21 | 1 | -8.53 | -76.829 | 12991 | -1 | 1 | UNDE4-B-5 | 22-304 | H2O | 1687515391.6 | 0.234 | 0.985 | 67.399 | 0.001 | -0.001 | 1.646 | 421.791 | 17.555 | 2.296 | 58.203 | 0.858 | -1.682 | 0 | 0 | 1 | 43.812 | 1 | 21 | 1 | GJB_Rinse | 0
106 | P-24823 | 2023/06/23 04:20:23 | 1-21 | 2 | -8.693 | -78.004 | 12226 | -1 | 1 | UNDE4-B-5 | 22-304 | H2O | 1687515927.95 | 0.209 | 0.983 | 68.12 | 0.001 | 0.003 | 1.693 | 398.317 | 16.201 | 2.156 | 55.661 | 0.856 | -1.523 | 0 | 0 | 1 | 43.688 | 1 | 21 | 1 | GJB_Rinse | 0
107 | P-24823 | 2023/06/23 04:29:19 | 1-21 | 3 | -8.676 | -78.249 | 11945 | -1 | 1 | UNDE4-B-5 | 22-304 | H2O | 1687516464.98 | 0.254 | 0.892 | 64.199 | 0.001 | 0.002 | 1.604 | 389.799 | 15.752 | 2.12 | 54.557 | 0.857 | -1.617 | 0 | 0 | 1 | 43.5 | 1 | 21 | 1 | GJB_Rinse | 0
108 | P-24823 | 2023/06/23 04:38:16 | 1-21 | 4 | -8.731 | -78.432 | 12183 | 0 | 1 | UNDE4-B-5 | 22-304 | H2O | 1687517001.45 | 0.226 | 0.918 | 65.545 | 0.001 | 0.003 | 1.627 | 396.897 | 16.584 | 2.149 | 54.836 | 0.857 | -1.562 | 0 | 0 | 1 | 43.438 | 1 | 21 | 1 | GJB_Rinse | 0
109 | P-24824 | 2023/06/23 04:48:01 | 1-22 | 1 | -8.922 | -68.275 | 14272 | -1 | 1 | EV | standard | H2O | 1687517585.81 | 0.211 | 0.684 | 66.838 | 0.002 | 0 | 1.632 | 48.965 | 34.171 | 2.406 | 69.832 | 0.857 | -0.953 | 0 | 0 | 1 | 43.375 | 1 | 22 | 1 | GJB_Rinse | 0
110 | P-24824 | 2023/06/23 04:56:57 | 1-22 | 2 | -8.925 | -67.216 | 12345 | -1 | 1 | EV | standard | H2O | 1687518122.9 | 0.223 | 0.798 | 64.764 | 0.002 | 0.003 | 1.604 | 42.09 | 28.625 | 2.101 | 62.028 | 0.855 | -0.951 | 0 | 0 | 1 | 43.438 | 1 | 22 | 1 | GJB_Rinse | 0
111 | P-24824 | 2023/06/23 05:05:53 | 1-22 | 3 | -8.9 | -66.736 | 12208 | -1 | 1 | EV | standard | H2O | 1687518659 | 0.202 | 0.721 | 62.757 | 0.001 | 0.002 | 1.551 | 40.977 | 27.914 | 2.089 | 63.364 | 0.855 | -0.808 | 0 | 0 | 1 | 43.375 | 1 | 22 | 1 | GJB_Rinse | 0
112 | P-24824 | 2023/06/23 05:14:51 | 1-22 | 4 | -8.959 | -66.866 | 12142 | 0 | 1 | EV | standard | H2O | 1687519196.52 | 0.21 | 0.769 | 64.569 | 0.002 | -0.001 | 1.596 | 40.524 | 28.281 | 2.067 | 61.174 | 0.856 | -1.058 | 0 | 0 | 1 | 43.438 | 1 | 22 | 1 | GJB_Rinse | 0
113 | P-24825 | 2023/06/23 05:24:36 | 1-23 | 1 | -8.132 | -71.199 | 14861 | -1 | 1 | UNDE4-C-3 | 22-304 | H2O | 1687519781.54 | 0.243 | 0.76 | 66.316 | 0.002 | -0.003 | 1.606 | 324.393 | 26.533 | 2.556 | 67.212 | 0.858 | -1.692 | 0 | 0 | 1 | 43.438 | 1 | 23 | 1 | GJB_Rinse | 0
114 | P-24825 | 2023/06/23 05:33:32 | 1-23 | 2 | -8.036 | -70.941 | 12199 | -1 | 1 | UNDE4-C-3 | 22-304 | H2O | 1687520318.58 | 0.215 | 0.83 | 65.441 | 0.001 | 0.005 | 1.625 | 269.445 | 20.371 | 2.146 | 58.203 | 0.856 | -1.309 | 0 | 0 | 1 | 43.438 | 1 | 23 | 1 | GJB_Rinse | 0
115 | P-24825 | 2023/06/23 05:42:28 | 1-23 | 3 | -8.018 | -71.448 | 12344 | -1 | 1 | UNDE4-C-3 | 22-304 | H2O | 1687520854.14 | 0.224 | 0.872 | 65.806 | 0.001 | -0.001 | 1.62 | 272.565 | 20.498 | 2.177 | 59.203 | 0.857 | -1.223 | 0 | 0 | 1 | 43.438 | 1 | 23 | 1 | GJB_Rinse | 0
116 | P-24825 | 2023/06/23 05:51:25 | 1-23 | 4 | -8.046 | -71.086 | 12220 | 0 | 1 | UNDE4-C-3 | 22-304 | H2O | 1687521390.37 | 0.218 | 0.862 | 65.769 | 0.001 | 0.003 | 1.636 | 270.233 | 20.339 | 2.156 | 58.175 | 0.856 | -1.374 | 0 | 0 | 1 | 43.312 | 1 | 23 | 1 | GJB_Rinse | 0
117 | P-24826 | 2023/06/23 06:01:12 | 1-24 | 1 | -2.724 | -60.298 | 14737 | -1 | 1 | UNDE4-D-1 | 22-304 | H2O | 1687521976.88 | 0.213 | 0.654 | 68.721 | 0.001 | 0 | 1.665 | 55.296 | 34.75 | 2.553 | 73.232 | 0.857 | -1.184 | 0 | 0 | 1 | 43.25 | 1 | 24 | 1 | GJB_Rinse | 0
118 | P-24826 | 2023/06/23 06:10:08 | 1-24 | 2 | -2.502 | -59.312 | 12500 | -1 | 1 | UNDE4-D-1 | 22-304 | H2O | 1687522513.96 | 0.202 | 0.78 | 67.426 | 0 | 0.002 | 1.667 | 47.675 | 28.407 | 2.185 | 63.188 | 0.856 | -1.234 | 0 | 0 | 1 | 43.312 | 1 | 24 | 1 | GJB_Rinse | 0
119 | P-24826 | 2023/06/23 06:19:05 | 1-24 | 3 | -2.39 | -59.253 | 12418 | -1 | 1 | UNDE4-D-1 | 22-304 | H2O | 1687523050.48 | 0.192 | 0.76 | 66.144 | 0.001 | -0.001 | 1.643 | 47.057 | 28.031 | 2.172 | 63.138 | 0.856 | -1.224 | 0 | 0 | 1 | 43.375 | 1 | 24 | 1 | GJB_Rinse | 0
120 | P-24826 | 2023/06/23 06:28:00 | 1-24 | 4 | -2.394 | -59.014 | 12363 | 0 | 1 | UNDE4-D-1 | 22-304 | H2O | 1687523586.52 | 0.195 | 0.855 | 63.506 | 0 | -0.001 | 1.569 | 46.572 | 27.914 | 2.172 | 63.595 | 0.856 | -1.135 | 0 | 0 | 1 | 43.312 | 1 | 24 | 1 | GJB_Rinse | 0
121 | P-24827 | 2023/06/23 06:37:47 | 1-25 | 1 | -5.156 | -63.753 | 15179 | -1 | 1 | UNDE4-C-5 | 22-304 | H2O | 1687524172.99 | 0.189 | 0.688 | 70.264 | 0 | 0 | 1.69 | 47.222 | 36.708 | 2.586 | 74.057 | 0.857 | -1.189 | 0 | 0 | 1 | 43.375 | 1 | 25 | 1 | GJB_Rinse | 0
122 | P-24827 | 2023/06/23 06:46:44 | 1-25 | 2 | -5.315 | -64.18 | 12503 | -1 | 1 | UNDE4-C-5 | 22-304 | H2O | 1687524710.24 | 0.192 | 0.716 | 66.681 | 0.001 | 0.002 | 1.644 | 41.733 | 28.876 | 2.161 | 63.177 | 0.855 | -1.239 | 0 | 0 | 1 | 43.375 | 1 | 25 | 1 | GJB_Rinse | 0
123 | P-24827 | 2023/06/23 06:55:40 | 1-25 | 3 | -5.235 | -64.163 | 12565 | -1 | 1 | UNDE4-C-5 | 22-304 | H2O | 1687525246.22 | 0.219 | 0.948 | 64.179 | 0.001 | -0.001 | 1.579 | 41.768 | 29.128 | 2.164 | 62.46 | 0.855 | -1.19 | 0 | 0 | 1 | 43.375 | 1 | 25 | 1 | GJB_Rinse | 0
124 | P-24827 | 2023/06/23 07:04:37 | 1-25 | 4 | -5.279 | -64.361 | 12305 | 0 | 1 | UNDE4-C-5 | 22-304 | H2O | 1687525782.99 | 0.207 | 0.789 | 65.976 | 0.001 | -0.001 | 1.621 | 41.113 | 28.17 | 2.122 | 62.521 | 0.856 | -1.138 | 0 | 0 | 1 | 43.5 | 1 | 25 | 1 | GJB_Rinse | 0
125 | P-24828 | 2023/06/23 07:14:25 | 1-26 | 1 | -8.562 | -78.302 | 14954 | -1 | 1 | UNDE4-B-1 | 22-304 | H2O | 1687526370.16 | 0.246 | 0.81 | 73.378 | 0.001 | 0.002 | 1.774 | 483.103 | 20.866 | 2.649 | 64.797 | 0.86 | -2.031 | 0 | 0 | 1 | 43.438 | 1 | 26 | 1 | GJB_Rinse | 0
126 | P-24828 | 2023/06/23 07:23:21 | 1-26 | 2 | -8.63 | -79.146 | 12336 | -1 | 1 | UNDE4-B-1 | 22-304 | H2O | 1687526907 | 0.217 | 1.022 | 66.816 | 0.001 | 0.003 | 1.65 | 400.646 | 16.265 | 2.216 | 56.586 | 0.857 | -1.675 | 0 | 0 | 1 | 43.5 | 1 | 26 | 1 | GJB_Rinse | 0
127 | P-24828 | 2023/06/23 07:32:18 | 1-26 | 3 | -8.698 | -79.573 | 12209 | -1 | 1 | UNDE4-B-1 | 22-304 | H2O | 1687527443.28 | 0.242 | 0.941 | 62.998 | 0.001 | -0.001 | 1.569 | 397.715 | 15.718 | 2.201 | 56.222 | 0.856 | -1.614 | 0 | 0 | 1 | 43.5 | 1 | 26 | 1 | GJB_Rinse | 0
128 | P-24828 | 2023/06/23 07:41:14 | 1-26 | 4 | -8.678 | -79.503 | 12019 | 0 | 1 | UNDE4-B-1 | 22-304 | H2O | 1687527980.01 | 0.225 | 0.973 | 62.973 | 0.001 | 0.003 | 1.576 | 392.776 | 15.194 | 2.164 | 55.045 | 0.857 | -1.758 | 0 | 0 | 1 | 43.5 | 1 | 26 | 1 | GJB_Rinse | 0
129 | P-24829 | 2023/06/23 07:51:02 | 1-27 | 1 | -4.769 | -63.815 | 14897 | -1 | 1 | UNDE4-A-2 | 22-304 | H2O | 1687528567.52 | 0.235 | 0.807 | 64.708 | 0 | 0 | 1.573 | 346.909 | 24.71 | 2.72 | 68.211 | 0.858 | -2.315 | 0 | 0 | 1 | 43.5 | 1 | 27 | 1 | GJB_Rinse | 0
130 | P-24829 | 2023/06/23 07:59:57 | 1-27 | 2 | -4.602 | -62.614 | 12389 | -1 | 1 | UNDE4-A-2 | 22-304 | H2O | 1687529103.23 | 0.216 | 0.913 | 66.075 | 0 | 0.002 | 1.643 | 288.991 | 19.556 | 2.301 | 58.838 | 0.856 | -1.84 | 0 | 0 | 1 | 43.5 | 1 | 27 | 1 | GJB_Rinse | 0
131 | P-24829 | 2023/06/23 08:08:54 | 1-27 | 3 | -4.589 | -62.252 | 11993 | -1 | 1 | UNDE4-A-2 | 22-304 | H2O | 1687529639.76 | 0.218 | 0.859 | 63.11 | 0 | 0 | 1.571 | 280.568 | 18.714 | 2.229 | 57.292 | 0.856 | -1.895 | 0 | 0 | 1 | 43.5 | 1 | 27 | 1 | GJB_Rinse | 0
132 | P-24829 | 2023/06/23 08:17:50 | 1-27 | 4 | -4.58 | -62.157 | 12154 | 0 | 1 | UNDE4-A-2 | 22-304 | H2O | 1687530175.89 | 0.214 | 0.98 | 63.344 | 0 | 0 | 1.563 | 284.088 | 18.596 | 2.258 | 58.539 | 0.856 | -1.959 | 0 | 0 | 1 | 43.438 | 1 | 27 | 1 | GJB_Rinse | 0
133 | P-24830 | 2023/06/23 08:27:38 | 1-28 | 1 | -5.054 | -62.839 | 14777 | -1 | 1 | UNDE4-D-3 | 22-304 | H2O | 1687530763.5 | 0.196 | 0.802 | 68.671 | 0 | 0.003 | 1.663 | 51.321 | 34.156 | 2.606 | 74.052 | 0.856 | -1.604 | 0 | 0 | 1 | 43.5 | 1 | 28 | 1 | GJB_Rinse | 0
134 | P-24830 | 2023/06/23 08:36:34 | 1-28 | 2 | -4.98 | -62.637 | 12029 | -1 | 1 | UNDE4-D-3 | 22-304 | H2O | 1687531299.9 | 0.203 | 0.862 | 64.172 | 0.002 | 0.002 | 1.6 | 42.974 | 26.966 | 2.146 | 61.242 | 0.855 | -1.427 | 0 | 0 | 1 | 43.688 | 1 | 28 | 1 | GJB_Rinse | 0
135 | P-24830 | 2023/06/23 08:45:31 | 1-28 | 3 | -5.048 | -62.531 | 11816 | -1 | 1 | UNDE4-D-3 | 22-304 | H2O | 1687531837.19 | 0.194 | 0.786 | 64.526 | 0.001 | -0.001 | 1.615 | 41.83 | 26.1 | 2.103 | 61.9 | 0.855 | -1.336 | 0 | 0 | 1 | 43.75 | 1 | 28 | 1 | GJB_Rinse | 0
136 | P-24830 | 2023/06/23 08:54:27 | 1-28 | 4 | -4.999 | -62.309 | 11756 | 0 | 1 | UNDE4-D-3 | 22-304 | H2O | 1687532373.42 | 0.191 | 0.849 | 66.537 | 0 | 0.005 | 1.659 | 41.647 | 25.931 | 2.098 | 60.529 | 0.856 | -1.534 | 0 | 0 | 1 | 43.688 | 1 | 28 | 1 | GJB_Rinse | 0
137 | P-24831 | 2023/06/23 09:04:13 | 1-29 | 1 | -8.559 | -71.1 | 16378 | -1 | 1 | UNDE4-deep-3 | 22-304 | H2O | 1687532957.9 | 0.2 | 0.63 | 73.739 | 0.002 | -0.002 | 1.763 | 48.86 | 40.702 | 2.73 | 78.22 | 0.858 | -1.226 | 0 | 0 | 1 | 43.75 | 1 | 29 | 1 | GJB_Rinse | 0
138 | P-24831 | 2023/06/23 09:13:08 | 1-29 | 2 | -8.678 | -71.573 | 12327 | -1 | 1 | UNDE4-deep-3 | 22-304 | H2O | 1687533493.68 | 0.201 | 0.832 | 66.023 | 0.001 | 0 | 1.643 | 40.591 | 28.368 | 2.101 | 62.611 | 0.855 | -0.906 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
139 | P-24831 | 2023/06/23 09:22:04 | 1-29 | 3 | -8.707 | -71.955 | 12293 | -1 | 1 | UNDE4-deep-3 | 22-304 | H2O | 1687534029.36 | 0.194 | 0.823 | 68.12 | 0.001 | 0.003 | 1.683 | 40.314 | 28.713 | 2.089 | 62.174 | 0.856 | -0.948 | 0 | 0 | 1 | 43.688 | 1 | 29 | 1 | GJB_Rinse | 0
140 | P-24831 | 2023/06/23 09:30:59 | 1-29 | 4 | -8.702 | -71.825 | 12238 | 0 | 1 | UNDE4-deep-3 | 22-304 | H2O | 1687534565.08 | 0.182 | 0.792 | 66.915 | 0.001 | 0.002 | 1.657 | 40.283 | 28.443 | 2.072 | 61.313 | 0.856 | -0.951 | 0 | 0 | 1 | 43.75 | 1 | 29 | 1 | GJB_Rinse | 0
141 | P-24832 | 2023/06/23 09:40:45 | 1-30 | 1 | -3.844 | -57.627 | 15402 | -1 | 1 | UNDE4-E-5 | 22-304 | H2O | 1687535149.39 | 0.201 | 0.751 | 71.623 | 0.001 | 0.002 | 1.736 | 107.279 | 34.447 | 2.732 | 74.1 | 0.857 | -1.895 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
142 | P-24832 | 2023/06/23 09:49:41 | 1-30 | 2 | -3.713 | -56.447 | 12403 | -1 | 1 | UNDE4-E-5 | 22-304 | H2O | 1687535686.82 | 0.204 | 0.859 | 68.683 | 0 | -0.001 | 1.702 | 89.553 | 26.142 | 2.242 | 62.306 | 0.855 | -1.577 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
143 | P-24832 | 2023/06/23 09:58:37 | 1-30 | 3 | -3.644 | -56.246 | 12188 | -1 | 1 | UNDE4-E-5 | 22-304 | H2O | 1687536222.46 | 0.216 | 0.845 | 70.064 | 0.001 | 0.003 | 1.735 | 88.398 | 25.171 | 2.212 | 62.811 | 0.856 | -1.413 | 0 | 0 | 1 | 43.75 | 1 | 30 | 1 | GJB_Rinse | 0
144 | P-24832 | 2023/06/23 10:07:33 | 1-30 | 4 | -3.609 | -55.952 | 12121 | 0 | 1 | UNDE4-E-5 | 22-304 | H2O | 1687536759.36 | 0.204 | 0.779 | 64.645 | 0 | 0.003 | 1.609 | 88.04 | 25.467 | 2.199 | 61.245 | 0.855 | -1.635 | 0 | 0 | 1 | 43.688 | 1 | 30 | 1 | GJB_Rinse | 0
145 | P-24833 | 2023/06/23 10:17:19 | 1-31 | 1 | -8.666 | -65.926 | 15382 | -1 | 1 | EV | standard | H2O | 1687537344.47 | 0.203 | 0.747 | 73.373 | 0.002 | 0.004 | 1.77 | 47.098 | 37.654 | 2.584 | 74.606 | 0.857 | -1.039 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
146 | P-24833 | 2023/06/23 10:26:15 | 1-31 | 2 | -8.848 | -66.072 | 12302 | -1 | 1 | EV | standard | H2O | 1687537881.11 | 0.185 | 0.692 | 65.827 | 0.001 | 0 | 1.624 | 40.268 | 28.68 | 2.098 | 62.396 | 0.856 | -0.849 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
147 | P-24833 | 2023/06/23 10:35:11 | 1-31 | 3 | -8.924 | -66.714 | 12188 | -1 | 1 | EV | standard | H2O | 1687538416.3 | 0.192 | 0.856 | 64.518 | 0.001 | -0.001 | 1.596 | 40.017 | 28.147 | 2.072 | 62.014 | 0.855 | -1.086 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
148 | P-24833 | 2023/06/23 10:44:06 | 1-31 | 4 | -8.854 | -66.385 | 12111 | 0 | 1 | EV | standard | H2O | 1687538951.97 | 0.197 | 0.798 | 66.441 | 0.001 | -0.001 | 1.656 | 39.677 | 28.311 | 2.06 | 61.249 | 0.855 | -0.938 | 0 | 0 | 1 | 43.688 | 1 | 31 | 1 | GJB_Rinse | 0
149 | P-24834 | 2023/06/23 10:53:52 | 1-32 | 1 | -7.714 | -73.297 | 14963 | -1 | 1 | UNDE4-B-2 | 22-304 | H2O | 1687539537.27 | 0.221 | 0.713 | 72.225 | 0.002 | 0.002 | 1.751 | 186.8 | 31.091 | 2.594 | 69.357 | 0.857 | -1.826 | 0 | 0 | 1 | 43.75 | 1 | 32 | 1 | GJB_Rinse | 0
150 | P-24834 | 2023/06/23 11:02:48 | 1-32 | 2 | -7.62 | -73.743 | 12387 | -1 | 1 | UNDE4-B-2 | 22-304 | H2O | 1687540073.58 | 0.22 | 0.871 | 69.573 | 0.001 | 0.003 | 1.719 | 157.46 | 24.42 | 2.185 | 60.229 | 0.856 | -1.51 | 0 | 0 | 1 | 43.812 | 1 | 32 | 1 | GJB_Rinse | 0
151 | P-24834 | 2023/06/23 11:11:45 | 1-32 | 3 | -7.638 | -73.997 | 12155 | -1 | 1 | UNDE4-B-2 | 22-304 | H2O | 1687540610.29 | 0.218 | 0.939 | 64.475 | 0 | 0.002 | 1.592 | 154.986 | 23.872 | 2.152 | 60.284 | 0.856 | -1.304 | 0 | 0 | 1 | 43.812 | 1 | 32 | 1 | GJB_Rinse | 0
152 | P-24834 | 2023/06/23 11:20:41 | 1-32 | 4 | -7.595 | -73.898 | 12164 | 0 | 1 | UNDE4-B-2 | 22-304 | H2O | 1687541146.65 | 0.207 | 0.884 | 63.48 | 0 | 0.003 | 1.579 | 155.013 | 23.803 | 2.152 | 59.448 | 0.855 | -1.38 | 0 | 0 | 1 | 43.812 | 1 | 32 | 1 | GJB_Rinse | 0
153 | P-24835 | 2023/06/23 11:30:28 | 1-33 | 1 | -4.95 | -49.862 | 15546 | -1 | 1 | UNDE4-1020-2 | 22-304 | H2O | 1687541732.82 | 0.221 | 0.787 | 68.56 | 0.001 | 0.001 | 1.653 | 48.462 | 38.016 | 2.64 | 75.947 | 0.857 | -1.141 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
154 | P-24835 | 2023/06/23 11:39:23 | 1-33 | 2 | -4.854 | -47.844 | 12101 | -1 | 1 | UNDE4-1020-2 | 22-304 | H2O | 1687542269.07 | 0.183 | 0.775 | 64.668 | 0 | -0.001 | 1.6 | 40.086 | 28.307 | 2.089 | 61.427 | 0.855 | -0.946 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
155 | P-24835 | 2023/06/23 11:48:19 | 1-33 | 3 | -4.775 | -47.09 | 12498 | -1 | 1 | UNDE4-1020-2 | 22-304 | H2O | 1687542804.94 | 0.213 | 0.81 | 62.498 | 0 | -0.003 | 1.539 | 40.774 | 29.299 | 2.157 | 63.614 | 0.855 | -1.163 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
156 | P-24835 | 2023/06/23 11:57:15 | 1-33 | 4 | -4.758 | -46.802 | 12523 | 0 | 1 | UNDE4-1020-2 | 22-304 | H2O | 1687543340.88 | 0.182 | 0.825 | 64.953 | 0.001 | 0.002 | 1.596 | 40.683 | 29.357 | 2.163 | 63.618 | 0.855 | -0.971 | 0 | 0 | 1 | 43.812 | 1 | 33 | 1 | GJB_Rinse | 0
157 | P-24836 | 2023/06/23 12:07:02 | 1-34 | 1 | -0.222 | -52.01 | 16754 | -1 | 1 | UNDE4-A-3 | 22-304 | H2O | 1687543927.37 | 0.201 | 0.719 | 74.074 | 0.001 | -0.001 | 1.765 | 49.513 | 41.402 | 2.829 | 81.102 | 0.858 | -1.113 | 0 | 0 | 1 | 43.875 | 1 | 34 | 1 | GJB_Rinse | 0
158 | P-24836 | 2023/06/23 12:15:58 | 1-34 | 2 | 0.002 | -51.714 | 12356 | -1 | 1 | UNDE4-A-3 | 22-304 | H2O | 1687544463.71 | 0.187 | 0.848 | 64.543 | 0.001 | 0.001 | 1.6 | 40.309 | 28.604 | 2.126 | 63.587 | 0.855 | -0.97 | 0 | 0 | 1 | 43.75 | 1 | 34 | 1 | GJB_Rinse | 0
159 | P-24836 | 2023/06/23 12:24:53 | 1-34 | 3 | 0.017 | -52.215 | 12504 | -1 | 1 | UNDE4-A-3 | 22-304 | H2O | 1687544998.84 | 0.189 | 0.802 | 64.575 | 0.001 | -0.001 | 1.6 | 40.589 | 29.239 | 2.16 | 63.189 | 0.855 | -1.072 | 0 | 0 | 1 | 43.812 | 1 | 34 | 1 | GJB_Rinse | 0
160 | P-24836 | 2023/06/23 12:33:49 | 1-34 | 4 | 0.09 | -51.981 | 12366 | 0 | 1 | UNDE4-A-3 | 22-304 | H2O | 1687545535.06 | 0.214 | 0.892 | 66.611 | 0.001 | -0.003 | 1.66 | 40.377 | 28.601 | 2.134 | 62.834 | 0.855 | -1.038 | 0 | 0 | 1 | 43.75 | 1 | 34 | 1 | GJB_Rinse | 0
161 | P-24837 | 2023/06/23 12:43:36 | 1-35 | 1 | -7.994 | -66.314 | 14992 | -1 | 1 | UNDE4-3040-3 | 22-304 | H2O | 1687546121.86 | 0.205 | 0.831 | 67.02 | 0.001 | 0.001 | 1.614 | 45.483 | 36.61 | 2.512 | 72.481 | 0.856 | -1.064 | 0 | 0 | 1 | 43.75 | 1 | 35 | 1 | GJB_Rinse | 0
162 | P-24837 | 2023/06/23 12:52:32 | 1-35 | 2 | -8.327 | -67.141 | 11806 | -1 | 1 | UNDE4-3040-3 | 22-304 | H2O | 1687546658.69 | 0.178 | 0.776 | 62.635 | 0.001 | 0.002 | 1.578 | 39.143 | 27.127 | 2.01 | 59.999 | 0.855 | -1.026 | 0 | 0 | 1 | 43.75 | 1 | 35 | 1 | GJB_Rinse | 0
163 | P-24837 | 2023/06/23 13:01:29 | 1-35 | 3 | -8.347 | -67.432 | 12059 | -1 | 1 | UNDE4-3040-3 | 22-304 | H2O | 1687547194.39 | 0.189 | 0.753 | 63.758 | 0.001 | 0.002 | 1.577 | 39.581 | 27.992 | 2.051 | 60.925 | 0.855 | -1.007 | 0 | 0 | 1 | 43.688 | 1 | 35 | 1 | GJB_Rinse | 0
164 | P-24837 | 2023/06/23 13:10:25 | 1-35 | 4 | -8.376 | -67.437 | 12027 | 0 | 1 | UNDE4-3040-3 | 22-304 | H2O | 1687547731.11 | 0.205 | 0.823 | 63.386 | 0.001 | 0.001 | 1.58 | 39.765 | 27.554 | 2.039 | 60.592 | 0.855 | -1.132 | 0 | 0 | 1 | 43.688 | 1 | 35 | 1 | GJB_Rinse | 0
165 | P-24838 | 2023/06/23 13:20:09 | 1-36 | 1 | -2.897 | -60.253 | 13337 | -1 | 1 | UNDE4-E-4 | 22-304 | H2O | 1687548314.72 | 0.177 | 0.832 | 67.427 | 0 | -0.002 | 1.657 | 43.023 | 31.312 | 2.302 | 65.174 | 0.856 | -1.503 | 0 | 0 | 1 | 43.5 | 1 | 36 | 1 | GJB_Rinse | 0
166 | P-24838 | 2023/06/23 13:29:06 | 1-36 | 2 | -2.605 | -59.111 | 10688 | -1 | 1 | UNDE4-E-4 | 22-304 | H2O | 1687548852.73 | 0.189 | 0.913 | 59.438 | 0.001 | 0.001 | 1.518 | 37.472 | 23.524 | 1.873 | 56.277 | 0.855 | -1.004 | 0 | 0 | 1 | 43.5 | 1 | 36 | 1 | GJB_Rinse | 0
167 | P-24838 | 2023/06/23 13:38:01 | 1-36 | 3 | -2.611 | -59.229 | 11557 | -1 | 1 | UNDE4-E-4 | 22-304 | H2O | 1687549387.13 | 0.205 | 0.798 | 61.867 | 0 | -0.002 | 1.546 | 39.348 | 25.755 | 2.018 | 59.743 | 0.855 | -1.163 | 0 | 0 | 1 | 43.5 | 1 | 36 | 1 | GJB_Rinse | 0
168 | P-24838 | 2023/06/23 13:46:57 | 1-36 | 4 | -2.596 | -59.064 | 11450 | 0 | 1 | UNDE4-E-4 | 22-304 | H2O | 1687549923.02 | 0.214 | 0.858 | 59.381 | 0.001 | 0.002 | 1.487 | 38.941 | 25.931 | 2.001 | 58.844 | 0.855 | -1.256 | 0 | 0 | 1 | 43.5 | 1 | 36 | 1 | GJB_Rinse | 0
169 | P-24839 | 2023/06/23 13:56:40 | 1-37 | 1 | 1.71 | -50.196 | 11920 | -1 | 1 | UNDE4-D-2 | 22-304 | H2O | 1687550506.08 | 0.192 | 0.786 | 67.39 | 0 | 0.001 | 1.678 | 40.018 | 27.318 | 2.079 | 60.373 | 0.855 | -1.317 | 0 | 0 | 1 | 43.438 | 1 | 37 | 1 | GJB_Rinse | 0
170 | P-24839 | 2023/06/23 14:05:36 | 1-37 | 2 | 2.006 | -49.055 | 12336 | -1 | 1 | UNDE4-D-2 | 22-304 | H2O | 1687551041.34 | 0.213 | 0.841 | 62.527 | 0.001 | 0.002 | 1.545 | 40.714 | 28.115 | 2.17 | 63.311 | 0.855 | -1.124 | -1e-04 | -1e-04 | 1 | 43.5 | 1 | 37 | 1 | GJB_Rinse | 0
171 | P-24839 | 2023/06/23 14:14:31 | 1-37 | 3 | 2.012 | -48.95 | 12223 | -1 | 1 | UNDE4-D-2 | 22-304 | H2O | 1687551577.08 | 0.207 | 0.678 | 63.847 | 0 | 0.004 | 1.578 | 40.467 | 27.982 | 2.16 | 63.338 | 0.855 | -1.078 | 0 | 0 | 1 | 43.5 | 1 | 37 | 1 | GJB_Rinse | 0
172 | P-24839 | 2023/06/23 14:23:26 | 1-37 | 4 | 2.068 | -48.549 | 12408 | 0 | 1 | UNDE4-D-2 | 22-304 | H2O | 1687552112.28 | 0.18 | 0.704 | 59.642 | 0 | -0.001 | 1.477 | 40.924 | 28.588 | 2.17 | 62.866 | 0.855 | -1.346 | 0 | 0 | 1 | 43.5 | 1 | 37 | 1 | GJB_Rinse | 0
173 | P-24840 | 2023/06/23 14:33:11 | 1-38 | 1 | -4.359 | -60.585 | 11484 | -1 | 1 | UNDE4-E-3 | 22-304 | H2O | 1687552696.97 | 0.204 | 0.848 | 62.845 | 0.001 | 0.001 | 1.576 | 42.313 | 24.141 | 2.193 | 59.788 | 0.855 | -2.183 | 0 | 0 | 1 | 43.5 | 1 | 38 | 1 | GJB_Rinse | 0
174 | P-24840 | 2023/06/23 14:42:07 | 1-38 | 2 | -4.858 | -62.235 | 12075 | -1 | 1 | UNDE4-E-3 | 22-304 | H2O | 1687553232.32 | 0.218 | 0.808 | 64.191 | 0.002 | 0.002 | 1.599 | 43.833 | 25.565 | 2.304 | 61.863 | 0.856 | -2.294 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
175 | P-24840 | 2023/06/23 14:51:03 | 1-38 | 3 | -4.932 | -62.581 | 11965 | -1 | 1 | UNDE4-E-3 | 22-304 | H2O | 1687553768.84 | 0.218 | 0.761 | 64.127 | 0.001 | -0.001 | 1.598 | 43.587 | 25.074 | 2.277 | 62.082 | 0.855 | -2.225 | 0 | 0 | 1 | 43.562 | 1 | 38 | 1 | GJB_Rinse | 0
176 | P-24840 | 2023/06/23 14:59:59 | 1-38 | 4 | -5.015 | -62.961 | 11980 | 0 | 1 | UNDE4-E-3 | 22-304 | H2O | 1687554304.5 | 0.179 | 0.89 | 64.749 | 0.001 | -0.001 | 1.612 | 43.564 | 25.026 | 2.299 | 62.638 | 0.855 | -2.227 | 0 | 0 | 1 | 43.625 | 1 | 38 | 1 | GJB_Rinse | 0
177 | P-24841 | 2023/06/23 15:09:44 | 1-39 | 1 | -2.888 | -45.261 | 12297 | -1 | 1 | UNDE4-1020-3 | 22-304 | H2O | 1687554889.33 | 0.206 | 0.843 | 63.868 | 0.001 | 0 | 1.581 | 40.047 | 28.6 | 2.141 | 62.185 | 0.855 | -1.078 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
178 | P-24841 | 2023/06/23 15:18:39 | 1-39 | 2 | -2.655 | -43.142 | 12267 | -1 | 1 | UNDE4-1020-3 | 22-304 | H2O | 1687555424.51 | 0.164 | 0.813 | 65.855 | 0 | 0.003 | 1.641 | 39.877 | 28.517 | 2.116 | 63.097 | 0.855 | -1.052 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
179 | P-24841 | 2023/06/23 15:27:35 | 1-39 | 3 | -2.668 | -42.757 | 12294 | -1 | 1 | UNDE4-1020-3 | 22-304 | H2O | 1687555961.4 | 0.217 | 0.809 | 64.875 | 0.002 | -0.003 | 1.605 | 40.038 | 28.57 | 2.128 | 62.592 | 0.855 | -1.131 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
180 | P-24841 | 2023/06/23 15:36:32 | 1-39 | 4 | -2.591 | -42.407 | 12266 | 0 | 1 | UNDE4-1020-3 | 22-304 | H2O | 1687556497.87 | 0.218 | 0.813 | 61.365 | 0.001 | 0 | 1.537 | 39.924 | 28.571 | 2.131 | 62.285 | 0.855 | -1.104 | 0 | 0 | 1 | 43.625 | 1 | 39 | 1 | GJB_Rinse | 0
181 | P-24842 | 2023/06/23 15:46:17 | 1-40 | 1 | -8.458 | -63.493 | 13417 | -1 | 1 | EV | standard | H2O | 1687557082.66 | 0.173 | 0.781 | 63.492 | 0.001 | 0 | 1.552 | 42.131 | 31.669 | 2.276 | 67.227 | 0.856 | -0.963 | 0 | 0 | 1 | 43.562 | 1 | 40 | 1 | GJB_Rinse | 0
182 | P-24842 | 2023/06/23 15:55:12 | 1-40 | 2 | -8.764 | -65.418 | 12616 | -1 | 1 | EV | standard | H2O | 1687557617.59 | 0.202 | 0.757 | 65.854 | 0.001 | 0 | 1.63 | 40.384 | 29.453 | 2.162 | 64.654 | 0.856 | -0.858 | 0 | 0 | 1 | 43.625 | 1 | 40 | 1 | GJB_Rinse | 0
183 | P-24842 | 2023/06/23 16:04:08 | 1-40 | 3 | -8.81 | -65.924 | 12412 | -1 | 1 | EV | standard | H2O | 1687558153.92 | 0.186 | 0.855 | 63.714 | 0.001 | 0.003 | 1.562 | 40.032 | 29.343 | 2.102 | 62.182 | 0.855 | -1.077 | 0 | 0 | 1 | 43.625 | 1 | 40 | 1 | GJB_Rinse | 0
184 | P-24842 | 2023/06/23 16:13:04 | 1-40 | 4 | -8.844 | -66.01 | 12258 | 0 | 1 | EV | standard | H2O | 1687558689.57 | 0.208 | 0.872 | 64.651 | 0.001 | -0.001 | 1.606 | 39.863 | 28.51 | 2.078 | 61.019 | 0.856 | -0.999 | 0 | 0 | 1 | 43.625 | 1 | 40 | 1 | GJB_Rinse | 0
185 | P-24843 | 2023/06/23 16:22:51 | 1-41 | 1 | -4.685 | -67.83 | 17443 | -1 | 1 | UNDE4-A-5 | 22-304 | H2O | 1687559276.07 | 0.191 | 0.699 | 77.344 | 0.001 | 0 | 1.835 | 53.102 | 43.04 | 2.933 | 82.948 | 0.858 | -1.485 | 0 | 0 | 1 | 43.625 | 1 | 41 | 1 | GJB_Rinse | 0
186 | P-24843 | 2023/06/23 16:31:47 | 1-41 | 2 | -4.503 | -67.681 | 11777 | -1 | 1 | UNDE4-A-5 | 22-304 | H2O | 1687559812.31 | 0.21 | 0.839 | 63.31 | 0.001 | 0.002 | 1.576 | 40.452 | 26.649 | 2.045 | 61.063 | 0.855 | -1.083 | 0 | 0 | 1 | 43.625 | 1 | 41 | 1 | GJB_Rinse | 0
187 | P-24843 | 2023/06/23 16:40:42 | 1-41 | 3 | -4.482 | -67.576 | 11630 | -1 | 1 | UNDE4-A-5 | 22-304 | H2O | 1687560347.46 | 0.189 | 0.761 | 63.745 | 0.001 | 0.002 | 1.597 | 40.264 | 26.208 | 2.018 | 59.877 | 0.855 | -1.09 | 0 | 0 | 1 | 43.562 | 1 | 41 | 1 | GJB_Rinse | 0
188 | P-24843 | 2023/06/23 16:49:36 | 1-41 | 4 | -4.431 | -67.585 | 11346 | 0 | 1 | UNDE4-A-5 | 22-304 | H2O | 1687560882.26 | 0.188 | 0.851 | 61.807 | 0.001 | 0.001 | 1.546 | 39.607 | 25.725 | 1.97 | 57.982 | 0.855 | -1.095 | 0 | 0 | 1 | 43.688 | 1 | 41 | 1 | GJB_Rinse | 0
189 | P-24844 | 2023/06/23 16:59:22 | 1-42 | 1 | 5.069 | -41.145 | 14438 | -1 | 1 | UNDE4-B-3 | 22-304 | H2O | 1687561467.3 | 0.2 | 0.718 | 69.535 | 0 | 0.001 | 1.688 | 46.187 | 34.307 | 2.516 | 71.771 | 0.856 | -1.304 | 0 | 0 | 1 | 43.625 | 1 | 42 | 1 | GJB_Rinse | 0
190 | P-24844 | 2023/06/23 17:08:18 | 1-42 | 2 | 5.463 | -38.515 | 11841 | -1 | 1 | UNDE4-B-3 | 22-304 | H2O | 1687562003.64 | 0.21 | 0.819 | 64.027 | 0 | 0.002 | 1.592 | 40.663 | 26.68 | 2.099 | 61.668 | 0.855 | -1.25 | 0 | 0 | 1 | 43.688 | 1 | 42 | 1 | GJB_Rinse | 0
191 | P-24844 | 2023/06/23 17:17:14 | 1-42 | 3 | 5.565 | -37.897 | 11984 | -1 | 1 | UNDE4-B-3 | 22-304 | H2O | 1687562540.28 | 0.22 | 0.855 | 60.785 | 0.001 | -0.001 | 1.521 | 40.864 | 27.172 | 2.122 | 61.946 | 0.856 | -1.245 | 0 | 0 | 1 | 43.688 | 1 | 42 | 1 | GJB_Rinse | 0
192 | P-24844 | 2023/06/23 17:26:10 | 1-42 | 4 | 5.638 | -37.541 | 12206 | 0 | 1 | UNDE4-B-3 | 22-304 | H2O | 1687563075.97 | 0.196 | 0.83 | 62.986 | 0.001 | -0.001 | 1.559 | 41.193 | 28.156 | 2.156 | 62.465 | 0.855 | -1.194 | 0 | 0 | 1 | 43.75 | 1 | 42 | 1 | GJB_Rinse | 0
193 | P-24845 | 2023/06/23 17:35:54 | 1-43 | 1 | -7.693 | -63.942 | 14717 | -1 | 1 | UNDE4-3040-2 | 22-304 | H2O | 1687563658.95 | 0.208 | 0.752 | 73.923 | 0.002 | -0.003 | 1.794 | 44.931 | 35.339 | 2.486 | 72.375 | 0.857 | -1.006 | 0 | 0 | 1 | 43.688 | 1 | 43 | 1 | GJB_Rinse | 0
194 | P-24845 | 2023/06/23 17:44:50 | 1-43 | 2 | -8.252 | -66.324 | 12407 | -1 | 1 | UNDE4-3040-2 | 22-304 | H2O | 1687564195.18 | 0.211 | 0.834 | 60.963 | 0.001 | 0.001 | 1.509 | 40.176 | 28.73 | 2.101 | 62.863 | 0.856 | -0.955 | 0 | 0 | 1 | 43.688 | 1 | 43 | 1 | GJB_Rinse | 0
195 | P-24845 | 2023/06/23 17:53:46 | 1-43 | 3 | -8.414 | -66.937 | 12018 | -1 | 1 | UNDE4-3040-2 | 22-304 | H2O | 1687564731.58 | 0.198 | 0.766 | 65.335 | 0.001 | 0.002 | 1.621 | 39.296 | 27.86 | 2.046 | 61.188 | 0.855 | -1.019 | 0 | 0 | 1 | 43.625 | 1 | 43 | 1 | GJB_Rinse | 0
196 | P-24845 | 2023/06/23 18:02:41 | 1-43 | 4 | -8.398 | -66.719 | 12064 | 0 | 1 | UNDE4-3040-2 | 22-304 | H2O | 1687565266.37 | 0.202 | 0.805 | 61.005 | 0.001 | 0.002 | 1.51 | 39.32 | 27.987 | 2.059 | 61.13 | 0.855 | -0.892 | 0 | 0 | 1 | 43.75 | 1 | 43 | 1 | GJB_Rinse | 0
197 | P-24846 | 2023/06/23 18:12:25 | 1-44 | 1 | -3.683 | -57.614 | 13142 | -1 | 1 | UNDE4-A-1 | 22-304 | H2O | 1687565850.94 | 0.173 | 0.786 | 68.609 | 0.001 | 0.002 | 1.676 | 41.526 | 31.112 | 2.253 | 65.816 | 0.856 | -0.985 | 0 | 0 | 1 | 43.75 | 1 | 44 | 1 | GJB_Rinse | 0
198 | P-24846 | 2023/06/23 18:21:21 | 1-44 | 2 | -3.333 | -56.503 | 12276 | -1 | 1 | UNDE4-A-1 | 22-304 | H2O | 1687566386.85 | 0.184 | 0.822 | 64.995 | 0.001 | 0.003 | 1.618 | 39.93 | 28.65 | 2.103 | 61.458 | 0.855 | -1.102 | 0 | 0 | 1 | 43.75 | 1 | 44 | 1 | GJB_Rinse | 0
199 | P-24846 | 2023/06/23 18:30:18 | 1-44 | 3 | -3.247 | -55.961 | 12028 | -1 | 1 | UNDE4-A-1 | 22-304 | H2O | 1687566923.88 | 0.194 | 0.783 | 63.296 | 0.001 | 0.003 | 1.577 | 39.356 | 27.763 | 2.074 | 62.152 | 0.856 | -0.874 | 0 | 0 | 1 | 43.688 | 1 | 44 | 1 | GJB_Rinse | 0
200 | P-24846 | 2023/06/23 18:39:13 | 1-44 | 4 | -3.215 | -55.988 | 12068 | 0 | 1 | UNDE4-A-1 | 22-304 | H2O | 1687567458.77 | 0.211 | 0.831 | 63.825 | 0.001 | -0.001 | 1.591 | 39.598 | 27.94 | 2.061 | 61.176 | 0.855 | -1.195 | 0 | 0 | 1 | 43.75 | 1 | 44 | 1 | GJB_Rinse | 0
201 | P-24847 | 2023/06/23 18:48:58 | 1-45 | 1 | 1.149 | -48.454 | 12257 | -1 | 1 | UNDE4-E-1 | 22-304 | H2O | 1687568043.89 | 0.215 | 0.771 | 62.499 | 0.001 | 0.002 | 1.556 | 40.61 | 28.078 | 2.153 | 63.692 | 0.855 | -0.987 | 0 | 0 | 1 | 43.75 | 1 | 45 | 1 | GJB_Rinse | 0
202 | P-24847 | 2023/06/23 18:57:54 | 1-45 | 2 | 1.335 | -47.705 | 11904 | -1 | 1 | UNDE4-E-1 | 22-304 | H2O | 1687568579.64 | 0.196 | 0.786 | 62.244 | 0 | 0 | 1.551 | 39.921 | 27.209 | 2.098 | 61.205 | 0.855 | -1.039 | 0 | 0 | 1 | 43.688 | 1 | 45 | 1 | GJB_Rinse | 0
203 | P-24847 | 2023/06/23 19:06:49 | 1-45 | 3 | 1.432 | -47.438 | 11966 | -1 | 1 | UNDE4-E-1 | 22-304 | H2O | 1687569115.18 | 0.22 | 0.765 | 65.444 | 0.001 | 0.002 | 1.627 | 40.183 | 27.202 | 2.096 | 61.419 | 0.856 | -1.2 | 0 | 0 | 1 | 43.625 | 1 | 45 | 1 | GJB_Rinse | 0
204 | P-24847 | 2023/06/23 19:15:45 | 1-45 | 4 | 1.465 | -47.428 | 11865 | 0 | 1 | UNDE4-E-1 | 22-304 | H2O | 1687569650.98 | 0.196 | 0.806 | 63.557 | 0.001 | -0.001 | 1.576 | 39.941 | 26.937 | 2.086 | 61.187 | 0.856 | -1.101 | 0 | 0 | 1 | 43.375 | 1 | 45 | 1 | GJB_Rinse | 0
205 | P-24848 | 2023/06/23 19:25:31 | 1-46 | 1 | -5.478 | -56.1 | 12267 | -1 | 1 | UNDE4-2030-3 | 22-304 | H2O | 1687570236.71 | 0.193 | 0.774 | 67.91 | 0.001 | 0.002 | 1.68 | 39.746 | 28.6 | 2.109 | 62.568 | 0.855 | -0.959 | 0 | 0 | 1 | 43.312 | 1 | 46 | 1 | GJB_Rinse | 0
206 | P-24848 | 2023/06/23 19:34:26 | 1-46 | 2 | -5.785 | -56.863 | 12182 | -1 | 1 | UNDE4-2030-3 | 22-304 | H2O | 1687570772.01 | 0.203 | 0.764 | 62.906 | 0.001 | 0 | 1.557 | 39.67 | 28.231 | 2.087 | 61.702 | 0.855 | -0.958 | 0 | 0 | 1 | 43.438 | 1 | 46 | 1 | GJB_Rinse | 0
207 | P-24848 | 2023/06/23 19:43:21 | 1-46 | 3 | -5.836 | -57.09 | 12173 | -1 | 1 | UNDE4-2030-3 | 22-304 | H2O | 1687571306.84 | 0.2 | 0.845 | 65.295 | 0.001 | -0.001 | 1.614 | 39.705 | 28.186 | 2.09 | 61.208 | 0.855 | -1.168 | 0 | 0 | 1 | 43.562 | 1 | 46 | 1 | GJB_Rinse | 0
208 | P-24848 | 2023/06/23 19:52:16 | 1-46 | 4 | -5.904 | -57.088 | 12171 | 0 | 1 | UNDE4-2030-3 | 22-304 | H2O | 1687571842.08 | 0.188 | 0.836 | 63.614 | 0.001 | -0.003 | 1.585 | 39.438 | 28.295 | 2.106 | 62.365 | 0.855 | -0.804 | 0 | 0 | 1 | 43.625 | 1 | 46 | 1 | GJB_Rinse | 0
209 | P-24849 | 2023/06/23 20:02:01 | 1-47 | 1 | -8.477 | -65.688 | 12368 | -1 | 1 | UNDE4-deep-1 | 22-304 | H2O | 1687572427.01 | 0.193 | 0.757 | 64.336 | 0.001 | 0.001 | 1.587 | 39.942 | 28.881 | 2.113 | 62.892 | 0.856 | -0.86 | 0 | 0 | 1 | 43.688 | 1 | 47 | 1 | GJB_Rinse | 0
210 | P-24849 | 2023/06/23 20:10:57 | 1-47 | 2 | -8.585 | -66.606 | 12444 | -1 | 1 | UNDE4-deep-1 | 22-304 | H2O | 1687572962.66 | 0.205 | 0.766 | 65.077 | 0.001 | 0.002 | 1.616 | 40.147 | 29.277 | 2.127 | 62.376 | 0.855 | -1.077 | 0 | 0 | 1 | 43.75 | 1 | 47 | 1 | GJB_Rinse | 0
211 | P-24849 | 2023/06/23 20:19:53 | 1-47 | 3 | -8.605 | -66.665 | 12132 | -1 | 1 | UNDE4-deep-1 | 22-304 | H2O | 1687573498.03 | 0.205 | 0.75 | 66.178 | 0.001 | 0.002 | 1.638 | 39.625 | 28.004 | 2.066 | 61.851 | 0.856 | -0.989 | 0 | 0 | 1 | 43.75 | 1 | 47 | 1 | GJB_Rinse | 0
212 | P-24849 | 2023/06/23 20:28:49 | 1-47 | 4 | -8.606 | -66.869 | 12138 | 0 | 1 | UNDE4-deep-1 | 22-304 | H2O | 1687574034.6 | 0.165 | 0.759 | 64.659 | 0.001 | 0.002 | 1.596 | 39.587 | 28.15 | 2.058 | 61.601 | 0.855 | -0.956 | 0 | 0 | 1 | 43.75 | 1 | 47 | 1 | GJB_Rinse | 0
213 | P-24850 | 2023/06/23 20:38:35 | 1-48 | 1 | -6.273 | -67.278 | 12869 | -1 | 1 | WOOD6-D-3 | 23-082 | H2O | 1687574620.08 | 0.172 | 0.845 | 67.363 | 0.001 | 0.002 | 1.649 | 49.761 | 29.395 | 2.248 | 64.912 | 0.856 | -1.273 | 0 | 0 | 1 | 43.625 | 1 | 48 | 1 | GJB_Rinse | 0
214 | P-24850 | 2023/06/23 20:47:31 | 1-48 | 2 | -6.125 | -67.429 | 12272 | -1 | 1 | WOOD6-D-3 | 23-082 | H2O | 1687575156.11 | 0.193 | 0.846 | 63.99 | 0 | 0.002 | 1.584 | 48.482 | 27.513 | 2.135 | 61.867 | 0.855 | -1.383 | 0 | 0 | 1 | 43.562 | 1 | 48 | 1 | GJB_Rinse | 0
215 | P-24850 | 2023/06/23 20:56:26 | 1-48 | 3 | -6.149 | -67.539 | 12175 | -1 | 1 | WOOD6-D-3 | 23-082 | H2O | 1687575692.31 | 0.21 | 0.88 | 64.698 | 0.002 | 0.003 | 1.61 | 48.06 | 27.171 | 2.132 | 62.885 | 0.855 | -1.308 | 0 | 0 | 1 | 43.562 | 1 | 48 | 1 | GJB_Rinse | 0
216 | P-24850 | 2023/06/23 21:05:22 | 1-48 | 4 | -6.129 | -67.65 | 12225 | 0 | 1 | WOOD6-D-3 | 23-082 | H2O | 1687576228.13 | 0.191 | 0.73 | 68.4 | 0.001 | 0.002 | 1.699 | 48.149 | 27.507 | 2.141 | 62.62 | 0.856 | -1.254 | 0 | 0 | 1 | 43.625 | 1 | 48 | 1 | GJB_Rinse | 0
217 | P-24851 | 2023/06/23 21:15:08 | 1-49 | 1 | -8.77 | -66.871 | 13623 | -1 | 1 | EV | standard | H2O | 1687576813.05 | 0.198 | 0.768 | 66.371 | 0.001 | 0.001 | 1.631 | 42.607 | 32.302 | 2.308 | 68.298 | 0.856 | -0.891 | 0 | 0 | 1 | 43.562 | 1 | 49 | 1 | GJB_Rinse | 0
218 | P-24851 | 2023/06/23 21:24:04 | 1-49 | 2 | -8.836 | -66.727 | 12280 | -1 | 1 | EV | standard | H2O | 1687577349.98 | 0.193 | 0.863 | 65.737 | 0.001 | 0.003 | 1.624 | 39.793 | 28.719 | 2.092 | 61.853 | 0.856 | -0.868 | 0 | 0 | 1 | 43.5 | 1 | 49 | 1 | GJB_Rinse | 0
219 | P-24851 | 2023/06/23 21:33:00 | 1-49 | 3 | -8.893 | -66.816 | 12316 | -1 | 1 | EV | standard | H2O | 1687577885.71 | 0.194 | 0.792 | 66.008 | 0.001 | 0.002 | 1.633 | 39.839 | 28.741 | 2.1 | 62.173 | 0.856 | -0.883 | 0 | 0 | 1 | 43.5 | 1 | 49 | 1 | GJB_Rinse | 0
220 | P-24851 | 2023/06/23 21:41:56 | 1-49 | 4 | -8.894 | -66.714 | 12090 | 0 | 1 | EV | standard | H2O | 1687578421.6 | 0.201 | 0.82 | 64.99 | 0.001 | -0.002 | 1.631 | 39.464 | 27.784 | 2.069 | 62.293 | 0.855 | -0.943 | 0 | 0 | 1 | 43.5 | 1 | 49 | 1 | GJB_Rinse | 0
221 | P-24852 | 2023/06/23 21:51:40 | 1-50 | 1 | -7.262 | -73.104 | 13623 | -1 | 1 | WOOD6-D-1 | 23-082 | H2O | 1687579005.59 | 0.192 | 0.7 | 67.822 | 0.001 | -0.002 | 1.656 | 48.685 | 31.749 | 2.336 | 67.891 | 0.856 | -1.239 | 0 | 0 | 1 | 43.438 | 1 | 50 | 1 | GJB_Rinse | 0
222 | P-24852 | 2023/06/23 22:00:35 | 1-50 | 2 | -7.155 | -73.558 | 12262 | -1 | 1 | WOOD6-D-1 | 23-082 | H2O | 1687579540.84 | 0.204 | 0.863 | 59.338 | 0 | 0.004 | 1.467 | 45.299 | 27.79 | 2.123 | 62.722 | 0.855 | -0.984 | 0 | 0 | 1 | 43.25 | 1 | 50 | 1 | GJB_Rinse | 0
223 | P-24852 | 2023/06/23 22:09:30 | 1-50 | 3 | -7.182 | -73.828 | 12143 | -1 | 1 | WOOD6-D-1 | 23-082 | H2O | 1687580076.11 | 0.185 | 0.759 | 62.098 | 0.001 | 0.003 | 1.545 | 44.955 | 27.599 | 2.108 | 62.128 | 0.855 | -0.997 | 0 | 0 | 1 | 43.188 | 1 | 50 | 1 | GJB_Rinse | 0
224 | P-24852 | 2023/06/23 22:18:26 | 1-50 | 4 | -7.189 | -73.971 | 12063 | 0 | 1 | WOOD6-D-1 | 23-082 | H2O | 1687580611.92 | 0.195 | 0.718 | 58.711 | 0 | -0.001 | 1.46 | 44.901 | 27.189 | 2.082 | 61.658 | 0.855 | -1.135 | 0 | 0 | 1 | 43.188 | 1 | 50 | 1 | GJB_Rinse | 0
225 | P-24853 | 2023/06/23 22:28:12 | 1-51 | 1 | -3.695 | -51.049 | 14998 | -1 | 1 | WOOD6-C-4 | 23-082 | H2O | 1687581197.38 | 0.184 | 0.694 | 66.16 | 0.001 | 0 | 1.6 | 80.334 | 34.513 | 2.632 | 73.161 | 0.856 | -1.6 | 0 | 0 | 1 | 43.25 | 1 | 51 | 1 | GJB_Rinse | 0
226 | P-24853 | 2023/06/23 22:37:07 | 1-51 | 2 | -3.508 | -48.887 | 11257 | -1 | 1 | WOOD6-C-4 | 23-082 | H2O | 1687581733.59 | 0.21 | 0.916 | 64.789 | 0.001 | 0.001 | 1.625 | 64.152 | 24.422 | 2.006 | 57.031 | 0.855 | -1.484 | 0 | 0 | 1 | 43.312 | 1 | 51 | 1 | GJB_Rinse | 0
227 | P-24853 | 2023/06/23 22:46:02 | 1-51 | 3 | -3.451 | -48.211 | 10801 | -1 | 1 | WOOD6-C-4 | 23-082 | H2O | 1687582268.8 | 0.206 | 0.903 | 61.839 | 0 | -0.001 | 1.556 | 62.189 | 22.879 | 1.939 | 56.208 | 0.855 | -1.208 | 0 | 0 | 1 | 43.312 | 1 | 51 | 1 | GJB_Rinse | 0
228 | P-24853 | 2023/06/23 22:54:57 | 1-51 | 4 | -3.425 | -47.954 | 10790 | 0 | 1 | WOOD6-C-4 | 23-082 | H2O | 1687582803.47 | 0.181 | 0.793 | 61.421 | 0 | 0.002 | 1.559 | 62.216 | 22.89 | 1.938 | 55.954 | 0.855 | -1.252 | 0 | 0 | 1 | 43.312 | 1 | 51 | 1 | GJB_Rinse | 0
229 | P-24854 | 2023/06/23 23:04:43 | 1-52 | 1 | -5.377 | -50.83 | 11534 | -1 | 1 | WOOD6-C-2 | 23-082 | H2O | 1687583389.23 | 0.235 | 0.919 | 63.118 | 0 | -0.001 | 1.581 | 417.821 | 13.129 | 2.139 | 54.443 | 0.856 | -1.508 | 0 | 0 | 1 | 43.375 | 1 | 52 | 1 | GJB_Rinse | 0
230 | P-24854 | 2023/06/23 23:13:37 | 1-52 | 2 | -5.468 | -51.212 | 11153 | -1 | 1 | WOOD6-C-2 | 23-082 | H2O | 1687583923.16 | 0.229 | 0.967 | 59.69 | 0 | -0.002 | 1.494 | 407.499 | 12.532 | 2.056 | 51.58 | 0.856 | -1.709 | 0 | 0 | 1 | 43.312 | 1 | 52 | 1 | GJB_Rinse | 0
231 | P-24854 | 2023/06/23 23:22:32 | 1-52 | 3 | -5.492 | -51.14 | 9931 | -1 | 1 | WOOD6-C-2 | 23-082 | H2O | 1687584459.86 | 0.236 | 1.045 | 59.373 | 0.001 | 0 | 1.527 | 364.269 | 10.361 | 1.844 | 47.278 | 0.856 | -1.652 | 0 | 0 | 1 | 43.375 | 1 | 52 | 1 | GJB_Rinse | 0
232 | P-24854 | 2023/06/23 23:31:28 | 1-52 | 4 | -5.494 | -51.012 | 9339 | 0 | 1 | WOOD6-C-2 | 23-082 | H2O | 1687584996.51 | 0.243 | 1.121 | 56.217 | 0 | 0 | 1.464 | 344.139 | 9.44 | 1.737 | 44.958 | 0.856 | -1.498 | 0 | 0 | 1 | 43.375 | 1 | 52 | 1 | GJB_Rinse | 0
233 | P-24855 | 2023/06/23 23:41:15 | 1-53 | 1 | -8.628 | -64.622 | 14047 | -1 | 1 | EV | standard | H2O | 1687585580.36 | 0.202 | 0.765 | 69.877 | 0.002 | 0.002 | 1.697 | 48.555 | 33.693 | 2.374 | 68.351 | 0.856 | -1.154 | 0 | 0 | 1 | 43.375 | 1 | 53 | 1 | GJB_Rinse | 0
234 | P-24855 | 2023/06/23 23:50:11 | 1-53 | 2 | -8.686 | -65.48 | 12151 | -1 | 1 | EV | standard | H2O | 1687586116.77 | 0.205 | 0.802 | 62.887 | 0.001 | 0.003 | 1.554 | 41.556 | 27.823 | 2.076 | 62.102 | 0.855 | -0.978 | 0 | 0 | 1 | 43.438 | 1 | 53 | 1 | GJB_Rinse | 0
235 | P-24855 | 2023/06/23 23:59:06 | 1-53 | 3 | -8.767 | -66.068 | 12134 | -1 | 1 | EV | standard | H2O | 1687586651.12 | 0.197 | 0.805 | 64.114 | 0 | 0.002 | 1.595 | 40.803 | 28.031 | 2.07 | 61.007 | 0.855 | -1.061 | 0 | 0 | 1 | 43.375 | 1 | 53 | 1 | GJB_Rinse | 0
236 | P-24855 | 2023/06/24 00:08:01 | 1-53 | 4 | -8.812 | -66.08 | 12035 | 0 | 1 | EV | standard | H2O | 1687587186.31 | 0.183 | 0.808 | 65.456 | 0.001 | 0.003 | 1.621 | 39.982 | 27.733 | 2.042 | 61.638 | 0.855 | -0.84 | 0 | 0 | 1 | 43.5 | 1 | 53 | 1 | GJB_Rinse | 0

## Instructions
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Variable | Instruction | Comment
StandardsFile | standards file.csv
OutputBase | ChemCorrect
CorrectionType | M
timecode | Time Code
identifier | Identifier 1
sample | Sample
injection | Inj Nr | Inj Nr
line | Line
job | Job
Variable | Instruction | Comment
baseline_shift_MEAN | mean_standards(baseline_shift) | The mean of the organic baseline of the standards.
baseline_shift_STD | std_standards(baseline_shift) | The standard deviation of the organic baseline of the standards.
ch4_ppm_MEAN | mean_standards(ch4_ppm) | The mean methane concentration of the standards.
ch4_ppm_STD | std_standards(ch4_ppm) | The standard deviation of the methane concentrations of the standards.
residuals_MEAN | mean_standards(residuals) | The mean difference between the measured and modeled spectra of the standards.
residuals_STD | std_standards(residuals) | The standard deviation of the difference between measured and modeled spectra of the standards.
slope_shift_MEAN | mean_standards(slope_shift) | The mean organic slope of the standards.
slope_shift_STD | std_standards(slope_shift) | The standard deviation of the organic slope of the standards.
baseline_curvature_MEAN | mean_standards(baseline_curvature) | The mean of the baseline curvature of the standards.
baseline_curvature_STD | std_standards(baseline_curvature) | The standard deviation of the baseline curvature of the standards.
residuals_MEAN_ABS | abs(residuals_MEAN) | The absolute value of the mean difference between the measured and modeled spectra of the standards.
baseline_shift_MEAN_ABS | abs(baseline_shift_MEAN) | The absolute value of the mean of the organic baseline of the standards.
ch4_ppm_MEAN_ABS | abs(ch4_ppm_MEAN) | The absolute value of the mean methane concentration of the standards.
slope_shift_MEAN_ABS | abs(slope_shift_MEAN) | The absolute value of the mean organic slope of the standards.
residuals_15SIGMA | evaluation(residuals_STD*1.5) | One and half times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_15SIGMA | evaluation(residuals_MEAN+residuals_15SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_2SIGMA | evaluation(residuals_STD*2) | Twice the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_2SIGMA | evaluation(residuals_MEAN+residuals_2SIGMA) | The mean difference between measured and modeled spectra of the standards plus two standard deviation.
residuals_3SIGMA | evaluation(residuals_STD*3) | Three times the standard deviation of the difference between measured and modeled spectra of the standards.
residuals_PLUS_3SIGMA | evaluation(residuals_MEAN+residuals_3SIGMA) | The mean difference between measured and modeled spectra of the standards plus three standard deviation.
residuals_10PCT | evaluation(residuals_MEAN*0.1) | 30% of the difference between measured and modeled spectra of the standards.
ch4_ppm_3SIGMA | evaluation(ch4_ppm_STD*3) | Three times the standard deviation of the methanol amplitudes of the standards.
ch4_ppm_PLUS_3SIGMA | evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA) | The mean methanol amplitude of the standards plus three standard deviation.
HALF | float(0.5) | A fixed value of 0.5 or half.
baseline_shift_HALFSIGMA | evaluation(baseline_shift_STD*HALF) | Half the standard deviation of the organic baseline of the standards.
baseline_shift_PLUS_HALFSIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA) | The absolute value of the mean organic baseline of the standards plus one-half standard deviation.
baseline_shift_PLUS_1SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD) | The absolute value of the mean organic baseline of the standards plus one standard deviation.
baseline_shift_3SIGMA | evaluation(baseline_shift_STD*3) | Three times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_3SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA) | The absolute value of the mean organic baseline of the standards plus three standard deviation.
baseline_shift_18SIGMA | evaluation(baseline_shift_STD*18) | Eighteen times the standard deviation of the organic base of the standards.
baseline_shift_PLUS_18SIGMA | evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA) | The absolute value of the mean organic baseline of the standards plus eighteen standard deviation.
slope_shift_5SIGMA | evaluation(slope_shift_STD*5) | Five times the standard deviation of the organic slope of the standards.
slope_shift_PLUS_5SIGMA | evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA) | The absolute value of the mean organic slope of the standards plus five standard deviation.
baseline_curvature_3SIGMA | evaluation(baseline_curvature_STD*3) | Three times the standard deviation of the baseline curvature of the standards.
baseline_curvature_PLUS_3SIGMA | evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
TRUE | evaluation(1>0)
FALSE | evaluation(1<0)
PROC_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the uncalibrated d18O value in order to calibrate it.
PROC_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the uncalibrated dD value in order to calibrate it.
Variable | Instruction | Comment
IS_STANDARD | is_standard_sample() | Standard recognition.
IS_NOT_STANDARD | bool_xor(IS_STANDARD,TRUE)
SAMP_residuals_MEAN | mean_all(residuals) | The mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_residuals_STD | std_all(residuals) | One standard deviation of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN | mean_all(ch4_ppm) | The mean methane concentration of a sample.
SAMP_ch4_ppm_STD | std_all(ch4_ppm) | The standard deviation of the methane concentrations of a sample.
SAMP_baseline_shift_MEAN | mean_all(baseline_shift) | The mean of the organic baseline of a sample.
SAMP_baseline_shift_STD | std_all(baseline_shift) | The standard deviation of the organic baseline of a sample.
SAMP_slope_shift_MEAN | mean_all(slope_shift) | The mean organic slope of a sample.
SAMP_slope_shift_STD | std_all(slope_shift) | The standard deviation of the organic slope of a sample.
SAMP_d18O_STD | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
SAMP_dD_STD | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_baseline_curvature_MEAN | mean_all(baseline_curvature) | The mean of the baseline curvature of the ssamples
SAMP_baseline_curvature_STD | std_all(baseline_curvature) | The standard deviation of the baseline curvature of the samples
SAMP_residuals_MEAN_ABS | abs(SAMP_residuals_MEAN) | Absolute value of the mean difference between the measured spectra of a sample and the modeled spectra of the standard.
SAMP_ch4_ppm_MEAN_ABS | abs(SAMP_ch4_ppm_MEAN) | Absolute value of the mean methane concentration of a sample.
SAMP_baseline_shift_MEAN_ABS | abs(SAMP_baseline_shift_MEAN) | Absolute value of the mean organic baseline of a sample.
SAMP_slope_shift_MEAN_ABS | abs(SAMP_slope_shift_MEAN) | Absolute value of the mean organic slope of a sample.
SAMP_residuals_DIFFERENTIAL | evaluation(SAMP_residuals_MEAN-residuals_MEAN)
SAMP_residuals_DIFFERENTIAL_ABS | abs(SAMP_residuals_DIFFERENTIAL) | Absolute value of the difference between the mean of the sample residual and the mean of all standards residuals
SAMP_residuals_DIFFERENTIAL_PCT | evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
SAMP_baseline_curvature_DIFFERENTIAL | evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
SAMP_STD_DELTA_18_16O_CURR_MEAN | current_standards_mean(d(18_16)Mean) | The d18O mean of the last run standard.
SAMP_STD_DELTA_DH_CURR_MEAN | current_standards_mean(d(D_H)Mean) | The dD mean of the last run standard.
known_dD | known_standard(d(D_H)Mean) | The dD value of the known Standard.
known_d18O | known_standard(d(18_16)Mean) | The d18O value of the known Standard.
d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
measured_d18O_mean | mean_all(d(18_16)Mean) | Mean of all d18O means.
measured_dD_mean | mean_all(d(D_H)Mean) | Mean of all dD means.
measured_d18O_stdev | std_all(d(18_16)Mean) | Standard deviation of all d18O means.
measured_dD_stdev | std_all(d(D_H)Mean) | Standard deviation of all dD means.
SAMP_DELTA_18_16O_CALIBRATE | fit_type1(d(18_16)Mean) | Linear fit run on the mean uncalibrated d18O value in order to calibrate it.
SAMP_DELTA_DH_CALIBRATE | fit_type1(d(D_H)Mean) | Linear fit run on the mean uncalibrated dD value in order to calibrate it.
SAMP_TEST_VALUE | evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
residuals_GT_3SIGMA | evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA) | Threshold of red trigger. If the mean difference between the measured spectra of the sample and the modeled spectra of the standard is greater than that of the mean difference of the standard plus three standard deviations, this will trigger.
residuals_GE_2SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GE_15SIGMA | evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA) | Threshold of yellow trigger. If the mean difference of the measured spectra of the sample and the modeled spectra is greater than or equal to the mean difference of that of the standard, this is triggered.
residuals_GT_10PCT | evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
ch4_ppm_GE_3SIGMA | evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA) | True if the mean methane concentration of the sample is greater than or equal to that of the standard plus three standard deviations, otherwise false.
baseline_shift_GT_1SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus one standard deviation, otherwise false.
baseline_shift_GT_3SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus three standard deviation, otherwise false. Sets yellow trigger for "other", trigger threshold is approximately sample to sample dD precision specification
baseline_shift_GT_18SIGMA | evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA) | True if the mean organic baseline of the sample is greater than that of the standard plus eighteen standard deviation, otherwise false. Sets red trigger for "other", trigger threshold is approximately 24 hour dD drift specification
slope_shift_GT_5SIGMA | evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA) | True if the mean organic slope of the sample is greater than that of the standard plus five standard deviations, otherwise false.
baseline_curvature_GT_3SIGMA | evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
RED | truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA)) | Red trigger. If the sample is not a standard and the threshold is triggered (measured-modeled > standard+3stdev's), then RED.
NOT_RED | bool_xor(RED,TRUE) | If the sample is determined to be "red" than it can not be "not red."
YELLOW | truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA))) | Truth evaluation used to determine whether or not a given sample is yellow.
NOT_YELLOW | bool_xor(YELLOW,TRUE) | If sample is determined to be "yellow" than it can not be "not yellow."
GREEN | truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW)) | This is the pre-requisite for a sample to be colored green. It must not be a standard, must not be yellow, and must not be red.
NOT_GREEN | bool_xor(GREEN,TRUE) | Pre-requisite for being "not-green."
CYAN | truth_evaluation((IS_STANDARD)) | This is the pre-requisite for a sample to be a standard and therefore cyan colored.
NOT_CYAN | bool_xor(CYAN,TRUE) | Pre-requisite for being not a standard.
METHANE | truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA)) | This is the threshold for amounts of methane being included in the main ChemCorrect spreadsheet. Methane will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured methane concentration must be greater than or equal to that of the standard plus three standard deviations.
OTHER1 | truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA)) | This is the threshold for amounts of "OTHER1" being included in the main ChemCorrect spreadsheet. This column (OTHER1) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic baseline must be greater than that of the standard plus one standard deviation.
OTHER2 | truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA)) | This is the threshold for amounts of "OTHER2" being included in the main ChemCorrect spreadsheet. This column (OTHER2) will only be reported to be present if the amount is significant enough to affect the measured d18O and dD values of the water. This line states that the sample must be tagged either red or yellow and the measured organic slope must be greater than that of the standard plus three standard deviations.
OTHER3 | truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | truth_evaluation(RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN)) | Organic baseline differential showing the difference between the measured baseline and the standard baseline.
Variable | Instruction | Comment
DTL_TEST | math_evaluation((1 + 2)) | Test evaluation of 1+2 being equal to 3.
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample | FALSE | FALSE | FALSE | FALSE | Testing Sample
B | identifier | TRUE | Name | RED | YELLOW | GREEN | CYAN | Testing Name
C | SAMP_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d<sup>18</sup>O Mean | RED | YELLOW | GREEN | CYAN | Testing18 16 Mean
D | SAMP_DELTA_DH_CALIBRATE | TRUE | Calibrated d<sup>2</sup>H Mean | RED | YELLOW | GREEN | CYAN | This is a test to see that I can use tool tips
E | SAMP_ch4_ppm_MEAN | METHANE | CH<sub>4</sub> | RED | YELLOW | FALSE | FALSE
F | baseline_shift_DIFF | OTHER1 | C<sub>2+</sub> alcohols | RED | YELLOW | FALSE | FALSE
G | SAMP_residuals_DIFFERENTIAL_PCT | OTHER3 | relative %deviation | RED | YELLOW | FALSE | FALSE
H | SAMP_d18O_STD | TRUE | Uncalibrated d<sup>18</sup>O precision | RED | YELLOW | GREEN | CYAN
I | SAMP_dD_STD | TRUE | Uncalibrated d<sup>2</sup>H precision | RED | YELLOW | GREEN | CYAN
J | SAMP_slope_shift_MEAN | OTHER2 | slope | RED | YELLOW | FALSE | FALSE
K | SAMP_baseline_curvature_DIFFERENTIAL | OTHER4 | curvature | RED | YELLOW | FALSE | FALSE
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | sample | TRUE | Sample
B | identifier | TRUE | Name
C | col(d(18_16)Mean) | TRUE | Measured d(18_16) Mean
D | col(d(D_H)Mean) | TRUE | Measured d(D_H) Mean
E | DTL_DELTA_18_16O_CALIBRATE | TRUE | Calibrated d(18_16) Mean
F | DTL_DELTA_DH_CALIBRATE | TRUE | Calibrated d(D_H) Mean
G | col(ch4_ppm) | METHANE | CH<sub>4</sub>
H | col(baseline_shift) | OTHER1 | Other
I | col(Time Code) | TRUE | Time Code
Variable | Instruction | Print? | Label | Red? | Yellow? | Green? | Cyan? | Comment
A | d18O_mean | TRUE | d(18_16) Mean | RED | YELLOW | GREEN | CYAN
B | dD_mean | TRUE | d(D_H) Mean | RED | YELLOW | GREEN | CYAN
C | XY_X(d18O_mean) XY_Y(dD_mean) | TRUE | d18O x dD | gmwl | best_fit

## Process Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Process Variables
Variable | Value | Operations
baseline_shift_MEAN | 41.5019
varray: [42.144, 41.883, 41.856, 41.65, 42.096, 41.526, 41.695, 41.594, 41.68, 41.289, 41.734, 41.615, 41.058, 41.273, 41.179, 41.062, 41.091, 41.099, 41.301, 41.213]
inst: mean_standards
operation: mean_standards(baseline_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_shift
baseline_shift_STD | 0.344129244387
varray: [42.144, 41.883, 41.856, 41.65, 42.096, 41.526, 41.695, 41.594, 41.68, 41.289, 41.734, 41.615, 41.058, 41.273, 41.179, 41.062, 41.091, 41.099, 41.301, 41.213]
inst: std_standards
operation: std_standards(baseline_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_shift
ch4_ppm_MEAN | -1.0056
varray: [-0.981, -1.134, -1.16, -1.12, -1.14, -1.142, -1.071, -1.005, -0.915, -1.111, -1.065, -1.056, -1.075, -0.915, -0.867, -0.978, -0.756, -0.913, -0.938, -0.77]
inst: mean_standards
operation: mean_standards(ch4_ppm)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: ch4_ppm
ch4_ppm_STD | 0.122323729247
varray: [-0.981, -1.134, -1.16, -1.12, -1.14, -1.142, -1.071, -1.005, -0.915, -1.111, -1.065, -1.056, -1.075, -0.915, -0.867, -0.978, -0.756, -0.913, -0.938, -0.77]
inst: std_standards
operation: std_standards(ch4_ppm)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: ch4_ppm
residuals_MEAN | 2.22325
varray: [2.359, 2.316, 2.327, 2.302, 2.338, 2.334, 2.326, 2.324, 2.345, 2.302, 2.179, 2.147, 2.106, 2.115, 2.109, 2.088, 2.121, 2.111, 2.099, 2.117]
inst: mean_standards
operation: mean_standards(residuals)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: residuals
residuals_STD | 0.108948986131
varray: [2.359, 2.316, 2.327, 2.302, 2.338, 2.334, 2.326, 2.324, 2.345, 2.302, 2.179, 2.147, 2.106, 2.115, 2.109, 2.088, 2.121, 2.111, 2.099, 2.117]
inst: std_standards
operation: std_standards(residuals)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: residuals
slope_shift_MEAN | 30.4391
varray: [31.57, 30.852, 30.793, 30.317, 30.803, 30.461, 30.315, 30.37, 30.555, 30.124, 31.43, 30.977, 30.386, 30.001, 29.859, 29.699, 29.878, 30.069, 30.11, 30.213]
inst: mean_standards
operation: mean_standards(slope_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: slope_shift
slope_shift_STD | 0.503505804075
varray: [31.57, 30.852, 30.793, 30.317, 30.803, 30.461, 30.315, 30.37, 30.555, 30.124, 31.43, 30.977, 30.386, 30.001, 29.859, 29.699, 29.878, 30.069, 30.11, 30.213]
inst: std_standards
operation: std_standards(slope_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: slope_shift
baseline_curvature_MEAN | 64.79045
varray: [67.948, 65.37, 66.304, 65.947, 66.619, 66.623, 66.394, 66.775, 68.148, 65.874, 63.967, 62.721, 61.556, 63.043, 63.154, 62.342, 63.531, 63.348, 62.974, 63.171]
inst: mean_standards
operation: mean_standards(baseline_curvature)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_curvature
baseline_curvature_STD | 2.00435201433
varray: [67.948, 65.37, 66.304, 65.947, 66.619, 66.623, 66.394, 66.775, 68.148, 65.874, 63.967, 62.721, 61.556, 63.043, 63.154, 62.342, 63.531, 63.348, 62.974, 63.171]
inst: std_standards
operation: std_standards(baseline_curvature)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_curvature
residuals_MEAN_ABS | 2.22325
operation: abs(residuals_MEAN)
inst: abs
source_col: residuals_MEAN
baseline_shift_MEAN_ABS | 41.5019
operation: abs(baseline_shift_MEAN)
inst: abs
source_col: baseline_shift_MEAN
ch4_ppm_MEAN_ABS | 1.0056
operation: abs(ch4_ppm_MEAN)
inst: abs
source_col: ch4_ppm_MEAN
slope_shift_MEAN_ABS | 30.4391
operation: abs(slope_shift_MEAN)
inst: abs
source_col: slope_shift_MEAN
residuals_15SIGMA | 0.163423479196
operation: evaluation(residuals_STD*1.5)
inst: evaluation
factors: residuals_STD*1.5
residuals_PLUS_15SIGMA | 2.3866734792
operation: evaluation(residuals_MEAN+residuals_15SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_15SIGMA
residuals_2SIGMA | 0.217897972262
operation: evaluation(residuals_STD*2)
inst: evaluation
factors: residuals_STD*2
residuals_PLUS_2SIGMA | 2.44114797226
operation: evaluation(residuals_MEAN+residuals_2SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_2SIGMA
residuals_3SIGMA | 0.326846958393
operation: evaluation(residuals_STD*3)
inst: evaluation
factors: residuals_STD*3
residuals_PLUS_3SIGMA | 2.55009695839
operation: evaluation(residuals_MEAN+residuals_3SIGMA)
inst: evaluation
factors: residuals_MEAN+residuals_3SIGMA
residuals_10PCT | 0.222325
operation: evaluation(residuals_MEAN*0.1)
inst: evaluation
factors: residuals_MEAN*0.1
ch4_ppm_3SIGMA | 0.36697118774
operation: evaluation(ch4_ppm_STD*3)
inst: evaluation
factors: ch4_ppm_STD*3
ch4_ppm_PLUS_3SIGMA | 1.37257118774
operation: evaluation(ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA)
inst: evaluation
factors: ch4_ppm_MEAN_ABS+ch4_ppm_3SIGMA
HALF | 0.5
operation: float(0.5)
inst: float
baseline_shift_HALFSIGMA | 0.172064622193
operation: evaluation(baseline_shift_STD*HALF)
inst: evaluation
factors: baseline_shift_STD*HALF
baseline_shift_PLUS_HALFSIGMA | 41.6739646222
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_HALFSIGMA
baseline_shift_PLUS_1SIGMA | 41.8460292444
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_STD)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_STD
baseline_shift_3SIGMA | 1.03238773316
operation: evaluation(baseline_shift_STD*3)
inst: evaluation
factors: baseline_shift_STD*3
baseline_shift_PLUS_3SIGMA | 42.5342877332
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_3SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_3SIGMA
baseline_shift_18SIGMA | 6.19432639896
operation: evaluation(baseline_shift_STD*18)
inst: evaluation
factors: baseline_shift_STD*18
baseline_shift_PLUS_18SIGMA | 47.696226399
operation: evaluation(baseline_shift_MEAN_ABS+baseline_shift_18SIGMA)
inst: evaluation
factors: baseline_shift_MEAN_ABS+baseline_shift_18SIGMA
slope_shift_5SIGMA | 2.51752902037
operation: evaluation(slope_shift_STD*5)
inst: evaluation
factors: slope_shift_STD*5
slope_shift_PLUS_5SIGMA | 32.9566290204
operation: evaluation(slope_shift_MEAN_ABS+slope_shift_5SIGMA)
inst: evaluation
factors: slope_shift_MEAN_ABS+slope_shift_5SIGMA
baseline_curvature_3SIGMA | 6.013056043
operation: evaluation(baseline_curvature_STD*3)
inst: evaluation
factors: baseline_curvature_STD*3
baseline_curvature_PLUS_3SIGMA | 70.803506043
operation: evaluation(baseline_curvature_3SIGMA+baseline_curvature_MEAN)
inst: evaluation
factors: baseline_curvature_3SIGMA+baseline_curvature_MEAN
TRUE | True
operation: evaluation(1>0)
inst: evaluation
factors: 1>0
FALSE | False
operation: evaluation(1<0)
inst: evaluation
factors: 1<0
PROC_DELTA_18_16O_CALIBRATE | -7.0033733008
actual_array: [ 1.93 -15.84]
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784, 2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986, -13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42, -9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942, -3.409, -3.136, -3.06, -3.049, -5.718, -5.891, -5.925, -5.914, -7.55, -7.643, -7.621, -7.629, -7.804, -7.786, -7.814, -7.827, -4.792, -4.716, -4.708, -4.695, -6.487, -6.448, -6.455, -6.502, -5.424, -5.373, -5.37, -5.358, -6.142, -6.166, -6.152, -6.117, -8.735, -8.853, -8.911, -8.925, -4.825, -4.569, -4.504, -4.525, -8.415, -8.61, -8.577, -8.63, -3.885, -3.65, -3.558, -3.549, -11.353, -11.759, -11.753, -11.829, 1.047, 1.694, 1.807, 1.87, -3.084, -3.257, -3.304, -3.342, -4.702, -4.839, -4.847, -4.805, -8.53, -8.693, -8.676, -8.731, -8.922, -8.925, -8.9, -8.959, -8.132, -8.036, -8.018, -8.046, -2.724, -2.502, -2.39, -2.394, -5.156, -5.315, -5.235, -5.279, -8.562, -8.63, -8.698, -8.678, -4.769, -4.602, -4.589, -4.58, -5.054, -4.98, -5.048, -4.999, -8.559, -8.678, -8.707, -8.702, -3.844, -3.713, -3.644, -3.609, -8.666, -8.848, -8.924, -8.854, -7.714, -7.62, -7.638, -7.595, -4.95, -4.854, -4.775, -4.758, -0.222, 0.002, 0.017, 0.09, -7.994, -8.327, -8.347, -8.376, -2.897, -2.605, -2.611, -2.596, 1.71, 2.006, 2.012, 2.068, -4.359, -4.858, -4.932, -5.015, -2.888, -2.655, -2.668, -2.591, -8.458, -8.764, -8.81, -8.844, -4.685, -4.503, -4.482, -4.431, 5.069, 5.463, 5.565, 5.638, -7.693, -8.252, -8.414, -8.398, -3.683, -3.333, -3.247, -3.215, 1.149, 1.335, 1.432, 1.465, -5.478, -5.785, -5.836, -5.904, -8.477, -8.585, -8.605, -8.606, -6.273, -6.125, -6.149, -6.129, -8.77, -8.836, -8.893, -8.894, -7.262, -7.155, -7.182, -7.189, -3.695, -3.508, -3.451, -3.425, -5.377, -5.468, -5.492, -5.494, -8.628, -8.686, -8.767, -8.812]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.73693220339
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0, 5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0, 16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0, 21682.0, 22218.0, 22754.0, 23291.0, 23882.0, 24419.0, 24955.0, 25491.0, 26082.0, 26619.0, 27155.0, 27692.0, 28279.0, 28816.0, 29353.0, 29890.0, 30478.0, 31015.0, 31551.0, 32086.0, 32675.0, 33212.0, 33747.0, 34284.0, 34874.0, 35411.0, 35947.0, 36484.0, 37073.0, 37610.0, 38148.0, 38684.0, 39274.0, 39811.0, 40348.0, 40885.0, 41474.0, 42010.0, 42546.0, 43083.0, 43669.0, 44206.0, 44742.0, 45279.0, 45866.0, 46402.0, 46939.0, 47475.0, 48062.0, 48598.0, 49134.0, 49671.0, 50259.0, 50796.0, 51332.0, 51869.0, 52457.0, 52993.0, 53530.0, 54067.0, 54656.0, 55193.0, 55729.0, 56265.0, 56854.0, 57391.0, 57927.0, 58464.0, 59049.0, 59585.0, 60121.0, 60659.0, 61244.0, 61780.0, 62316.0, 62853.0, 63440.0, 63976.0, 64513.0, 65048.0, 65635.0, 66172.0, 66708.0, 67245.0, 67833.0, 68369.0, 68906.0, 69442.0, 70030.0, 70565.0, 71102.0, 71638.0, 72226.0, 72762.0, 73299.0, 73835.0, 74421.0, 74956.0, 75492.0, 76027.0, 76613.0, 77149.0, 77685.0, 78221.0, 78807.0, 79343.0, 79879.0, 80414.0, 81000.0, 81536.0, 82073.0, 82609.0, 83196.0, 83731.0, 84267.0, 84803.0, 85390.0, 85926.0, 86461.0, 86997.0, 87584.0, 88120.0, 88657.0, 89193.0, 89777.0, 90314.0, 90849.0, 91385.0, 91968.0, 92504.0, 93039.0, 93574.0, 94159.0, 94695.0, 95231.0, 95767.0, 96352.0, 96887.0, 97423.0, 97960.0, 98545.0, 99080.0, 99616.0, 100152.0, 100739.0, 101275.0, 101810.0, 102344.0, 102930.0, 103466.0, 104002.0, 104538.0, 105122.0, 105658.0, 106194.0, 106729.0, 107313.0, 107849.0, 108386.0, 108921.0, 109506.0, 110042.0, 110577.0, 111113.0, 111699.0, 112234.0, 112769.0, 113304.0, 113889.0, 114425.0, 114961.0, 115497.0, 116083.0, 116619.0, 117154.0, 117690.0, 118276.0, 118812.0, 119348.0, 119884.0, 120468.0, 121003.0, 121538.0, 122074.0, 122660.0, 123195.0, 123730.0, 124265.0, 124851.0, 125385.0, 125920.0, 126456.0, 127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
PROC_DELTA_DH_CALIBRATE | -67.5986346088
actual_array: [ 18.1 -119. ]
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756, 7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951, -93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888, -72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994, -60.25, -59.653, -59.229, -59.214, -52.055, -51.381, -51.431, -51.275, -61.447, -62.229, -62.591, -62.405, -72.461, -73.45, -73.663, -73.865, -69.431, -69.178, -68.811, -69.065, -66.168, -65.487, -65.264, -65.225, -66.081, -66.077, -66.162, -65.949, -55.83, -54.896, -54.408, -54.053, -65.085, -66.291, -66.534, -66.588, -45.274, -43.04, -42.213, -42.173, -65.61, -68.034, -68.558, -68.832, -61.094, -60.493, -60.415, -60.184, -85.58, -88.369, -88.666, -88.918, -50.222, -46.046, -45.302, -44.877, -41.581, -40.662, -40.495, -40.631, -61.642, -63.869, -64.506, -64.411, -76.829, -78.004, -78.249, -78.432, -68.275, -67.216, -66.736, -66.866, -71.199, -70.941, -71.448, -71.086, -60.298, -59.312, -59.253, -59.014, -63.753, -64.18, -64.163, -64.361, -78.302, -79.146, -79.573, -79.503, -63.815, -62.614, -62.252, -62.157, -62.839, -62.637, -62.531, -62.309, -71.1, -71.573, -71.955, -71.825, -57.627, -56.447, -56.246, -55.952, -65.926, -66.072, -66.714, -66.385, -73.297, -73.743, -73.997, -73.898, -49.862, -47.844, -47.09, -46.802, -52.01, -51.714, -52.215, -51.981, -66.314, -67.141, -67.432, -67.437, -60.253, -59.111, -59.229, -59.064, -50.196, -49.055, -48.95, -48.549, -60.585, -62.235, -62.581, -62.961, -45.261, -43.142, -42.757, -42.407, -63.493, -65.418, -65.924, -66.01, -67.83, -67.681, -67.576, -67.585, -41.145, -38.515, -37.897, -37.541, -63.942, -66.324, -66.937, -66.719, -57.614, -56.503, -55.961, -55.988, -48.454, -47.705, -47.438, -47.428, -56.1, -56.863, -57.09, -57.088, -65.688, -66.606, -66.665, -66.869, -67.278, -67.429, -67.539, -67.65, -66.871, -66.727, -66.816, -66.714, -73.104, -73.558, -73.828, -73.971, -51.049, -48.887, -48.211, -47.954, -50.83, -51.212, -51.14, -51.012, -64.622, -65.48, -66.068, -66.08]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.6072966102
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0, 5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0, 10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0, 16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0, 21682.0, 22218.0, 22754.0, 23291.0, 23882.0, 24419.0, 24955.0, 25491.0, 26082.0, 26619.0, 27155.0, 27692.0, 28279.0, 28816.0, 29353.0, 29890.0, 30478.0, 31015.0, 31551.0, 32086.0, 32675.0, 33212.0, 33747.0, 34284.0, 34874.0, 35411.0, 35947.0, 36484.0, 37073.0, 37610.0, 38148.0, 38684.0, 39274.0, 39811.0, 40348.0, 40885.0, 41474.0, 42010.0, 42546.0, 43083.0, 43669.0, 44206.0, 44742.0, 45279.0, 45866.0, 46402.0, 46939.0, 47475.0, 48062.0, 48598.0, 49134.0, 49671.0, 50259.0, 50796.0, 51332.0, 51869.0, 52457.0, 52993.0, 53530.0, 54067.0, 54656.0, 55193.0, 55729.0, 56265.0, 56854.0, 57391.0, 57927.0, 58464.0, 59049.0, 59585.0, 60121.0, 60659.0, 61244.0, 61780.0, 62316.0, 62853.0, 63440.0, 63976.0, 64513.0, 65048.0, 65635.0, 66172.0, 66708.0, 67245.0, 67833.0, 68369.0, 68906.0, 69442.0, 70030.0, 70565.0, 71102.0, 71638.0, 72226.0, 72762.0, 73299.0, 73835.0, 74421.0, 74956.0, 75492.0, 76027.0, 76613.0, 77149.0, 77685.0, 78221.0, 78807.0, 79343.0, 79879.0, 80414.0, 81000.0, 81536.0, 82073.0, 82609.0, 83196.0, 83731.0, 84267.0, 84803.0, 85390.0, 85926.0, 86461.0, 86997.0, 87584.0, 88120.0, 88657.0, 89193.0, 89777.0, 90314.0, 90849.0, 91385.0, 91968.0, 92504.0, 93039.0, 93574.0, 94159.0, 94695.0, 95231.0, 95767.0, 96352.0, 96887.0, 97423.0, 97960.0, 98545.0, 99080.0, 99616.0, 100152.0, 100739.0, 101275.0, 101810.0, 102344.0, 102930.0, 103466.0, 104002.0, 104538.0, 105122.0, 105658.0, 106194.0, 106729.0, 107313.0, 107849.0, 108386.0, 108921.0, 109506.0, 110042.0, 110577.0, 111113.0, 111699.0, 112234.0, 112769.0, 113304.0, 113889.0, 114425.0, 114961.0, 115497.0, 116083.0, 116619.0, 117154.0, 117690.0, 118276.0, 118812.0, 119348.0, 119884.0, 120468.0, 121003.0, 121538.0, 122074.0, 122660.0, 123195.0, 123730.0, 124265.0, 124851.0, 125385.0, 125920.0, 126456.0, 127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean

## 1 - Sample 1 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 1
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2426
varray: [2.547, 2.197, 2.201, 2.192, 2.214, 2.232, 2.236, 2.209, 2.195, 2.203]
inst: mean_all
operation: mean_all(residuals)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: residuals
SAMP_residuals_STD | 0.107995061615
varray: [2.547, 2.197, 2.201, 2.192, 2.214, 2.232, 2.236, 2.209, 2.195, 2.203]
inst: std_all
operation: std_all(residuals)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0635
varray: [-1.148, -1.195, -0.992, -1.032, -1.078, -0.97, -1.039, -0.889, -1.106, -1.186]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0985181878301
varray: [-1.148, -1.195, -0.992, -1.032, -1.078, -0.97, -1.039, -0.889, -1.106, -1.186]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 53.2085
varray: [111.837, 55.478, 47.186, 47.063, 46.843, 46.513, 45.494, 44.315, 43.768, 43.588]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 20.8756198644
varray: [111.837, 55.478, 47.186, 47.063, 46.843, 46.513, 45.494, 44.315, 43.768, 43.588]
inst: std_all
operation: std_all(baseline_shift)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.9921
varray: [34.293, 29.983, 30.621, 30.63, 30.988, 30.463, 31.008, 30.518, 30.581, 30.836]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.19635617606
varray: [34.293, 29.983, 30.621, 30.63, 30.988, 30.463, 31.008, 30.518, 30.581, 30.836]
inst: std_all
operation: std_all(slope_shift)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: slope_shift
SAMP_d18O_STD | 0.292737443985
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.19583259077
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.3069
varray: [72.062, 64.164, 64.619, 63.386, 64.464, 65.887, 65.141, 65.168, 64.205, 63.973]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.47628630772
varray: [72.062, 64.164, 64.619, 63.386, 64.464, 65.887, 65.141, 65.168, 64.205, 63.973]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2426
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0635
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 53.2085
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.9921
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01935
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01935
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.870347464298
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.51645
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 0
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 0
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9501
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_mean | -67.8931
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
d18O_stdev | 0.292737443985
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
dD_stdev | 2.19583259077
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9501
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.8931
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.292737443985
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.19583259077
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3580925019
actual_array: [ 1.93 -15.84]
varray: [-9.755, -9.058, -8.837, -8.857, -8.846, -8.895, -8.81, -8.822, -8.837, -8.784]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.9501
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.6037351069
actual_array: [ 18.1 -119. ]
varray: [-73.851, -68.957, -67.42, -67.358, -67.007, -67.268, -66.746, -66.708, -66.86, -66.756]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.8931
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [1.0, 539.0, 1075.0, 1612.0, 2149.0, 2686.0, 3222.0, 3758.0, 4295.0, 4832.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2593
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 11.7066
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 2 - Sample 2 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 2
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3273
varray: [2.359, 2.316, 2.327, 2.302, 2.338, 2.334, 2.326, 2.324, 2.345, 2.302]
inst: mean_all
operation: mean_all(residuals)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: residuals
SAMP_residuals_STD | 0.0179446680413
varray: [2.359, 2.316, 2.327, 2.302, 2.338, 2.334, 2.326, 2.324, 2.345, 2.302]
inst: std_all
operation: std_all(residuals)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.0779
varray: [-0.981, -1.134, -1.16, -1.12, -1.14, -1.142, -1.071, -1.005, -0.915, -1.111]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0830052876495
varray: [-0.981, -1.134, -1.16, -1.12, -1.14, -1.142, -1.071, -1.005, -0.915, -1.111]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.7413
varray: [42.144, 41.883, 41.856, 41.65, 42.096, 41.526, 41.695, 41.594, 41.68, 41.289]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.259931208848
varray: [42.144, 41.883, 41.856, 41.65, 42.096, 41.526, 41.695, 41.594, 41.68, 41.289]
inst: std_all
operation: std_all(baseline_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.616
varray: [31.57, 30.852, 30.793, 30.317, 30.803, 30.461, 30.315, 30.37, 30.555, 30.124]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.413983091442
varray: [31.57, 30.852, 30.793, 30.317, 30.803, 30.461, 30.315, 30.37, 30.555, 30.124]
inst: std_all
operation: std_all(slope_shift)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: slope_shift
SAMP_d18O_STD | 0.259611782818
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 3.73841557199
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.6002
varray: [67.948, 65.37, 66.304, 65.947, 66.619, 66.623, 66.394, 66.775, 68.148, 65.874]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.871866427079
varray: [67.948, 65.37, 66.304, 65.947, 66.619, 66.623, 66.394, 66.775, 68.148, 65.874]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3273
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.0779
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.7413
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.616
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.10405
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10405
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.68008546047
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.80975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | 2.8195
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | 17.3909
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
known_dD | 18.1
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 1.93
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 2.8195
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_mean | 17.3909
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
d18O_stdev | 0.259611782818
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
dD_stdev | 3.73841557199
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_mean | 2.8195
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_mean | 17.3909
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.259611782818
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
measured_dD_stdev | 3.73841557199
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.93
actual_array: [ 1.93 -15.84]
varray: [2.133, 2.673, 2.829, 2.861, 2.884, 2.925, 2.953, 2.94, 3.011, 2.986]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.8195
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | 18.1
actual_array: [ 18.1 -119. ]
varray: [7.583, 15.076, 17.295, 18.004, 18.459, 19.137, 19.25, 19.371, 19.783, 19.951]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 17.3909
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [5421.0, 5958.0, 6495.0, 7032.0, 7568.0, 8105.0, 8642.0, 9178.0, 9715.0, 10252.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.1746
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.2394
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 3 - Sample 3 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 3
Variable | Value | Operations
IS_STANDARD | True
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | False
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1192
varray: [2.179, 2.147, 2.106, 2.115, 2.109, 2.088, 2.121, 2.111, 2.099, 2.117]
inst: mean_all
operation: mean_all(residuals)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: residuals
SAMP_residuals_STD | 0.0260375796793
varray: [2.179, 2.147, 2.106, 2.115, 2.109, 2.088, 2.121, 2.111, 2.099, 2.117]
inst: std_all
operation: std_all(residuals)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9333
varray: [-1.065, -1.056, -1.075, -0.915, -0.867, -0.978, -0.756, -0.913, -0.938, -0.77]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.114379533911
varray: [-1.065, -1.056, -1.075, -0.915, -0.867, -0.978, -0.756, -0.913, -0.938, -0.77]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.2625
varray: [41.734, 41.615, 41.058, 41.273, 41.179, 41.062, 41.091, 41.099, 41.301, 41.213]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.234698032942
varray: [41.734, 41.615, 41.058, 41.273, 41.179, 41.062, 41.091, 41.099, 41.301, 41.213]
inst: std_all
operation: std_all(baseline_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.2622
varray: [31.43, 30.977, 30.386, 30.001, 29.859, 29.699, 29.878, 30.069, 30.11, 30.213]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.542477198702
varray: [31.43, 30.977, 30.386, 30.001, 29.859, 29.699, 29.878, 30.069, 30.11, 30.213]
inst: std_all
operation: std_all(slope_shift)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: slope_shift
SAMP_d18O_STD | 0.376482860758
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 5.41509620311
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.9807
varray: [63.967, 62.721, 61.556, 63.043, 63.154, 62.342, 63.531, 63.348, 62.974, 63.171]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.665486630636
varray: [63.967, 62.721, 61.556, 63.043, 63.154, 62.342, 63.531, 63.348, 62.974, 63.171]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1192
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9333
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.2625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.2622
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.10405
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.10405
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.68008546047
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.80975
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
known_dD | -119.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | -15.84
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -14.2007
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
dD_mean | -107.39
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
d18O_stdev | 0.376482860758
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
dD_stdev | 5.41509620311
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
measured_d18O_mean | -14.2007
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
measured_dD_mean | -107.39
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.376482860758
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
measured_dD_stdev | 5.41509620311
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -15.84
actual_array: [ 1.93 -15.84]
varray: [-13.221, -13.964, -14.202, -14.22, -14.277, -14.394, -14.404, -14.449, -14.456, -14.42]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.2007
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -119.0
actual_array: [ 18.1 -119. ]
varray: [-93.149, -104.054, -107.311, -108.263, -109.013, -109.935, -110.138, -110.462, -110.687, -110.888]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -107.39
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [10841.0, 11378.0, 11914.0, 12452.0, 12988.0, 13525.0, 14061.0, 14598.0, 15133.0, 15670.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3827
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | True
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | False
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.2394
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 4 - Sample 4 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 4
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1808
varray: [2.189, 2.184, 2.163, 2.179, 2.203, 2.159, 2.17, 2.182, 2.174, 2.205]
inst: mean_all
operation: mean_all(residuals)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: residuals
SAMP_residuals_STD | 0.0153318839895
varray: [2.189, 2.184, 2.163, 2.179, 2.203, 2.159, 2.17, 2.182, 2.174, 2.205]
inst: std_all
operation: std_all(residuals)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9673
varray: [-0.903, -0.878, -0.955, -0.986, -0.881, -0.935, -1.009, -1.115, -1.04, -0.971]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0745013795547
varray: [-0.903, -0.878, -0.955, -0.986, -0.881, -0.935, -1.009, -1.115, -1.04, -0.971]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.1482
varray: [41.13, 41.174, 41.058, 41.286, 41.311, 40.715, 41.004, 41.17, 41.234, 41.4]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.1927351666
varray: [41.13, 41.174, 41.058, 41.286, 41.311, 40.715, 41.004, 41.17, 41.234, 41.4]
inst: std_all
operation: std_all(baseline_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.1214
varray: [30.233, 30.203, 29.565, 30.083, 30.197, 29.815, 30.139, 30.192, 30.106, 30.681]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.288376374437
varray: [30.233, 30.203, 29.565, 30.083, 30.197, 29.815, 30.139, 30.192, 30.106, 30.681]
inst: std_all
operation: std_all(slope_shift)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: slope_shift
SAMP_d18O_STD | 0.121250933559
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.79494854151
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.5094
varray: [65.112, 65.104, 63.859, 65.043, 65.425, 64.221, 64.15, 63.659, 64.08, 64.441]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.613490414667
varray: [65.112, 65.104, 63.859, 65.043, 65.425, 64.221, 64.15, 63.659, 64.08, 64.441]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1808
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9673
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.1482
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.1214
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.04245
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04245
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.9093669178
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.28105
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -9.0273
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
dD_mean | -68.0666
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
d18O_stdev | 0.121250933559
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
dD_stdev | 1.79494854151
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
measured_d18O_mean | -9.0273
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
measured_dD_mean | -68.0666
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.121250933559
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.79494854151
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.4386934349
actual_array: [ 1.93 -15.84]
varray: [-9.351, -9.085, -9.023, -8.944, -9.004, -8.992, -8.997, -8.96, -8.975, -8.942]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.0273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.7943640413
actual_array: [ 18.1 -119. ]
varray: [-72.894, -69.011, -67.995, -67.433, -67.396, -67.375, -67.196, -67.127, -67.245, -66.994]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.0666
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [16260.0, 16797.0, 17334.0, 17870.0, 18407.0, 18943.0, 19480.0, 20017.0, 20555.0, 21092.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3211
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.3537
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 5 - Sample 5 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 5
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.24525
varray: [2.254, 2.237, 2.234, 2.256]
inst: mean_all
operation: mean_all(residuals)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: residuals
SAMP_residuals_STD | 0.0113541475535
varray: [2.254, 2.237, 2.234, 2.256]
inst: std_all
operation: std_all(residuals)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.2895
varray: [-1.392, -1.213, -1.266, -1.287]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0750932753314
varray: [-1.392, -1.213, -1.266, -1.287]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.14325
varray: [42.449, 42.041, 42.002, 42.081]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.206369207974
varray: [42.449, 42.041, 42.002, 42.081]
inst: std_all
operation: std_all(baseline_shift)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.68375
varray: [30.107, 29.436, 29.461, 29.731]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.312180903751
varray: [30.107, 29.436, 29.461, 29.731]
inst: std_all
operation: std_all(slope_shift)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: slope_shift
SAMP_d18O_STD | 0.168175503567
varray: [-3.409, -3.136, -3.06, -3.049]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.486900742794
varray: [-60.25, -59.653, -59.229, -59.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.92775
varray: [64.738, 64.712, 65.415, 64.846]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.329974115146
varray: [64.738, 64.712, 65.415, 64.846]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.24525
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.2895
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.14325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.68375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.022
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.022
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.989542336669
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.1373
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.409, -3.136, -3.06, -3.049]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-60.25, -59.653, -59.229, -59.214]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.1635
varray: [-3.409, -3.136, -3.06, -3.049]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
dD_mean | -59.5865
varray: [-60.25, -59.653, -59.229, -59.214]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
d18O_stdev | 0.168175503567
varray: [-3.409, -3.136, -3.06, -3.049]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
dD_stdev | 0.486900742794
varray: [-60.25, -59.653, -59.229, -59.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.1635
varray: [-3.409, -3.136, -3.06, -3.049]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.5865
varray: [-60.25, -59.653, -59.229, -59.214]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.168175503567
varray: [-3.409, -3.136, -3.06, -3.049]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.486900742794
varray: [-60.25, -59.653, -59.229, -59.214]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.3165723082
actual_array: [ 1.93 -15.84]
varray: [-3.409, -3.136, -3.06, -3.049]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.1635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.477058989
actual_array: [ 18.1 -119. ]
varray: [-60.25, -59.653, -59.229, -59.214]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.5865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [21682.0, 22218.0, 22754.0, 23291.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.25665
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.64135
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 6 - Sample 6 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 6
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.24275
varray: [2.272, 2.236, 2.258, 2.205]
inst: mean_all
operation: mean_all(residuals)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: residuals
SAMP_residuals_STD | 0.0292047370587
varray: [2.272, 2.236, 2.258, 2.205]
inst: std_all
operation: std_all(residuals)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.98175
varray: [-1.043, -1.055, -0.814, -1.015]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.113082197243
varray: [-1.043, -1.055, -0.814, -1.015]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.44875
varray: [41.862, 41.308, 41.531, 41.094]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.328226928613
varray: [41.862, 41.308, 41.531, 41.094]
inst: std_all
operation: std_all(baseline_shift)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.70625
varray: [31.348, 30.446, 30.671, 30.36]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.447476163239
varray: [31.348, 30.446, 30.671, 30.36]
inst: std_all
operation: std_all(slope_shift)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0970395108534
varray: [-5.718, -5.891, -5.925, -5.914]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.352387570723
varray: [-52.055, -51.381, -51.431, -51.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.72025
varray: [66.268, 65.511, 66.501, 64.601]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.857536539552
varray: [66.268, 65.511, 66.501, 64.601]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.24275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.98175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.44875
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.70625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0195
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0195
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.877094343866
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.9298
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.718, -5.891, -5.925, -5.914]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-52.055, -51.381, -51.431, -51.275]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.862
varray: [-5.718, -5.891, -5.925, -5.914]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
dD_mean | -51.5355
varray: [-52.055, -51.381, -51.431, -51.275]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0970395108534
varray: [-5.718, -5.891, -5.925, -5.914]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
dD_stdev | 0.352387570723
varray: [-52.055, -51.381, -51.431, -51.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.862
varray: [-5.718, -5.891, -5.925, -5.914]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.5355
varray: [-52.055, -51.381, -51.431, -51.275]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0970395108534
varray: [-5.718, -5.891, -5.925, -5.914]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.352387570723
varray: [-52.055, -51.381, -51.431, -51.275]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.13395077614
actual_array: [ 1.93 -15.84]
varray: [-5.718, -5.891, -5.925, -5.914]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.862
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.6312171975
actual_array: [ 18.1 -119. ]
varray: [-52.055, -51.381, -51.431, -51.275]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.5355
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [23882.0, 24419.0, 24955.0, 25491.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.25915
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.05315
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 7 - Sample 7 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 7
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.23
varray: [2.281, 2.194, 2.242, 2.203]
inst: mean_all
operation: mean_all(residuals)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: residuals
SAMP_residuals_STD | 0.0398748040748
varray: [2.281, 2.194, 2.242, 2.203]
inst: std_all
operation: std_all(residuals)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00625
varray: [-1.05, -0.944, -1.048, -0.983]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.051874046176
varray: [-1.05, -0.944, -1.048, -0.983]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.588
varray: [42.189, 41.009, 41.734, 41.42]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.498658868032
varray: [42.189, 41.009, 41.734, 41.42]
inst: std_all
operation: std_all(baseline_shift)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.986
varray: [31.948, 30.212, 31.629, 30.155]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.936043090176
varray: [31.948, 30.212, 31.629, 30.155]
inst: std_all
operation: std_all(slope_shift)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0415080313514
varray: [-7.55, -7.643, -7.621, -7.629]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.502878381586
varray: [-61.447, -62.229, -62.591, -62.405]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.33725
varray: [66.297, 64.884, 65.037, 65.131]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.647879811385
varray: [66.297, 64.884, 65.037, 65.131]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.23
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.588
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.986
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.00675
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.00675
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.303609580569
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.5468
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.55, -7.643, -7.621, -7.629]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-61.447, -62.229, -62.591, -62.405]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.61075
varray: [-7.55, -7.643, -7.621, -7.629]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
dD_mean | -62.168
varray: [-61.447, -62.229, -62.591, -62.405]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0415080313514
varray: [-7.55, -7.643, -7.621, -7.629]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
dD_stdev | 0.502878381586
varray: [-61.447, -62.229, -62.591, -62.405]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.61075
varray: [-7.55, -7.643, -7.621, -7.629]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.168
varray: [-61.447, -62.229, -62.591, -62.405]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0415080313514
varray: [-7.55, -7.643, -7.621, -7.629]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.502878381586
varray: [-61.447, -62.229, -62.591, -62.405]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.95973939789
actual_array: [ 1.93 -15.84]
varray: [-7.55, -7.643, -7.621, -7.629]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.61075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.3134197622
actual_array: [ 18.1 -119. ]
varray: [-61.447, -62.229, -62.591, -62.405]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.168
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [26082.0, 26619.0, 27155.0, 27692.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2719
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.0861
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 8 - Sample 8 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 8
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3855
varray: [2.306, 2.332, 2.422, 2.482]
inst: mean_all
operation: mean_all(residuals)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: residuals
SAMP_residuals_STD | 0.0812957563468
varray: [2.306, 2.332, 2.422, 2.482]
inst: std_all
operation: std_all(residuals)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.57875
varray: [-1.513, -1.569, -1.607, -1.626]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.049828873825
varray: [-1.513, -1.569, -1.607, -1.626]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 44.6815
varray: [43.555, 43.78, 45.221, 46.17]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.23671567738
varray: [43.555, 43.78, 45.221, 46.17]
inst: std_all
operation: std_all(baseline_shift)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.71875
varray: [29.061, 29.613, 31.387, 32.814]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.71342237933
varray: [29.061, 29.613, 31.387, 32.814]
inst: std_all
operation: std_all(slope_shift)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0172892066523
varray: [-7.804, -7.786, -7.814, -7.827]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.62266490453
varray: [-72.461, -73.45, -73.663, -73.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 67.5805
varray: [66.164, 66.312, 67.585, 70.261]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.8974250095
varray: [66.164, 66.312, 67.585, 70.261]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3855
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.57875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 44.6815
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.71875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.16225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.16225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 7.29787473294
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.79005
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.804, -7.786, -7.814, -7.827]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-72.461, -73.45, -73.663, -73.865]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.80775
varray: [-7.804, -7.786, -7.814, -7.827]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
dD_mean | -73.35975
varray: [-72.461, -73.45, -73.663, -73.865]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0172892066523
varray: [-7.804, -7.786, -7.814, -7.827]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
dD_stdev | 0.62266490453
varray: [-72.461, -73.45, -73.663, -73.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.80775
varray: [-7.804, -7.786, -7.814, -7.827]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.35975
varray: [-72.461, -73.45, -73.663, -73.865]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0172892066523
varray: [-7.804, -7.786, -7.814, -7.827]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.62266490453
varray: [-72.461, -73.45, -73.663, -73.865]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.16541794456
actual_array: [ 1.93 -15.84]
varray: [-7.804, -7.786, -7.814, -7.827]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.80775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.6100847566
actual_array: [ 18.1 -119. ]
varray: [-72.461, -73.45, -73.663, -73.865]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.35975
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [28279.0, 28816.0, 29353.0, 29890.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.1164
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 3.1796
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 9 - Sample 9 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 9
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.55275
varray: [2.542, 2.571, 2.567, 2.531]
inst: mean_all
operation: mean_all(residuals)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: residuals
SAMP_residuals_STD | 0.0193627649541
varray: [2.542, 2.571, 2.567, 2.531]
inst: std_all
operation: std_all(residuals)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -3.068
varray: [-3.109, -3.158, -3.096, -2.909]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.109309956851
varray: [-3.109, -3.158, -3.096, -2.909]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 557.278
varray: [556.171, 561.894, 557.361, 553.686]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.43716249252
varray: [556.171, 561.894, 557.361, 553.686]
inst: std_all
operation: std_all(baseline_shift)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 9.04475
varray: [9.144, 9.25, 9.093, 8.692]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.24408929377
varray: [9.144, 9.25, 9.093, 8.692]
inst: std_all
operation: std_all(slope_shift)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0436987795405
varray: [-4.792, -4.716, -4.708, -4.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.257283727948
varray: [-69.431, -69.178, -68.811, -69.065]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.962
varray: [55.808, 56.284, 55.429, 56.327]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.426111878893
varray: [55.808, 56.284, 55.429, 56.327]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.55275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 3.068
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 557.278
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 9.04475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.3295
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.3295
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 14.8206454515
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.82845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.792, -4.716, -4.708, -4.695]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-69.431, -69.178, -68.811, -69.065]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.72775
varray: [-4.792, -4.716, -4.708, -4.695]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
dD_mean | -69.12125
varray: [-69.431, -69.178, -68.811, -69.065]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0436987795405
varray: [-4.792, -4.716, -4.708, -4.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
dD_stdev | 0.257283727948
varray: [-69.431, -69.178, -68.811, -69.065]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.72775
varray: [-4.792, -4.716, -4.708, -4.695]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
measured_dD_mean | -69.12125
varray: [-69.431, -69.178, -68.811, -69.065]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0436987795405
varray: [-4.792, -4.716, -4.708, -4.695]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.257283727948
varray: [-69.431, -69.178, -68.811, -69.065]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.94973305249
actual_array: [ 1.93 -15.84]
varray: [-4.792, -4.716, -4.708, -4.695]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.72775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -76.9531352555
actual_array: [ 18.1 -119. ]
varray: [-69.431, -69.178, -68.811, -69.065]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -69.12125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [30478.0, 31015.0, 31551.0, 32086.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 38.94915
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | True
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | True
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | True
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 515.7761
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 10 - Sample 10 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 10
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.26775
varray: [2.283, 2.278, 2.261, 2.249]
inst: mean_all
operation: mean_all(residuals)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: residuals
SAMP_residuals_STD | 0.0156498136304
varray: [2.283, 2.278, 2.261, 2.249]
inst: std_all
operation: std_all(residuals)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.56325
varray: [-1.426, -1.704, -1.485, -1.638]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.129561761334
varray: [-1.426, -1.704, -1.485, -1.638]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 46.3165
varray: [49.509, 46.662, 44.988, 44.107]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.37754978917
varray: [49.509, 46.662, 44.988, 44.107]
inst: std_all
operation: std_all(baseline_shift)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.6415
varray: [28.646, 28.772, 28.678, 28.47]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.126220706172
varray: [28.646, 28.772, 28.678, 28.47]
inst: std_all
operation: std_all(slope_shift)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0257293606605
varray: [-6.487, -6.448, -6.455, -6.502]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.436856192967
varray: [-66.168, -65.487, -65.264, -65.225]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.60025
varray: [65.385, 64.321, 65.236, 63.459]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.894509688787
varray: [65.385, 64.321, 65.236, 63.459]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.26775
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.56325
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 46.3165
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.6415
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0445
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0445
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.0015742719
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.1902
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-6.487, -6.448, -6.455, -6.502]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-66.168, -65.487, -65.264, -65.225]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.473
varray: [-6.487, -6.448, -6.455, -6.502]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
dD_mean | -65.536
varray: [-66.168, -65.487, -65.264, -65.225]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0257293606605
varray: [-6.487, -6.448, -6.455, -6.502]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
dD_stdev | 0.436856192967
varray: [-66.168, -65.487, -65.264, -65.225]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.473
varray: [-6.487, -6.448, -6.455, -6.502]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.536
varray: [-66.168, -65.487, -65.264, -65.225]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0257293606605
varray: [-6.487, -6.448, -6.455, -6.502]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.436856192967
varray: [-66.168, -65.487, -65.264, -65.225]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.77186748687
actual_array: [ 1.93 -15.84]
varray: [-6.487, -6.448, -6.455, -6.502]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.473
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.0139284137
actual_array: [ 18.1 -119. ]
varray: [-66.168, -65.487, -65.264, -65.225]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.536
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [32675.0, 33212.0, 33747.0, 34284.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.23415
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.8146
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 11 - Sample 11 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 11
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.201
varray: [2.24, 2.18, 2.193, 2.191]
inst: mean_all
operation: mean_all(residuals)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: residuals
SAMP_residuals_STD | 0.0266207938775
varray: [2.24, 2.18, 2.193, 2.191]
inst: std_all
operation: std_all(residuals)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.33825
varray: [-1.342, -1.275, -1.247, -1.489]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.108115293399
varray: [-1.342, -1.275, -1.247, -1.489]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 168.61275
varray: [170.908, 166.666, 167.803, 169.074]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.8190157366
varray: [170.908, 166.666, 167.803, 169.074]
inst: std_all
operation: std_all(baseline_shift)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 24.639
varray: [25.469, 24.096, 24.46, 24.531]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.585221895239
varray: [25.469, 24.096, 24.46, 24.531]
inst: std_all
operation: std_all(slope_shift)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0292275554914
varray: [-5.424, -5.373, -5.37, -5.358]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.088024144415
varray: [-66.081, -66.077, -66.162, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.93675
varray: [61.707, 60.371, 60.904, 60.765]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.560930402694
varray: [61.707, 60.371, 60.904, 60.765]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.201
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.33825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 168.61275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 24.639
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.02225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.00078713595
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.8537
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.424, -5.373, -5.37, -5.358]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-66.081, -66.077, -66.162, -65.949]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.38125
varray: [-5.424, -5.373, -5.37, -5.358]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
dD_mean | -66.06725
varray: [-66.081, -66.077, -66.162, -65.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0292275554914
varray: [-5.424, -5.373, -5.37, -5.358]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
dD_stdev | 0.088024144415
varray: [-66.081, -66.077, -66.162, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.38125
varray: [-5.424, -5.373, -5.37, -5.358]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.06725
varray: [-66.081, -66.077, -66.162, -65.949]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0292275554914
varray: [-5.424, -5.373, -5.37, -5.358]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.088024144415
varray: [-66.081, -66.077, -66.162, -65.949]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.63202203852
actual_array: [ 1.93 -15.84]
varray: [-5.424, -5.373, -5.37, -5.358]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.38125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.5976265198
actual_array: [ 18.1 -119. ]
varray: [-66.081, -66.077, -66.162, -65.949]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.06725
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [34874.0, 35411.0, 35947.0, 36484.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3009
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 127.11085
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 12 - Sample 12 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 12
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.107
varray: [2.13, 2.103, 2.119, 2.076]
inst: mean_all
operation: mean_all(residuals)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: residuals
SAMP_residuals_STD | 0.0234520787991
varray: [2.13, 2.103, 2.119, 2.076]
inst: std_all
operation: std_all(residuals)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.012
varray: [-1.06, -0.9, -1.106, -0.982]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0905243981845
varray: [-1.06, -0.9, -1.106, -0.982]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.1195
varray: [42.502, 40.784, 40.877, 40.315]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.953920506821
varray: [42.502, 40.784, 40.877, 40.315]
inst: std_all
operation: std_all(baseline_shift)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.359
varray: [28.915, 27.817, 28.981, 27.723]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.681733085599
varray: [28.915, 27.817, 28.981, 27.723]
inst: std_all
operation: std_all(slope_shift)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0206619618301
varray: [-6.142, -6.166, -6.152, -6.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.770659187536
varray: [-55.83, -54.896, -54.408, -54.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.505
varray: [62.989, 63.668, 62.027, 61.336]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.02986568704
varray: [62.989, 63.668, 62.027, 61.336]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.107
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.012
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.1195
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.359
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11625
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11625
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.22883166535
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.28545
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-6.142, -6.166, -6.152, -6.117]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-55.83, -54.896, -54.408, -54.053]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.14425
varray: [-6.142, -6.166, -6.152, -6.117]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
dD_mean | -54.79675
varray: [-55.83, -54.896, -54.408, -54.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0206619618301
varray: [-6.142, -6.166, -6.152, -6.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
dD_stdev | 0.770659187536
varray: [-55.83, -54.896, -54.408, -54.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.14425
varray: [-6.142, -6.166, -6.152, -6.117]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
measured_dD_mean | -54.79675
varray: [-55.83, -54.896, -54.408, -54.053]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0206619618301
varray: [-6.142, -6.166, -6.152, -6.117]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.770659187536
varray: [-55.83, -54.896, -54.408, -54.053]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.42863488678
actual_array: [ 1.93 -15.84]
varray: [-6.142, -6.166, -6.152, -6.117]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.14425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -61.2144368649
actual_array: [ 18.1 -119. ]
varray: [-55.83, -54.896, -54.408, -54.053]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -54.79675
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [37073.0, 37610.0, 38148.0, 38684.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3949
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.3824
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 13 - Sample 13 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 13
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.158
varray: [2.212, 2.154, 2.134, 2.132]
inst: mean_all
operation: mean_all(residuals)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: residuals
SAMP_residuals_STD | 0.0373452361978
varray: [2.212, 2.154, 2.134, 2.132]
inst: std_all
operation: std_all(residuals)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89
varray: [-0.89, -0.893, -0.947, -0.83]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0478121323515
varray: [-0.89, -0.893, -0.947, -0.83]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.013
varray: [41.585, 41.101, 40.778, 40.588]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.436187268651
varray: [41.585, 41.101, 40.778, 40.588]
inst: std_all
operation: std_all(baseline_shift)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.4125
varray: [30.241, 29.289, 28.932, 29.188]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.572405741877
varray: [30.241, 29.289, 28.932, 29.188]
inst: std_all
operation: std_all(slope_shift)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0864792846101
varray: [-8.735, -8.853, -8.911, -8.925]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.704936167323
varray: [-65.085, -66.291, -66.534, -66.588]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.1215
varray: [65.587, 63.83, 63.902, 63.167]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.03148969942
varray: [65.587, 63.83, 63.902, 63.167]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.158
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.013
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.4125
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.93489261217
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.66895
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.735, -8.853, -8.911, -8.925]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-65.085, -66.291, -66.534, -66.588]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.856
varray: [-8.735, -8.853, -8.911, -8.925]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
dD_mean | -66.1245
varray: [-65.085, -66.291, -66.534, -66.588]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0864792846101
varray: [-8.735, -8.853, -8.911, -8.925]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
dD_stdev | 0.704936167323
varray: [-65.085, -66.291, -66.534, -66.588]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.856
varray: [-8.735, -8.853, -8.911, -8.925]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.1245
varray: [-65.085, -66.291, -66.534, -66.588]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0864792846101
varray: [-8.735, -8.853, -8.911, -8.925]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.704936167323
varray: [-65.085, -66.291, -66.534, -66.588]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2598470641
actual_array: [ 1.93 -15.84]
varray: [-8.735, -8.853, -8.911, -8.925]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.856
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.6605285745
actual_array: [ 18.1 -119. ]
varray: [-65.085, -66.291, -66.534, -66.588]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.1245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [39274.0, 39811.0, 40348.0, 40885.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3439
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.4889
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 14 - Sample 14 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 14
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1315
varray: [2.115, 2.185, 2.109, 2.117]
inst: mean_all
operation: mean_all(residuals)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: residuals
SAMP_residuals_STD | 0.0358282942193
varray: [2.115, 2.185, 2.109, 2.117]
inst: std_all
operation: std_all(residuals)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.8995
varray: [-0.889, -0.986, -0.837, -0.886]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0623992521323
varray: [-0.889, -0.986, -0.837, -0.886]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.01575
varray: [39.956, 40.689, 39.52, 39.898]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.488695116271
varray: [39.956, 40.689, 39.52, 39.898]
inst: std_all
operation: std_all(baseline_shift)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.46575
varray: [28.299, 29.424, 27.84, 28.3]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.67455781813
varray: [28.299, 29.424, 27.84, 28.3]
inst: std_all
operation: std_all(slope_shift)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: slope_shift
SAMP_d18O_STD | 0.14865480147
varray: [-4.825, -4.569, -4.504, -4.525]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.4552747736
varray: [-45.274, -43.04, -42.213, -42.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.08
varray: [62.144, 64.25, 63.02, 62.906]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.871564111239
varray: [62.144, 64.25, 63.02, 62.906]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1315
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.8995
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.01575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.46575
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.12684133588
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.71045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.825, -4.569, -4.504, -4.525]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-45.274, -43.04, -42.213, -42.173]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.60575
varray: [-4.825, -4.569, -4.504, -4.525]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
dD_mean | -43.175
varray: [-45.274, -43.04, -42.213, -42.173]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
d18O_stdev | 0.14865480147
varray: [-4.825, -4.569, -4.504, -4.525]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
dD_stdev | 1.4552747736
varray: [-45.274, -43.04, -42.213, -42.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.60575
varray: [-4.825, -4.569, -4.504, -4.525]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.175
varray: [-45.274, -43.04, -42.213, -42.173]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.14865480147
varray: [-4.825, -4.569, -4.504, -4.525]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.4552747736
varray: [-45.274, -43.04, -42.213, -42.173]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.82235852105
actual_array: [ 1.93 -15.84]
varray: [-4.825, -4.569, -4.504, -4.525]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.60575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.4453197565
actual_array: [ 18.1 -119. ]
varray: [-45.274, -43.04, -42.213, -42.173]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -43.175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [41474.0, 42010.0, 42546.0, 43083.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3704
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.48615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 15 - Sample 15 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 15
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.13625
varray: [2.165, 2.168, 2.092, 2.12]
inst: mean_all
operation: mean_all(residuals)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: residuals
SAMP_residuals_STD | 0.0367729520164
varray: [2.165, 2.168, 2.092, 2.12]
inst: std_all
operation: std_all(residuals)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.06125
varray: [-1.055, -1.057, -1.087, -1.046]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0178208679175
varray: [-1.055, -1.057, -1.087, -1.046]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.655
varray: [40.946, 41.124, 40.143, 40.407]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.457635225917
varray: [40.946, 41.124, 40.143, 40.407]
inst: std_all
operation: std_all(baseline_shift)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.4095
varray: [29.98, 30.049, 28.488, 29.121]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.745391396069
varray: [29.98, 30.049, 28.488, 29.121]
inst: std_all
operation: std_all(slope_shift)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0978059302906
varray: [-8.415, -8.61, -8.577, -8.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.47009693558
varray: [-65.61, -68.034, -68.558, -68.832]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.192
varray: [63.959, 63.798, 62.791, 62.22]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.828876749986
varray: [63.959, 63.798, 62.791, 62.22]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.13625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.06125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.655
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.4095
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.087
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.087
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.91319014956
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.59845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.415, -8.61, -8.577, -8.63]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-65.61, -68.034, -68.558, -68.832]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.558
varray: [-8.415, -8.61, -8.577, -8.63]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
dD_mean | -67.7585
varray: [-65.61, -68.034, -68.558, -68.832]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0978059302906
varray: [-8.415, -8.61, -8.577, -8.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
dD_stdev | 1.47009693558
varray: [-65.61, -68.034, -68.558, -68.832]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.558
varray: [-8.415, -8.61, -8.577, -8.63]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.7585
varray: [-65.61, -68.034, -68.558, -68.832]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0978059302906
varray: [-8.415, -8.61, -8.577, -8.63]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.47009693558
varray: [-65.61, -68.034, -68.558, -68.832]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.94871911023
actual_array: [ 1.93 -15.84]
varray: [-8.415, -8.61, -8.577, -8.63]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.558
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.4558466079
actual_array: [ 18.1 -119. ]
varray: [-65.61, -68.034, -68.558, -68.832]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.7585
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [43669.0, 44206.0, 44742.0, 45279.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.36565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8469
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 16 - Sample 16 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 16
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.138
varray: [2.173, 2.138, 2.115, 2.126]
inst: mean_all
operation: mean_all(residuals)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: residuals
SAMP_residuals_STD | 0.0251528659732
varray: [2.173, 2.138, 2.115, 2.126]
inst: std_all
operation: std_all(residuals)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.25
varray: [-1.203, -1.137, -1.292, -1.368]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.101103906947
varray: [-1.203, -1.137, -1.292, -1.368]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 127.0565
varray: [127.682, 126.835, 127.125, 126.584]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.471967866138
varray: [127.682, 126.835, 127.125, 126.584]
inst: std_all
operation: std_all(baseline_shift)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.04375
varray: [25.217, 25.03, 24.834, 25.094]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.159921595373
varray: [25.217, 25.03, 24.834, 25.094]
inst: std_all
operation: std_all(slope_shift)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: slope_shift
SAMP_d18O_STD | 0.156470444494
varray: [-3.885, -3.65, -3.558, -3.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.38786466712
varray: [-61.094, -60.493, -60.415, -60.184]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.4395
varray: [61.422, 60.577, 59.988, 59.771]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.73823144519
varray: [61.422, 60.577, 59.988, 59.771]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.138
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.25
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 127.0565
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.04375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.83447655459
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.35095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.885, -3.65, -3.558, -3.549]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-61.094, -60.493, -60.415, -60.184]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.6605
varray: [-3.885, -3.65, -3.558, -3.549]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
dD_mean | -60.5465
varray: [-61.094, -60.493, -60.415, -60.184]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
d18O_stdev | 0.156470444494
varray: [-3.885, -3.65, -3.558, -3.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
dD_stdev | 0.38786466712
varray: [-61.094, -60.493, -60.415, -60.184]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.6605
varray: [-3.885, -3.65, -3.558, -3.549]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
measured_dD_mean | -60.5465
varray: [-61.094, -60.493, -60.415, -60.184]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.156470444494
varray: [-3.885, -3.65, -3.558, -3.549]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.38786466712
varray: [-61.094, -60.493, -60.415, -60.184]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.83546691578
actual_array: [ 1.93 -15.84]
varray: [-3.885, -3.65, -3.558, -3.549]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.6605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -67.5318358018
actual_array: [ 18.1 -119. ]
varray: [-61.094, -60.493, -60.415, -60.184]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.5465
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [45866.0, 46402.0, 46939.0, 47475.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3639
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 85.5546
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 17 - Sample 17 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 17
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.06575
varray: [2.103, 2.039, 2.088, 2.033]
inst: mean_all
operation: mean_all(residuals)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: residuals
SAMP_residuals_STD | 0.0349797560502
varray: [2.103, 2.039, 2.088, 2.033]
inst: std_all
operation: std_all(residuals)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9865
varray: [-1.122, -0.885, -0.982, -0.957]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0992522040058
varray: [-1.122, -0.885, -0.982, -0.957]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.54525
varray: [42.035, 39.903, 40.548, 39.695]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.05747698951
varray: [42.035, 39.903, 40.548, 39.695]
inst: std_all
operation: std_all(baseline_shift)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.49075
varray: [29.581, 27.854, 28.636, 27.892]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.811109271307
varray: [29.581, 27.854, 28.636, 27.892]
inst: std_all
operation: std_all(slope_shift)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: slope_shift
SAMP_d18O_STD | 0.216433977616
varray: [-11.353, -11.759, -11.753, -11.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.55180741395
varray: [-85.58, -88.369, -88.666, -88.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.52225
varray: [62.247, 60.875, 62.842, 60.125]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.24341502189
varray: [62.247, 60.875, 62.842, 60.125]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.06575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9865
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.54525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.49075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.08422354661
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.2682
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-11.353, -11.759, -11.753, -11.829]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-85.58, -88.369, -88.666, -88.918]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -11.6735
varray: [-11.353, -11.759, -11.753, -11.829]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
dD_mean | -87.88325
varray: [-85.58, -88.369, -88.666, -88.918]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
d18O_stdev | 0.216433977616
varray: [-11.353, -11.759, -11.753, -11.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
dD_stdev | 1.55180741395
varray: [-85.58, -88.369, -88.666, -88.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
measured_d18O_mean | -11.6735
varray: [-11.353, -11.759, -11.753, -11.829]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
measured_dD_mean | -87.88325
varray: [-85.58, -88.369, -88.666, -88.918]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.216433977616
varray: [-11.353, -11.759, -11.753, -11.829]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.55180741395
varray: [-85.58, -88.369, -88.666, -88.918]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -13.2014679028
actual_array: [ 1.93 -15.84]
varray: [-11.353, -11.759, -11.753, -11.829]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -11.6735
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -97.5674295906
actual_array: [ 18.1 -119. ]
varray: [-85.58, -88.369, -88.666, -88.918]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -87.88325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [48062.0, 48598.0, 49134.0, 49671.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.43615
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.95665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 18 - Sample 18 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 18
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1875
varray: [2.234, 2.161, 2.166, 2.189]
inst: mean_all
operation: mean_all(residuals)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: residuals
SAMP_residuals_STD | 0.0333116596204
varray: [2.234, 2.161, 2.166, 2.189]
inst: std_all
operation: std_all(residuals)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.211
varray: [-1.09, -1.378, -1.208, -1.168]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121638809596
varray: [-1.09, -1.378, -1.208, -1.168]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 49.757
varray: [50.542, 49.593, 49.337, 49.556]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.535388332584
varray: [50.542, 49.593, 49.337, 49.556]
inst: std_all
operation: std_all(baseline_shift)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.49175
varray: [29.338, 28.294, 27.914, 28.421]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.603888717673
varray: [29.338, 28.294, 27.914, 28.421]
inst: std_all
operation: std_all(slope_shift)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: slope_shift
SAMP_d18O_STD | 0.378731655221
varray: [1.047, 1.694, 1.807, 1.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 2.45484424149
varray: [-50.222, -46.046, -45.302, -44.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.375
varray: [65.25, 61.325, 63.249, 63.676]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.61499246644
varray: [65.25, 61.325, 63.249, 63.676]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.211
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 49.757
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.49175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.03575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.03575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.60800629709
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.41545
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [1.047, 1.694, 1.807, 1.87]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-50.222, -46.046, -45.302, -44.877]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.6045
varray: [1.047, 1.694, 1.807, 1.87]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
dD_mean | -46.61175
varray: [-50.222, -46.046, -45.302, -44.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
d18O_stdev | 0.378731655221
varray: [1.047, 1.694, 1.807, 1.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
dD_stdev | 2.45484424149
varray: [-50.222, -46.046, -45.302, -44.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.6045
varray: [1.047, 1.694, 1.807, 1.87]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
measured_dD_mean | -46.61175
varray: [-50.222, -46.046, -45.302, -44.877]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.378731655221
varray: [1.047, 1.694, 1.807, 1.87]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
measured_dD_stdev | 2.45484424149
varray: [-50.222, -46.046, -45.302, -44.877]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.661474953291
actual_array: [ 1.93 -15.84]
varray: [1.047, 1.694, 1.807, 1.87]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.6045
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -52.22136581
actual_array: [ 18.1 -119. ]
varray: [-50.222, -46.046, -45.302, -44.877]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -46.61175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [50259.0, 50796.0, 51332.0, 51869.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3144
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 8.2551
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 19 - Sample 19 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 19
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1115
varray: [2.1, 2.115, 2.112, 2.119]
inst: mean_all
operation: mean_all(residuals)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: residuals
SAMP_residuals_STD | 0.00818535277187
varray: [2.1, 2.115, 2.112, 2.119]
inst: std_all
operation: std_all(residuals)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.95975
varray: [-1.103, -0.914, -0.799, -1.023]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.132230543622
varray: [-1.103, -0.914, -0.799, -1.023]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.70325
varray: [39.88, 39.572, 39.705, 39.656]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.130003525593
varray: [39.88, 39.572, 39.705, 39.656]
inst: std_all
operation: std_all(baseline_shift)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.24525
varray: [28.241, 28.174, 28.271, 28.295]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0523855896216
varray: [28.241, 28.174, 28.271, 28.295]
inst: std_all
operation: std_all(slope_shift)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: slope_shift
SAMP_d18O_STD | 0.113933825823
varray: [-3.084, -3.257, -3.304, -3.342]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.497812129891
varray: [-41.581, -40.662, -40.495, -40.631]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.35025
varray: [61.811, 62.851, 62.682, 62.057]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.495847674863
varray: [61.811, 62.851, 62.682, 62.057]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1115
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.95975
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.70325
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.24525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.11175
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.11175
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.02642527831
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.4402
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.084, -3.257, -3.304, -3.342]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-41.581, -40.662, -40.495, -40.631]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.24675
varray: [-3.084, -3.257, -3.304, -3.342]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
dD_mean | -40.84225
varray: [-41.581, -40.662, -40.495, -40.631]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
d18O_stdev | 0.113933825823
varray: [-3.084, -3.257, -3.304, -3.342]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
dD_stdev | 0.497812129891
varray: [-41.581, -40.662, -40.495, -40.631]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.24675
varray: [-3.084, -3.257, -3.304, -3.342]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
measured_dD_mean | -40.84225
varray: [-41.581, -40.662, -40.495, -40.631]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.113933825823
varray: [-3.084, -3.257, -3.304, -3.342]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.497812129891
varray: [-41.581, -40.662, -40.495, -40.631]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.4034897651
actual_array: [ 1.93 -15.84]
varray: [-3.084, -3.257, -3.304, -3.342]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.24675
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -45.8822670377
actual_array: [ 18.1 -119. ]
varray: [-41.581, -40.662, -40.495, -40.631]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -40.84225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [52457.0, 52993.0, 53530.0, 54067.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3904
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.79865
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 20 - Sample 20 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 20
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.21175
varray: [2.244, 2.195, 2.217, 2.191]
inst: mean_all
operation: mean_all(residuals)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: residuals
SAMP_residuals_STD | 0.0243498802187
varray: [2.244, 2.195, 2.217, 2.191]
inst: std_all
operation: std_all(residuals)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.46925
varray: [-1.734, -1.34, -1.48, -1.323]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.190002412265
varray: [-1.734, -1.34, -1.48, -1.323]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 268.60175
varray: [272.56, 265.883, 270.474, 265.49]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.47599639192
varray: [272.56, 265.883, 270.474, 265.49]
inst: std_all
operation: std_all(baseline_shift)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.8765
varray: [21.758, 20.377, 21.095, 20.276]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.691588750631
varray: [21.758, 20.377, 21.095, 20.276]
inst: std_all
operation: std_all(slope_shift)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0667001998998
varray: [-4.702, -4.839, -4.847, -4.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.33971215814
varray: [-61.642, -63.869, -64.506, -64.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 59.107
varray: [59.004, 59.046, 59.312, 59.066]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.139087502434
varray: [59.004, 59.046, 59.312, 59.066]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.21175
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.46925
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 268.60175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.8765
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.517260766895
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -5.68345
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.702, -4.839, -4.847, -4.805]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-61.642, -63.869, -64.506, -64.411]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.79825
varray: [-4.702, -4.839, -4.847, -4.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
dD_mean | -63.607
varray: [-61.642, -63.869, -64.506, -64.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0667001998998
varray: [-4.702, -4.839, -4.847, -4.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
dD_stdev | 1.33971215814
varray: [-61.642, -63.869, -64.506, -64.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.79825
varray: [-4.702, -4.839, -4.847, -4.805]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
measured_dD_mean | -63.607
varray: [-61.642, -63.869, -64.506, -64.411]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0667001998998
varray: [-4.702, -4.839, -4.847, -4.805]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.33971215814
varray: [-61.642, -63.869, -64.506, -64.411]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.02333882681
actual_array: [ 1.93 -15.84]
varray: [-4.702, -4.839, -4.847, -4.805]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.79825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -70.8944862555
actual_array: [ 18.1 -119. ]
varray: [-61.642, -63.869, -64.506, -64.411]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.607
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [54656.0, 55193.0, 55729.0, 56265.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.29015
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 227.09985
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 21 - Sample 21 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 21
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.18025
varray: [2.296, 2.156, 2.12, 2.149]
inst: mean_all
operation: mean_all(residuals)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: residuals
SAMP_residuals_STD | 0.0787247313958
varray: [2.296, 2.156, 2.12, 2.149]
inst: std_all
operation: std_all(residuals)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.596
varray: [-1.682, -1.523, -1.617, -1.562]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0690941386805
varray: [-1.682, -1.523, -1.617, -1.562]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 401.701
varray: [421.791, 398.317, 389.799, 396.897]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 13.901985182
varray: [421.791, 398.317, 389.799, 396.897]
inst: std_all
operation: std_all(baseline_shift)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 16.523
varray: [17.555, 16.201, 15.752, 16.584]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.767435122122
varray: [17.555, 16.201, 15.752, 16.584]
inst: std_all
operation: std_all(slope_shift)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0880549071129
varray: [-8.53, -8.693, -8.676, -8.731]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.721302756592
varray: [-76.829, -78.004, -78.249, -78.432]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 55.81425
varray: [58.203, 55.661, 54.557, 54.836]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.66004646019
varray: [58.203, 55.661, 54.557, 54.836]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.18025
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.596
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 401.701
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 16.523
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.043
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.043
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -1.93410547622
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -8.9762
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.53, -8.693, -8.676, -8.731]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-76.829, -78.004, -78.249, -78.432]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.6575
varray: [-8.53, -8.693, -8.676, -8.731]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
dD_mean | -77.8785
varray: [-76.829, -78.004, -78.249, -78.432]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0880549071129
varray: [-8.53, -8.693, -8.676, -8.731]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
dD_stdev | 0.721302756592
varray: [-76.829, -78.004, -78.249, -78.432]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.6575
varray: [-8.53, -8.693, -8.676, -8.731]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
measured_dD_mean | -77.8785
varray: [-76.829, -78.004, -78.249, -78.432]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0880549071129
varray: [-8.53, -8.693, -8.676, -8.731]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.721302756592
varray: [-76.829, -78.004, -78.249, -78.432]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0526024371
actual_array: [ 1.93 -15.84]
varray: [-8.53, -8.693, -8.676, -8.731]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.6575
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -86.5749521762
actual_array: [ 18.1 -119. ]
varray: [-76.829, -78.004, -78.249, -78.432]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -77.8785
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [56854.0, 57391.0, 57927.0, 58464.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.32165
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 360.1991
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 22 - Sample 22 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 22
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16575
varray: [2.406, 2.101, 2.089, 2.067]
inst: mean_all
operation: mean_all(residuals)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: residuals
SAMP_residuals_STD | 0.160784275765
varray: [2.406, 2.101, 2.089, 2.067]
inst: std_all
operation: std_all(residuals)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9425
varray: [-0.953, -0.951, -0.808, -1.058]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.102653137637
varray: [-0.953, -0.951, -0.808, -1.058]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.139
varray: [48.965, 42.09, 40.977, 40.524]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.9393377616
varray: [48.965, 42.09, 40.977, 40.524]
inst: std_all
operation: std_all(baseline_shift)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.74775
varray: [34.171, 28.625, 27.914, 28.281]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.96308975846
varray: [34.171, 28.625, 27.914, 28.281]
inst: std_all
operation: std_all(slope_shift)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0243652758381
varray: [-8.922, -8.925, -8.9, -8.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.697918512435
varray: [-68.275, -67.216, -66.736, -66.866]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.0995
varray: [69.832, 62.028, 63.364, 61.174]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.92649882211
varray: [69.832, 62.028, 63.364, 61.174]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16575
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.139
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.74775
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0575
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0575
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.58630383448
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.69095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.922, -8.925, -8.9, -8.959]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-68.275, -67.216, -66.736, -66.866]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.9265
varray: [-8.922, -8.925, -8.9, -8.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
dD_mean | -67.27325
varray: [-68.275, -67.216, -66.736, -66.866]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0243652758381
varray: [-8.922, -8.925, -8.9, -8.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
dD_stdev | 0.697918512435
varray: [-68.275, -67.216, -66.736, -66.866]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.9265
varray: [-8.922, -8.925, -8.9, -8.959]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.27325
varray: [-68.275, -67.216, -66.736, -66.866]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0243652758381
varray: [-8.922, -8.925, -8.9, -8.959]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.697918512435
varray: [-68.275, -67.216, -66.736, -66.866]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.3334528384
actual_array: [ 1.93 -15.84]
varray: [-8.922, -8.925, -8.9, -8.959]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.9265
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.9226898908
actual_array: [ 18.1 -119. ]
varray: [-68.275, -67.216, -66.736, -66.866]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.27325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [59049.0, 59585.0, 60121.0, 60659.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.33615
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.6371
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 23 - Sample 23 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 23
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.25875
varray: [2.556, 2.146, 2.177, 2.156]
inst: mean_all
operation: mean_all(residuals)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: residuals
SAMP_residuals_STD | 0.19858730238
varray: [2.556, 2.146, 2.177, 2.156]
inst: std_all
operation: std_all(residuals)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.3995
varray: [-1.692, -1.309, -1.223, -1.374]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.204571910747
varray: [-1.692, -1.309, -1.223, -1.374]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 284.159
varray: [324.393, 269.445, 272.565, 270.233]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 26.8553586459
varray: [324.393, 269.445, 272.565, 270.233]
inst: std_all
operation: std_all(baseline_shift)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 21.93525
varray: [26.533, 20.371, 20.498, 20.339]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.06593567828
varray: [26.533, 20.371, 20.498, 20.339]
inst: std_all
operation: std_all(slope_shift)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0506754378373
varray: [-8.132, -8.036, -8.018, -8.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.214175161959
varray: [-71.199, -70.941, -71.448, -71.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.69825
varray: [67.212, 58.203, 59.203, 58.175]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.36874408917
varray: [67.212, 58.203, 59.203, 58.175]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.25875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.3995
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 284.159
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 21.93525
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0355
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0355
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.59676149781
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.0922
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.132, -8.036, -8.018, -8.046]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-71.199, -70.941, -71.448, -71.086]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.058
varray: [-8.132, -8.036, -8.018, -8.046]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
dD_mean | -71.1685
varray: [-71.199, -70.941, -71.448, -71.086]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0506754378373
varray: [-8.132, -8.036, -8.018, -8.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
dD_stdev | 0.214175161959
varray: [-71.199, -70.941, -71.448, -71.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.058
varray: [-8.132, -8.036, -8.018, -8.046]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.1685
varray: [-71.199, -70.941, -71.448, -71.086]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0506754378373
varray: [-8.132, -8.036, -8.018, -8.046]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.214175161959
varray: [-71.199, -70.941, -71.448, -71.086]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.42669234204
actual_array: [ 1.93 -15.84]
varray: [-8.132, -8.036, -8.018, -8.046]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.2025017451
actual_array: [ 18.1 -119. ]
varray: [-71.199, -70.941, -71.448, -71.086]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.1685
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [61244.0, 61780.0, 62316.0, 62853.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.24315
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 242.6571
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 24 - Sample 24 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 24
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2705
varray: [2.553, 2.185, 2.172, 2.172]
inst: mean_all
operation: mean_all(residuals)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: residuals
SAMP_residuals_STD | 0.18843301197
varray: [2.553, 2.185, 2.172, 2.172]
inst: std_all
operation: std_all(residuals)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.19425
varray: [-1.184, -1.234, -1.224, -1.135]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0450212912594
varray: [-1.184, -1.234, -1.224, -1.135]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 49.15
varray: [55.296, 47.675, 47.057, 46.572]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.12212218806
varray: [55.296, 47.675, 47.057, 46.572]
inst: std_all
operation: std_all(baseline_shift)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.7755
varray: [34.75, 28.407, 28.031, 27.914]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.32299588725
varray: [34.75, 28.407, 28.031, 27.914]
inst: std_all
operation: std_all(slope_shift)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: slope_shift
SAMP_d18O_STD | 0.156515174983
varray: [-2.724, -2.502, -2.39, -2.394]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.567324348734
varray: [-60.298, -59.312, -59.253, -59.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.78825
varray: [73.232, 63.188, 63.138, 63.595]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.96671872736
varray: [73.232, 63.188, 63.138, 63.595]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2705
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.19425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 49.15
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.7755
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.04725
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.04725
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.12526706398
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.9978
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-2.724, -2.502, -2.39, -2.394]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-60.298, -59.312, -59.253, -59.014]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.5025
varray: [-2.724, -2.502, -2.39, -2.394]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
dD_mean | -59.46925
varray: [-60.298, -59.312, -59.253, -59.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
d18O_stdev | 0.156515174983
varray: [-2.724, -2.502, -2.39, -2.394]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
dD_stdev | 0.567324348734
varray: [-60.298, -59.312, -59.253, -59.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.5025
varray: [-2.724, -2.502, -2.39, -2.394]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.46925
varray: [-60.298, -59.312, -59.253, -59.014]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.156515174983
varray: [-2.724, -2.502, -2.39, -2.394]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.567324348734
varray: [-60.298, -59.312, -59.253, -59.014]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.62645292065
actual_array: [ 1.93 -15.84]
varray: [-2.724, -2.502, -2.39, -2.394]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.5025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.3482333835
actual_array: [ 18.1 -119. ]
varray: [-60.298, -59.312, -59.253, -59.014]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.46925
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [63440.0, 63976.0, 64513.0, 65048.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2314
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.6481
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 25 - Sample 25 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 25
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.25825
varray: [2.586, 2.161, 2.164, 2.122]
inst: mean_all
operation: mean_all(residuals)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: residuals
SAMP_residuals_STD | 0.219335929569
varray: [2.586, 2.161, 2.164, 2.122]
inst: std_all
operation: std_all(residuals)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.189
varray: [-1.189, -1.239, -1.19, -1.138]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0412391399846
varray: [-1.189, -1.239, -1.19, -1.138]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.959
varray: [47.222, 41.733, 41.768, 41.113]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.85788045003
varray: [47.222, 41.733, 41.768, 41.113]
inst: std_all
operation: std_all(baseline_shift)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.7205
varray: [36.708, 28.876, 29.128, 28.17]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: slope_shift
SAMP_slope_shift_STD | 4.01220816176
varray: [36.708, 28.876, 29.128, 28.17]
inst: std_all
operation: std_all(slope_shift)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0684853999039
varray: [-5.156, -5.315, -5.235, -5.279]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.256960924396
varray: [-63.753, -64.18, -64.163, -64.361]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.55375
varray: [74.057, 63.177, 62.46, 62.521]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.67811772656
varray: [74.057, 63.177, 62.46, 62.521]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.25825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.189
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.959
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.7205
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.035
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.035
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.57427189925
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.7633
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.156, -5.315, -5.235, -5.279]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-63.753, -64.18, -64.163, -64.361]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.24625
varray: [-5.156, -5.315, -5.235, -5.279]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
dD_mean | -64.11425
varray: [-63.753, -64.18, -64.163, -64.361]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0684853999039
varray: [-5.156, -5.315, -5.235, -5.279]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
dD_stdev | 0.256960924396
varray: [-63.753, -64.18, -64.163, -64.361]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.24625
varray: [-5.156, -5.315, -5.235, -5.279]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
measured_dD_mean | -64.11425
varray: [-63.753, -64.18, -64.163, -64.361]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0684853999039
varray: [-5.156, -5.315, -5.235, -5.279]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.256960924396
varray: [-63.753, -64.18, -64.163, -64.361]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.49107481111
actual_array: [ 1.93 -15.84]
varray: [-5.156, -5.315, -5.235, -5.279]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.24625
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -71.4518149412
actual_array: [ 18.1 -119. ]
varray: [-63.753, -64.18, -64.163, -64.361]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.11425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [65635.0, 66172.0, 66708.0, 67245.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.24365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.4571
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 26 - Sample 26 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 26
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.3075
varray: [2.649, 2.216, 2.201, 2.164]
inst: mean_all
operation: mean_all(residuals)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: residuals
SAMP_residuals_STD | 0.228713066235
varray: [2.649, 2.216, 2.201, 2.164]
inst: std_all
operation: std_all(residuals)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.7695
varray: [-2.031, -1.675, -1.614, -1.758]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.184051623193
varray: [-2.031, -1.675, -1.614, -1.758]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 418.56
varray: [483.103, 400.646, 397.715, 392.776]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 43.1510483689
varray: [483.103, 400.646, 397.715, 392.776]
inst: std_all
operation: std_all(baseline_shift)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 17.01075
varray: [20.866, 16.265, 15.718, 15.194]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.60709792362
varray: [20.866, 16.265, 15.718, 15.194]
inst: std_all
operation: std_all(slope_shift)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0604869131851
varray: [-8.562, -8.63, -8.698, -8.678]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.583442085101
varray: [-78.302, -79.146, -79.573, -79.503]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 58.1625
varray: [64.797, 56.586, 56.222, 55.045]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.471624947
varray: [64.797, 56.586, 56.222, 55.045]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.3075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.7695
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 418.56
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 17.01075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.08425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 3.78949735747
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -6.62795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.562, -8.63, -8.698, -8.678]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-78.302, -79.146, -79.573, -79.503]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.642
varray: [-8.562, -8.63, -8.698, -8.678]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
dD_mean | -79.131
varray: [-78.302, -79.146, -79.573, -79.503]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0604869131851
varray: [-8.562, -8.63, -8.698, -8.678]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
dD_stdev | 0.583442085101
varray: [-78.302, -79.146, -79.573, -79.503]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.642
varray: [-8.562, -8.63, -8.698, -8.678]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
measured_dD_mean | -79.131
varray: [-78.302, -79.146, -79.573, -79.503]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0604869131851
varray: [-8.562, -8.63, -8.698, -8.678]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.583442085101
varray: [-78.302, -79.146, -79.573, -79.503]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0364196073
actual_array: [ 1.93 -15.84]
varray: [-8.562, -8.63, -8.698, -8.678]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.642
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -87.9511062991
actual_array: [ 18.1 -119. ]
varray: [-78.302, -79.146, -79.573, -79.503]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -79.131
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [67833.0, 68369.0, 68906.0, 69442.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.1944
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 377.0581
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 27 - Sample 27 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 27
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.377
varray: [2.72, 2.301, 2.229, 2.258]
inst: mean_all
operation: mean_all(residuals)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: residuals
SAMP_residuals_STD | 0.230571753112
varray: [2.72, 2.301, 2.229, 2.258]
inst: std_all
operation: std_all(residuals)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.00225
varray: [-2.315, -1.84, -1.895, -1.959]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.214095578345
varray: [-2.315, -1.84, -1.895, -1.959]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 300.139
varray: [346.909, 288.991, 280.568, 284.088]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 31.3707371394
varray: [346.909, 288.991, 280.568, 284.088]
inst: std_all
operation: std_all(baseline_shift)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 20.394
varray: [24.71, 19.556, 18.714, 18.596]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.90891182403
varray: [24.71, 19.556, 18.714, 18.596]
inst: std_all
operation: std_all(slope_shift)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0897886407069
varray: [-4.769, -4.602, -4.589, -4.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.762848390355
varray: [-63.815, -62.614, -62.252, -62.157]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.72
varray: [68.211, 58.838, 57.292, 58.539]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.03868203667
varray: [68.211, 58.838, 57.292, 58.539]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.377
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.00225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 300.139
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 20.394
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.15375
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.15375
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 6.9155515574
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.07045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.769, -4.602, -4.589, -4.58]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-63.815, -62.614, -62.252, -62.157]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.635
varray: [-4.769, -4.602, -4.589, -4.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
dD_mean | -62.7095
varray: [-63.815, -62.614, -62.252, -62.157]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0897886407069
varray: [-4.769, -4.602, -4.589, -4.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
dD_stdev | 0.762848390355
varray: [-63.815, -62.614, -62.252, -62.157]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.635
varray: [-4.769, -4.602, -4.589, -4.58]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.7095
varray: [-63.815, -62.614, -62.252, -62.157]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0897886407069
varray: [-4.769, -4.602, -4.589, -4.58]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.762848390355
varray: [-63.815, -62.614, -62.252, -62.157]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.85289708699
actual_array: [ 1.93 -15.84]
varray: [-4.769, -4.602, -4.589, -4.58]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.635
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.9083798081
actual_array: [ 18.1 -119. ]
varray: [-63.815, -62.614, -62.252, -62.157]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.7095
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [70030.0, 70565.0, 71102.0, 71638.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.1249
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 258.6371
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 28 - Sample 28 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 28
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.23825
varray: [2.606, 2.146, 2.103, 2.098]
inst: mean_all
operation: mean_all(residuals)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: residuals
SAMP_residuals_STD | 0.24611159393
varray: [2.606, 2.146, 2.103, 2.098]
inst: std_all
operation: std_all(residuals)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.47525
varray: [-1.604, -1.427, -1.336, -1.534]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.117964330767
varray: [-1.604, -1.427, -1.336, -1.534]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 44.443
varray: [51.321, 42.974, 41.83, 41.647]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.62277802481
varray: [51.321, 42.974, 41.83, 41.647]
inst: std_all
operation: std_all(baseline_shift)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.28825
varray: [34.156, 26.966, 26.1, 25.931]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.93801560645
varray: [34.156, 26.966, 26.1, 25.931]
inst: std_all
operation: std_all(slope_shift)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0364268673738
varray: [-5.054, -4.98, -5.048, -4.999]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.220732115168
varray: [-62.839, -62.637, -62.531, -62.309]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.43075
varray: [74.052, 61.242, 61.9, 60.529]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.43855384254
varray: [74.052, 61.242, 61.9, 60.529]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.23825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.47525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 44.443
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.28825
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.015
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.015
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.67468795682
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.3597
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.054, -4.98, -5.048, -4.999]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-62.839, -62.637, -62.531, -62.309]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.02025
varray: [-5.054, -4.98, -5.048, -4.999]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
dD_mean | -62.579
varray: [-62.839, -62.637, -62.531, -62.309]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0364268673738
varray: [-5.054, -4.98, -5.048, -4.999]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
dD_stdev | 0.220732115168
varray: [-62.839, -62.637, -62.531, -62.309]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.02025
varray: [-5.054, -4.98, -5.048, -4.999]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.579
varray: [-62.839, -62.637, -62.531, -62.309]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0364268673738
varray: [-5.054, -4.98, -5.048, -4.999]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.220732115168
varray: [-62.839, -62.637, -62.531, -62.309]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.25511871188
actual_array: [ 1.93 -15.84]
varray: [-5.054, -4.98, -5.048, -4.999]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.02025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.7649960851
actual_array: [ 18.1 -119. ]
varray: [-62.839, -62.637, -62.531, -62.309]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.579
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [72226.0, 72762.0, 73299.0, 73835.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.26365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 2.9411
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 29 - Sample 29 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 29
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.248
varray: [2.73, 2.101, 2.089, 2.072]
inst: mean_all
operation: mean_all(residuals)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: residuals
SAMP_residuals_STD | 0.321553520688
varray: [2.73, 2.101, 2.089, 2.072]
inst: std_all
operation: std_all(residuals)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00775
varray: [-1.226, -0.906, -0.948, -0.951]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.146943016166
varray: [-1.226, -0.906, -0.948, -0.951]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.512
varray: [48.86, 40.591, 40.314, 40.283]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.23426459573
varray: [48.86, 40.591, 40.314, 40.283]
inst: std_all
operation: std_all(baseline_shift)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 31.5565
varray: [40.702, 28.368, 28.713, 28.443]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: slope_shift
SAMP_slope_shift_STD | 6.09879979996
varray: [40.702, 28.368, 28.713, 28.443]
inst: std_all
operation: std_all(slope_shift)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0694958032306
varray: [-8.559, -8.678, -8.707, -8.702]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.377128072852
varray: [-71.1, -71.573, -71.955, -71.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.0795
varray: [78.22, 62.611, 62.174, 61.313]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.1116107114
varray: [78.22, 62.611, 62.174, 61.313]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.248
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00775
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.512
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 31.5565
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.02475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.02475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.11323512875
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.28905
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.559, -8.678, -8.707, -8.702]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-71.1, -71.573, -71.955, -71.825]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.6615
varray: [-8.559, -8.678, -8.707, -8.702]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
dD_mean | -71.61325
varray: [-71.1, -71.573, -71.955, -71.825]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0694958032306
varray: [-8.559, -8.678, -8.707, -8.702]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
dD_stdev | 0.377128072852
varray: [-71.1, -71.573, -71.955, -71.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.6615
varray: [-8.559, -8.678, -8.707, -8.702]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
measured_dD_mean | -71.61325
varray: [-71.1, -71.573, -71.955, -71.825]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0694958032306
varray: [-8.559, -8.678, -8.707, -8.702]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.377128072852
varray: [-71.1, -71.573, -71.955, -71.825]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.0567786512
actual_array: [ 1.93 -15.84]
varray: [-8.559, -8.678, -8.707, -8.702]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.6615
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -79.6911600654
actual_array: [ 18.1 -119. ]
varray: [-71.1, -71.573, -71.955, -71.825]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.61325
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [74421.0, 74956.0, 75492.0, 76027.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2539
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.0101
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 30 - Sample 30 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 30
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.34625
varray: [2.732, 2.242, 2.212, 2.199]
inst: mean_all
operation: mean_all(residuals)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: residuals
SAMP_residuals_STD | 0.257796269691
varray: [2.732, 2.242, 2.212, 2.199]
inst: std_all
operation: std_all(residuals)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.63
varray: [-1.895, -1.577, -1.413, -1.635]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.200123295329
varray: [-1.895, -1.577, -1.413, -1.635]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 93.3175
varray: [107.279, 89.553, 88.398, 88.04]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 9.33003088598
varray: [107.279, 89.553, 88.398, 88.04]
inst: std_all
operation: std_all(baseline_shift)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.80675
varray: [34.447, 26.142, 25.171, 25.467]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: slope_shift
SAMP_slope_shift_STD | 4.44544413792
varray: [34.447, 26.142, 25.171, 25.467]
inst: std_all
operation: std_all(slope_shift)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: slope_shift
SAMP_d18O_STD | 0.103757730636
varray: [-3.844, -3.713, -3.644, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.734679521969
varray: [-57.627, -56.447, -56.246, -55.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.1155
varray: [74.1, 62.306, 62.811, 61.245]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.02511468771
varray: [74.1, 62.306, 62.811, 61.245]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.34625
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.63
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 93.3175
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.80675
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.123
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.123
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 5.53244124592
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.32505
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.844, -3.713, -3.644, -3.609]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-57.627, -56.447, -56.246, -55.952]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.7025
varray: [-3.844, -3.713, -3.644, -3.609]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
dD_mean | -56.568
varray: [-57.627, -56.447, -56.246, -55.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
d18O_stdev | 0.103757730636
varray: [-3.844, -3.713, -3.644, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
dD_stdev | 0.734679521969
varray: [-57.627, -56.447, -56.246, -55.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.7025
varray: [-3.844, -3.713, -3.644, -3.609]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.568
varray: [-57.627, -56.447, -56.246, -55.952]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.103757730636
varray: [-3.844, -3.713, -3.644, -3.609]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.734679521969
varray: [-57.627, -56.447, -56.246, -55.952]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.87931716431
actual_array: [ 1.93 -15.84]
varray: [-3.844, -3.713, -3.644, -3.609]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.7025
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.1605550208
actual_array: [ 18.1 -119. ]
varray: [-57.627, -56.447, -56.246, -55.952]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.568
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [76613.0, 77149.0, 77685.0, 78221.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.15565
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 51.8156
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 31 - Sample 31 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 31
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2035
varray: [2.584, 2.098, 2.072, 2.06]
inst: mean_all
operation: mean_all(residuals)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: residuals
SAMP_residuals_STD | 0.254162021815
varray: [2.584, 2.098, 2.072, 2.06]
inst: std_all
operation: std_all(residuals)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.978
varray: [-1.039, -0.849, -1.086, -0.938]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105870990676
varray: [-1.039, -0.849, -1.086, -0.938]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 41.765
varray: [47.098, 40.268, 40.017, 39.677]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.56357245846
varray: [47.098, 40.268, 40.017, 39.677]
inst: std_all
operation: std_all(baseline_shift)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.698
varray: [37.654, 28.68, 28.147, 28.311]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: slope_shift
SAMP_slope_shift_STD | 4.64268708113
varray: [37.654, 28.68, 28.147, 28.311]
inst: std_all
operation: std_all(slope_shift)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: slope_shift
SAMP_d18O_STD | 0.110205867962
varray: [-8.666, -8.848, -8.924, -8.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.350156512624
varray: [-65.926, -66.072, -66.714, -66.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.06625
varray: [74.606, 62.396, 62.014, 61.249]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.37768747928
varray: [74.606, 62.396, 62.014, 61.249]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2035
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.978
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 41.765
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.698
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.01975
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01975
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -0.888339143146
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.2758
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.666, -8.848, -8.924, -8.854]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-65.926, -66.072, -66.714, -66.385]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.823
varray: [-8.666, -8.848, -8.924, -8.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
dD_mean | -66.27425
varray: [-65.926, -66.072, -66.714, -66.385]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
d18O_stdev | 0.110205867962
varray: [-8.666, -8.848, -8.924, -8.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
dD_stdev | 0.350156512624
varray: [-65.926, -66.072, -66.714, -66.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.823
varray: [-8.666, -8.848, -8.924, -8.854]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.27425
varray: [-65.926, -66.072, -66.714, -66.385]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.110205867962
varray: [-8.666, -8.848, -8.924, -8.854]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.350156512624
varray: [-65.926, -66.072, -66.714, -66.385]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2253932974
actual_array: [ 1.93 -15.84]
varray: [-8.666, -8.848, -8.924, -8.854]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.823
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.82506277
actual_array: [ 18.1 -119. ]
varray: [-65.926, -66.072, -66.714, -66.385]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.27425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [78807.0, 79343.0, 79879.0, 80414.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2984
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.2631
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 32 - Sample 32 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 32
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.27075
varray: [2.594, 2.185, 2.152, 2.152]
inst: mean_all
operation: mean_all(residuals)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: residuals
SAMP_residuals_STD | 0.216060755344
varray: [2.594, 2.185, 2.152, 2.152]
inst: std_all
operation: std_all(residuals)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.505
varray: [-1.826, -1.51, -1.304, -1.38]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.230283882777
varray: [-1.826, -1.51, -1.304, -1.38]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 163.56475
varray: [186.8, 157.46, 154.986, 155.013]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 15.5335357077
varray: [186.8, 157.46, 154.986, 155.013]
inst: std_all
operation: std_all(baseline_shift)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 25.7965
varray: [31.091, 24.42, 23.872, 23.803]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.54044371042
varray: [31.091, 24.42, 23.872, 23.803]
inst: std_all
operation: std_all(slope_shift)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0512924620843
varray: [-7.714, -7.62, -7.638, -7.595]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.309362112526
varray: [-73.297, -73.743, -73.997, -73.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.3295
varray: [69.357, 60.229, 60.284, 59.448]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.70053078563
varray: [69.357, 60.229, 60.284, 59.448]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.27075
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.505
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 163.56475
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 25.7965
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.0475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.13651186326
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.46095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.714, -7.62, -7.638, -7.595]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-73.297, -73.743, -73.997, -73.898]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.64175
varray: [-7.714, -7.62, -7.638, -7.595]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
dD_mean | -73.73375
varray: [-73.297, -73.743, -73.997, -73.898]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0512924620843
varray: [-7.714, -7.62, -7.638, -7.595]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
dD_stdev | 0.309362112526
varray: [-73.297, -73.743, -73.997, -73.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.64175
varray: [-7.714, -7.62, -7.638, -7.595]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.73375
varray: [-73.297, -73.743, -73.997, -73.898]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0512924620843
varray: [-7.714, -7.62, -7.638, -7.595]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.309362112526
varray: [-73.297, -73.743, -73.997, -73.898]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.99210505752
actual_array: [ 1.93 -15.84]
varray: [-7.714, -7.62, -7.638, -7.595]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.64175
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -82.0210082232
actual_array: [ 18.1 -119. ]
varray: [-73.297, -73.743, -73.997, -73.898]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.73375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [81000.0, 81536.0, 82073.0, 82609.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.23115
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 122.06285
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 33 - Sample 33 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 33
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.26225
varray: [2.64, 2.089, 2.157, 2.163]
inst: mean_all
operation: mean_all(residuals)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: residuals
SAMP_residuals_STD | 0.254059540266
varray: [2.64, 2.089, 2.157, 2.163]
inst: std_all
operation: std_all(residuals)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05525
varray: [-1.141, -0.946, -1.163, -0.971]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.112541473837
varray: [-1.141, -0.946, -1.163, -0.971]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.50125
varray: [48.462, 40.086, 40.774, 40.683]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.98553211629
varray: [48.462, 40.086, 40.774, 40.683]
inst: std_all
operation: std_all(baseline_shift)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 31.24475
varray: [38.016, 28.307, 29.299, 29.357]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: slope_shift
SAMP_slope_shift_STD | 4.53981441434
varray: [38.016, 28.307, 29.299, 29.357]
inst: std_all
operation: std_all(slope_shift)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0877738571558
varray: [-4.95, -4.854, -4.775, -4.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.3801307909
varray: [-49.862, -47.844, -47.09, -46.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 66.1515
varray: [75.947, 61.427, 63.614, 63.618]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.6113601475
varray: [75.947, 61.427, 63.614, 63.618]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.26225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05525
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.50125
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 31.24475
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.039
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.039
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.75418868773
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 1.36105
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.95, -4.854, -4.775, -4.758]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-49.862, -47.844, -47.09, -46.802]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.83425
varray: [-4.95, -4.854, -4.775, -4.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
dD_mean | -47.8995
varray: [-49.862, -47.844, -47.09, -46.802]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0877738571558
varray: [-4.95, -4.854, -4.775, -4.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
dD_stdev | 1.3801307909
varray: [-49.862, -47.844, -47.09, -46.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.83425
varray: [-4.95, -4.854, -4.775, -4.758]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
measured_dD_mean | -47.8995
varray: [-49.862, -47.844, -47.09, -46.802]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0877738571558
varray: [-4.95, -4.854, -4.775, -4.758]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.3801307909
varray: [-49.862, -47.844, -47.09, -46.802]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.06092475412
actual_array: [ 1.93 -15.84]
varray: [-4.95, -4.854, -4.775, -4.758]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.83425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.636250019
actual_array: [ 18.1 -119. ]
varray: [-49.862, -47.844, -47.09, -46.802]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.8995
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [83196.0, 83731.0, 84267.0, 84803.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.23965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.99935
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 34 - Sample 34 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 34
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.31225
varray: [2.829, 2.126, 2.16, 2.134]
inst: mean_all
operation: mean_all(residuals)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: residuals
SAMP_residuals_STD | 0.344805621571
varray: [2.829, 2.126, 2.16, 2.134]
inst: std_all
operation: std_all(residuals)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.04825
varray: [-1.113, -0.97, -1.072, -1.038]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0605110182804
varray: [-1.113, -0.97, -1.072, -1.038]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.697
varray: [49.513, 40.309, 40.589, 40.377]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 4.54556428474
varray: [49.513, 40.309, 40.589, 40.377]
inst: std_all
operation: std_all(baseline_shift)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 31.9615
varray: [41.402, 28.604, 29.239, 28.601]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: slope_shift
SAMP_slope_shift_STD | 6.300815106
varray: [41.402, 28.604, 29.239, 28.601]
inst: std_all
operation: std_all(slope_shift)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: slope_shift
SAMP_d18O_STD | 0.134764918778
varray: [-0.222, 0.002, 0.017, 0.09]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.205655050996
varray: [-52.01, -51.714, -52.215, -51.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 67.678
varray: [81.102, 63.587, 63.189, 62.834]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.9546173192
varray: [81.102, 63.587, 63.189, 62.834]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.31225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.04825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.697
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 31.9615
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.089
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.089
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 4.0031485438
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 2.88755
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-0.222, 0.002, 0.017, 0.09]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-52.01, -51.714, -52.215, -51.981]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -0.02825
varray: [-0.222, 0.002, 0.017, 0.09]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
dD_mean | -51.98
varray: [-52.01, -51.714, -52.215, -51.981]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
d18O_stdev | 0.134764918778
varray: [-0.222, 0.002, 0.017, 0.09]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
dD_stdev | 0.205655050996
varray: [-52.01, -51.714, -52.215, -51.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
measured_d18O_mean | -0.02825
varray: [-0.222, 0.002, 0.017, 0.09]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.98
varray: [-52.01, -51.714, -52.215, -51.981]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.134764918778
varray: [-0.222, 0.002, 0.017, 0.09]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.205655050996
varray: [-52.01, -51.714, -52.215, -51.981]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -1.04320345824
actual_array: [ 1.93 -15.84]
varray: [-0.222, 0.002, 0.017, 0.09]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -0.02825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -58.1196008363
actual_array: [ 18.1 -119. ]
varray: [-52.01, -51.714, -52.215, -51.981]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.98
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [85390.0, 85926.0, 86461.0, 86997.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.18965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.1951
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 35 - Sample 35 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 35
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.153
varray: [2.512, 2.01, 2.051, 2.039]
inst: mean_all
operation: mean_all(residuals)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: residuals
SAMP_residuals_STD | 0.239951383965
varray: [2.512, 2.01, 2.051, 2.039]
inst: std_all
operation: std_all(residuals)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.05725
varray: [-1.064, -1.026, -1.007, -1.132]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0551807635564
varray: [-1.064, -1.026, -1.007, -1.132]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.993
varray: [45.483, 39.143, 39.581, 39.765]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.00468123656
varray: [45.483, 39.143, 39.581, 39.765]
inst: std_all
operation: std_all(baseline_shift)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.82075
varray: [36.61, 27.127, 27.992, 27.554]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: slope_shift
SAMP_slope_shift_STD | 4.53992242041
varray: [36.61, 27.127, 27.992, 27.554]
inst: std_all
operation: std_all(slope_shift)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: slope_shift
SAMP_d18O_STD | 0.17913309763
varray: [-7.994, -8.327, -8.347, -8.376]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.529725085933
varray: [-66.314, -67.141, -67.432, -67.437]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.49925
varray: [72.481, 59.999, 60.925, 60.592]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 6.00006802045
varray: [72.481, 59.999, 60.925, 60.592]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.153
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.05725
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.993
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.82075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.07025
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.07025
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.15978859777
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.2912
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.994, -8.327, -8.347, -8.376]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-66.314, -67.141, -67.432, -67.437]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.261
varray: [-7.994, -8.327, -8.347, -8.376]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
dD_mean | -67.081
varray: [-66.314, -67.141, -67.432, -67.437]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
d18O_stdev | 0.17913309763
varray: [-7.994, -8.327, -8.347, -8.376]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
dD_stdev | 0.529725085933
varray: [-66.314, -67.141, -67.432, -67.437]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.261
varray: [-7.994, -8.327, -8.347, -8.376]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.081
varray: [-66.314, -67.141, -67.432, -67.437]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.17913309763
varray: [-7.994, -8.327, -8.347, -8.376]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.529725085933
varray: [-66.314, -67.141, -67.432, -67.437]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.63863520993
actual_array: [ 1.93 -15.84]
varray: [-7.994, -8.327, -8.347, -8.376]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.261
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.7114598468
actual_array: [ 18.1 -119. ]
varray: [-66.314, -67.141, -67.432, -67.437]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [87584.0, 88120.0, 88657.0, 89193.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3489
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.5089
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 36 - Sample 36 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 36
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.0485
varray: [2.302, 1.873, 2.018, 2.001]
inst: mean_all
operation: mean_all(residuals)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: residuals
SAMP_residuals_STD | 0.180968689741
varray: [2.302, 1.873, 2.018, 2.001]
inst: std_all
operation: std_all(residuals)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.2315
varray: [-1.503, -1.004, -1.163, -1.256]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.208774998503
varray: [-1.503, -1.004, -1.163, -1.256]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.696
varray: [43.023, 37.472, 39.348, 38.941]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.35981877835
varray: [43.023, 37.472, 39.348, 38.941]
inst: std_all
operation: std_all(baseline_shift)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.6305
varray: [31.312, 23.524, 25.755, 25.931]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.30769733601
varray: [31.312, 23.524, 25.755, 25.931]
inst: std_all
operation: std_all(slope_shift)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: slope_shift
SAMP_d18O_STD | 0.146629635477
varray: [-2.897, -2.605, -2.611, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.563457999026
varray: [-60.253, -59.111, -59.229, -59.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.0095
varray: [65.174, 56.277, 59.743, 58.844]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.74312832624
varray: [65.174, 56.277, 59.743, 58.844]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.0485
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.2315
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.696
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.6305
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.17475
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.17475
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -7.86011469695
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.78095
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-2.897, -2.605, -2.611, -2.596]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-60.253, -59.111, -59.229, -59.064]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.67725
varray: [-2.897, -2.605, -2.611, -2.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
dD_mean | -59.41425
varray: [-60.253, -59.111, -59.229, -59.064]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
d18O_stdev | 0.146629635477
varray: [-2.897, -2.605, -2.611, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
dD_stdev | 0.563457999026
varray: [-60.253, -59.111, -59.229, -59.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.67725
varray: [-2.897, -2.605, -2.611, -2.596]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
measured_dD_mean | -59.41425
varray: [-60.253, -59.111, -59.229, -59.064]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.146629635477
varray: [-2.897, -2.605, -2.611, -2.596]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.563457999026
varray: [-60.253, -59.111, -59.229, -59.064]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.80890127613
actual_array: [ 1.93 -15.84]
varray: [-2.897, -2.605, -2.611, -2.596]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.67725
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -66.2878034619
actual_array: [ 18.1 -119. ]
varray: [-60.253, -59.111, -59.229, -59.064]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.41425
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [89777.0, 90314.0, 90849.0, 91385.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.4534
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.8059
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 37 - Sample 37 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 37
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14475
varray: [2.079, 2.17, 2.16, 2.17]
inst: mean_all
operation: mean_all(residuals)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: residuals
SAMP_residuals_STD | 0.0440860900209
varray: [2.079, 2.17, 2.16, 2.17]
inst: std_all
operation: std_all(residuals)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.21625
varray: [-1.317, -1.124, -1.078, -1.346]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.134918185085
varray: [-1.317, -1.124, -1.078, -1.346]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.53075
varray: [40.018, 40.714, 40.467, 40.924]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.389530807511
varray: [40.018, 40.714, 40.467, 40.924]
inst: std_all
operation: std_all(baseline_shift)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.00075
varray: [27.318, 28.115, 27.982, 28.588]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.524218386426
varray: [27.318, 28.115, 27.982, 28.588]
inst: std_all
operation: std_all(slope_shift)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: slope_shift
SAMP_d18O_STD | 0.161761140781
varray: [1.71, 2.006, 2.012, 2.068]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.706804310494
varray: [-50.196, -49.055, -48.95, -48.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.472
varray: [60.373, 63.311, 63.338, 62.866]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.41597010326
varray: [60.373, 63.311, 63.338, 62.866]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14475
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.21625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.53075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.00075
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0785
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0785
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.53086697402
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.31845
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [1.71, 2.006, 2.012, 2.068]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-50.196, -49.055, -48.95, -48.549]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.949
varray: [1.71, 2.006, 2.012, 2.068]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
dD_mean | -49.1875
varray: [-50.196, -49.055, -48.95, -48.549]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
d18O_stdev | 0.161761140781
varray: [1.71, 2.006, 2.012, 2.068]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
dD_stdev | 0.706804310494
varray: [-50.196, -49.055, -48.95, -48.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.949
varray: [1.71, 2.006, 2.012, 2.068]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.1875
varray: [-50.196, -49.055, -48.95, -48.549]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.161761140781
varray: [1.71, 2.006, 2.012, 2.068]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.706804310494
varray: [-50.196, -49.055, -48.95, -48.549]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 1.02115139658
actual_array: [ 1.93 -15.84]
varray: [1.71, 2.006, 2.012, 2.068]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.949
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -55.0514089095
actual_array: [ 18.1 -119. ]
varray: [-50.196, -49.055, -48.95, -48.549]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -49.1875
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [91968.0, 92504.0, 93039.0, 93574.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.35715
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.97115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 38 - Sample 38 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 38
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.26825
varray: [2.193, 2.304, 2.277, 2.299]
inst: mean_all
operation: mean_all(residuals)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: residuals
SAMP_residuals_STD | 0.0515194138165
varray: [2.193, 2.304, 2.277, 2.299]
inst: std_all
operation: std_all(residuals)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -2.23225
varray: [-2.183, -2.294, -2.225, -2.227]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0458938993767
varray: [-2.183, -2.294, -2.225, -2.227]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.32425
varray: [42.313, 43.833, 43.587, 43.564]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.685071954858
varray: [42.313, 43.833, 43.587, 43.564]
inst: std_all
operation: std_all(baseline_shift)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 24.9515
varray: [24.141, 25.565, 25.074, 25.026]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.592691319322
varray: [24.141, 25.565, 25.074, 25.026]
inst: std_all
operation: std_all(slope_shift)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: slope_shift
SAMP_d18O_STD | 0.295053667435
varray: [-4.359, -4.858, -4.932, -5.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.04654526897
varray: [-60.585, -62.235, -62.581, -62.961]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.59275
varray: [59.788, 61.863, 62.082, 62.638]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.24660482244
varray: [59.788, 61.863, 62.082, 62.638]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.26825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 2.23225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.32425
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 24.9515
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.045
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.045
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 2.02406387046
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -3.1977
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.359, -4.858, -4.932, -5.015]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-60.585, -62.235, -62.581, -62.961]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.791
varray: [-4.359, -4.858, -4.932, -5.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
dD_mean | -62.0905
varray: [-60.585, -62.235, -62.581, -62.961]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
d18O_stdev | 0.295053667435
varray: [-4.359, -4.858, -4.932, -5.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
dD_stdev | 1.04654526897
varray: [-60.585, -62.235, -62.581, -62.961]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.791
varray: [-4.359, -4.858, -4.932, -5.015]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
measured_dD_mean | -62.0905
varray: [-60.585, -62.235, -62.581, -62.961]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.295053667435
varray: [-4.359, -4.858, -4.932, -5.015]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.04654526897
varray: [-60.585, -62.235, -62.581, -62.961]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.01576943867
actual_array: [ 1.93 -15.84]
varray: [-4.359, -4.858, -4.932, -5.015]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.791
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -69.228268509
actual_array: [ 18.1 -119. ]
varray: [-60.585, -62.235, -62.581, -62.961]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.0905
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [94159.0, 94695.0, 95231.0, 95767.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.23365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.82235
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 39 - Sample 39 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 39
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.129
varray: [2.141, 2.116, 2.128, 2.131]
inst: mean_all
operation: mean_all(residuals)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: residuals
SAMP_residuals_STD | 0.010295630141
varray: [2.141, 2.116, 2.128, 2.131]
inst: std_all
operation: std_all(residuals)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.09125
varray: [-1.078, -1.052, -1.131, -1.104]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0339546265871
varray: [-1.078, -1.052, -1.131, -1.104]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.9715
varray: [40.047, 39.877, 40.038, 39.924]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.0842792975766
varray: [40.047, 39.877, 40.038, 39.924]
inst: std_all
operation: std_all(baseline_shift)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.5645
varray: [28.6, 28.517, 28.57, 28.571]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.0345880518869
varray: [28.6, 28.517, 28.57, 28.571]
inst: std_all
operation: std_all(slope_shift)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: slope_shift
SAMP_d18O_STD | 0.129451406584
varray: [-2.888, -2.655, -2.668, -2.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.28181001582
varray: [-45.261, -43.142, -42.757, -42.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.53975
varray: [62.185, 63.097, 62.592, 62.285]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.409879148856
varray: [62.185, 63.097, 62.592, 62.285]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.129
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.09125
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.9715
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.5645
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.09425
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.09425
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.23928932869
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.2507
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-2.888, -2.655, -2.668, -2.591]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-45.261, -43.142, -42.757, -42.407]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -2.7005
varray: [-2.888, -2.655, -2.668, -2.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
dD_mean | -43.39175
varray: [-45.261, -43.142, -42.757, -42.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
d18O_stdev | 0.129451406584
varray: [-2.888, -2.655, -2.668, -2.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
dD_stdev | 1.28181001582
varray: [-45.261, -43.142, -42.757, -42.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
measured_d18O_mean | -2.7005
varray: [-2.888, -2.655, -2.668, -2.591]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
measured_dD_mean | -43.39175
varray: [-45.261, -43.142, -42.757, -42.407]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.129451406584
varray: [-2.888, -2.655, -2.668, -2.591]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.28181001582
varray: [-45.261, -43.142, -42.757, -42.407]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -3.83317552085
actual_array: [ 1.93 -15.84]
varray: [-2.888, -2.655, -2.668, -2.591]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.7005
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -48.6834685837
actual_array: [ 18.1 -119. ]
varray: [-45.261, -43.142, -42.757, -42.407]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -43.39175
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [96352.0, 96887.0, 97423.0, 97960.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3729
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.5304
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 40 - Sample 40 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 40
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1545
varray: [2.276, 2.162, 2.102, 2.078]
inst: mean_all
operation: mean_all(residuals)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: residuals
SAMP_residuals_STD | 0.0883685464405
varray: [2.276, 2.162, 2.102, 2.078]
inst: std_all
operation: std_all(residuals)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97425
varray: [-0.963, -0.858, -1.077, -0.999]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0909409148843
varray: [-0.963, -0.858, -1.077, -0.999]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.6025
varray: [42.131, 40.384, 40.032, 39.863]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.04185491632
varray: [42.131, 40.384, 40.032, 39.863]
inst: std_all
operation: std_all(baseline_shift)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.74375
varray: [31.669, 29.453, 29.343, 28.51]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.35078529629
varray: [31.669, 29.453, 29.343, 28.51]
inst: std_all
operation: std_all(slope_shift)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: slope_shift
SAMP_d18O_STD | 0.177061194695
varray: [-8.458, -8.764, -8.81, -8.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.17489613016
varray: [-63.493, -65.418, -65.924, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.7705
varray: [67.227, 64.654, 62.182, 61.019]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.75814243045
varray: [67.227, 64.654, 62.182, 61.019]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1545
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97425
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.6025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.74375
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.06875
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.06875
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.09231980209
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.01995
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.458, -8.764, -8.81, -8.844]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-63.493, -65.418, -65.924, -66.01]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.719
varray: [-8.458, -8.764, -8.81, -8.844]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
dD_mean | -65.21125
varray: [-63.493, -65.418, -65.924, -66.01]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
d18O_stdev | 0.177061194695
varray: [-8.458, -8.764, -8.81, -8.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
dD_stdev | 1.17489613016
varray: [-63.493, -65.418, -65.924, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.719
varray: [-8.458, -8.764, -8.81, -8.844]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.21125
varray: [-63.493, -65.418, -65.924, -66.01]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.177061194695
varray: [-8.458, -8.764, -8.81, -8.844]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.17489613016
varray: [-63.493, -65.418, -65.924, -66.01]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1168117296
actual_array: [ 1.93 -15.84]
varray: [-8.458, -8.764, -8.81, -8.844]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.719
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -72.657117195
actual_array: [ 18.1 -119. ]
varray: [-63.493, -65.418, -65.924, -66.01]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.21125
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [98545.0, 99080.0, 99616.0, 100152.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3474
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.8994
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 41 - Sample 41 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 41
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.2415
varray: [2.933, 2.045, 2.018, 1.97]
inst: mean_all
operation: mean_all(residuals)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: residuals
SAMP_residuals_STD | 0.462042205864
varray: [2.933, 2.045, 2.018, 1.97]
inst: std_all
operation: std_all(residuals)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.18825
varray: [-1.485, -1.083, -1.09, -1.095]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.197894542623
varray: [-1.485, -1.083, -1.09, -1.095]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 43.35625
varray: [53.102, 40.452, 40.264, 39.607]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 6.5072574036
varray: [53.102, 40.452, 40.264, 39.607]
inst: std_all
operation: std_all(baseline_shift)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 30.4055
varray: [43.04, 26.649, 26.208, 25.725]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: slope_shift
SAMP_slope_shift_STD | 8.43144845208
varray: [43.04, 26.649, 26.208, 25.725]
inst: std_all
operation: std_all(slope_shift)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: slope_shift
SAMP_d18O_STD | 0.110707949127
varray: [-4.685, -4.503, -4.482, -4.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.117991525119
varray: [-67.83, -67.681, -67.576, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 65.4675
varray: [82.948, 61.063, 59.877, 57.982]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 11.7225409788
varray: [82.948, 61.063, 59.877, 57.982]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.2415
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.18825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 43.35625
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 30.4055
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 0.01825
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.01825
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 0.820870347464
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | 0.67705
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-4.685, -4.503, -4.482, -4.431]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-67.83, -67.681, -67.576, -67.585]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -4.52525
varray: [-4.685, -4.503, -4.482, -4.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
dD_mean | -67.668
varray: [-67.83, -67.681, -67.576, -67.585]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
d18O_stdev | 0.110707949127
varray: [-4.685, -4.503, -4.482, -4.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
dD_stdev | 0.117991525119
varray: [-67.83, -67.681, -67.576, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
measured_d18O_mean | -4.52525
varray: [-4.685, -4.503, -4.482, -4.431]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.668
varray: [-67.83, -67.681, -67.576, -67.585]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.110707949127
varray: [-4.685, -4.503, -4.482, -4.431]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.117991525119
varray: [-67.83, -67.681, -67.576, -67.585]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -5.73831221137
actual_array: [ 1.93 -15.84]
varray: [-4.685, -4.503, -4.482, -4.431]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.52525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.3564119188
actual_array: [ 18.1 -119. ]
varray: [-67.83, -67.681, -67.576, -67.585]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.668
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [100739.0, 101275.0, 101810.0, 102344.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.2604
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.85435
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 42 - Sample 42 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 42
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.22325
varray: [2.516, 2.099, 2.122, 2.156]
inst: mean_all
operation: mean_all(residuals)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: residuals
SAMP_residuals_STD | 0.196566146628
varray: [2.516, 2.099, 2.122, 2.156]
inst: std_all
operation: std_all(residuals)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.24825
varray: [-1.304, -1.25, -1.245, -1.194]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0449620210103
varray: [-1.304, -1.25, -1.245, -1.194]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.22675
varray: [46.187, 40.663, 40.864, 41.193]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.64918986044
varray: [46.187, 40.663, 40.864, 41.193]
inst: std_all
operation: std_all(baseline_shift)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.07875
varray: [34.307, 26.68, 27.172, 28.156]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.53910359413
varray: [34.307, 26.68, 27.172, 28.156]
inst: std_all
operation: std_all(slope_shift)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: slope_shift
SAMP_d18O_STD | 0.25353681521
varray: [5.069, 5.463, 5.565, 5.638]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.63076045656
varray: [-41.145, -38.515, -37.897, -37.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.4625
varray: [71.771, 61.668, 61.946, 62.465]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.88351584414
varray: [71.771, 61.668, 61.946, 62.465]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.22325
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.24825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.22675
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.07875
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | 4.4408920985e-16
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 4.4408920985e-16
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | 1.99747761093e-14
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.32795
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [5.069, 5.463, 5.565, 5.638]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-41.145, -38.515, -37.897, -37.541]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 5.43375
varray: [5.069, 5.463, 5.565, 5.638]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
dD_mean | -38.7745
varray: [-41.145, -38.515, -37.897, -37.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
d18O_stdev | 0.25353681521
varray: [5.069, 5.463, 5.565, 5.638]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
dD_stdev | 1.63076045656
varray: [-41.145, -38.515, -37.897, -37.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
measured_d18O_mean | 5.43375
varray: [5.069, 5.463, 5.565, 5.638]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
measured_dD_mean | -38.7745
varray: [-41.145, -38.515, -37.897, -37.541]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.25353681521
varray: [5.069, 5.463, 5.565, 5.638]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.63076045656
varray: [-41.145, -38.515, -37.897, -37.541]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 4.6594169575
actual_array: [ 1.93 -15.84]
varray: [5.069, 5.463, 5.565, 5.638]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 5.43375
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -43.6103766682
actual_array: [ 18.1 -119. ]
varray: [-41.145, -38.515, -37.897, -37.541]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -38.7745
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [102930.0, 103466.0, 104002.0, 104538.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.27865
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 0.72485
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 43 - Sample 43 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 43
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.173
varray: [2.486, 2.101, 2.046, 2.059]
inst: mean_all
operation: mean_all(residuals)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: residuals
SAMP_residuals_STD | 0.209982538957
varray: [2.486, 2.101, 2.046, 2.059]
inst: std_all
operation: std_all(residuals)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.968
varray: [-1.006, -0.955, -1.019, -0.892]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.057706152185
varray: [-1.006, -0.955, -1.019, -0.892]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.93075
varray: [44.931, 40.176, 39.296, 39.32]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 2.69805922532
varray: [44.931, 40.176, 39.296, 39.32]
inst: std_all
operation: std_all(baseline_shift)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.979
varray: [35.339, 28.73, 27.86, 27.987]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: slope_shift
SAMP_slope_shift_STD | 3.59387562389
varray: [35.339, 28.73, 27.86, 27.987]
inst: std_all
operation: std_all(slope_shift)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: slope_shift
SAMP_d18O_STD | 0.338767742463
varray: [-7.693, -8.252, -8.414, -8.398]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.38247953571
varray: [-63.942, -66.324, -66.937, -66.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 64.389
varray: [72.375, 62.863, 61.188, 61.13]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 5.38430911693
varray: [72.375, 62.863, 61.188, 61.13]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.173
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.968
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.93075
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.979
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05025
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05025
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.26020465535
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -0.40145
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.693, -8.252, -8.414, -8.398]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-63.942, -66.324, -66.937, -66.719]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.18925
varray: [-7.693, -8.252, -8.414, -8.398]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
dD_mean | -65.9805
varray: [-63.942, -66.324, -66.937, -66.719]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
d18O_stdev | 0.338767742463
varray: [-7.693, -8.252, -8.414, -8.398]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
dD_stdev | 1.38247953571
varray: [-63.942, -66.324, -66.937, -66.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.18925
varray: [-7.693, -8.252, -8.414, -8.398]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.9805
varray: [-63.942, -66.324, -66.937, -66.719]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.338767742463
varray: [-7.693, -8.252, -8.414, -8.398]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.38247953571
varray: [-63.942, -66.324, -66.937, -66.719]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.56372436869
actual_array: [ 1.93 -15.84]
varray: [-7.693, -8.252, -8.414, -8.398]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.18925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.5023120526
actual_array: [ 18.1 -119. ]
varray: [-63.942, -66.324, -66.937, -66.719]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.9805
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [105122.0, 105658.0, 106194.0, 106729.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3289
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -0.57115
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 44 - Sample 44 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 44
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12275
varray: [2.253, 2.103, 2.074, 2.061]
inst: mean_all
operation: mean_all(residuals)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: residuals
SAMP_residuals_STD | 0.0885903493615
varray: [2.253, 2.103, 2.074, 2.061]
inst: std_all
operation: std_all(residuals)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.039
varray: [-0.985, -1.102, -0.874, -1.195]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.139577935219
varray: [-0.985, -1.102, -0.874, -1.195]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.1025
varray: [41.526, 39.93, 39.356, 39.598]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.977733944725
varray: [41.526, 39.93, 39.356, 39.598]
inst: std_all
operation: std_all(baseline_shift)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.86625
varray: [31.112, 28.65, 27.763, 27.94]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.54545103988
varray: [31.112, 28.65, 27.763, 27.94]
inst: std_all
operation: std_all(slope_shift)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: slope_shift
SAMP_d18O_STD | 0.214857316996
varray: [-3.683, -3.333, -3.247, -3.215]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.772998706338
varray: [-57.614, -56.503, -55.961, -55.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.6505
varray: [65.816, 61.458, 62.152, 61.176]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.14981386791
varray: [65.816, 61.458, 62.152, 61.176]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12275
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.039
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.1025
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.86625
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.1005
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.1005
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.52040931069
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.13995
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.683, -3.333, -3.247, -3.215]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-57.614, -56.503, -55.961, -55.988]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.3695
varray: [-3.683, -3.333, -3.247, -3.215]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
dD_mean | -56.5165
varray: [-57.614, -56.503, -55.961, -55.988]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
d18O_stdev | 0.214857316996
varray: [-3.683, -3.333, -3.247, -3.215]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
dD_stdev | 0.772998706338
varray: [-57.614, -56.503, -55.961, -55.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.3695
varray: [-3.683, -3.333, -3.247, -3.215]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.5165
varray: [-57.614, -56.503, -55.961, -55.988]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.214857316996
varray: [-3.683, -3.333, -3.247, -3.215]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.772998706338
varray: [-57.614, -56.503, -55.961, -55.988]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.53164733669
actual_array: [ 1.93 -15.84]
varray: [-3.683, -3.333, -3.247, -3.215]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.3695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.1039706397
actual_array: [ 18.1 -119. ]
varray: [-57.614, -56.503, -55.961, -55.988]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.5165
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [107313.0, 107849.0, 108386.0, 108921.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.37915
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.3994
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 45 - Sample 45 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 45
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.10825
varray: [2.153, 2.098, 2.096, 2.086]
inst: mean_all
operation: mean_all(residuals)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: residuals
SAMP_residuals_STD | 0.0302916380101
varray: [2.153, 2.098, 2.096, 2.086]
inst: std_all
operation: std_all(residuals)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.08175
varray: [-0.987, -1.039, -1.2, -1.101]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0915764707772
varray: [-0.987, -1.039, -1.2, -1.101]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.16375
varray: [40.61, 39.921, 40.183, 39.941]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.320444873054
varray: [40.61, 39.921, 40.183, 39.941]
inst: std_all
operation: std_all(baseline_shift)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.3565
varray: [28.078, 27.209, 27.202, 26.937]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.49738281702
varray: [28.078, 27.209, 27.202, 26.937]
inst: std_all
operation: std_all(slope_shift)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: slope_shift
SAMP_d18O_STD | 0.141991490355
varray: [1.149, 1.335, 1.432, 1.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.482532468683
varray: [-48.454, -47.705, -47.438, -47.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.87575
varray: [63.692, 61.205, 61.419, 61.187]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.21541032166
varray: [63.692, 61.205, 61.419, 61.187]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.10825
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.08175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.16375
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.3565
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.115
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.115
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.17260766895
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.9147
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [1.149, 1.335, 1.432, 1.465]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-48.454, -47.705, -47.438, -47.428]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | 1.34525
varray: [1.149, 1.335, 1.432, 1.465]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
dD_mean | -47.75625
varray: [-48.454, -47.705, -47.438, -47.428]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
d18O_stdev | 0.141991490355
varray: [1.149, 1.335, 1.432, 1.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
dD_stdev | 0.482532468683
varray: [-48.454, -47.705, -47.438, -47.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
measured_d18O_mean | 1.34525
varray: [1.149, 1.335, 1.432, 1.465]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
measured_dD_mean | -47.75625
varray: [-48.454, -47.705, -47.438, -47.428]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.141991490355
varray: [1.149, 1.335, 1.432, 1.465]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.482532468683
varray: [-48.454, -47.705, -47.438, -47.428]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | 0.390804073983
actual_array: [ 1.93 -15.84]
varray: [1.149, 1.335, 1.432, 1.465]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.34525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -53.4788575415
actual_array: [ 18.1 -119. ]
varray: [-48.454, -47.705, -47.438, -47.428]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.75625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [109506.0, 110042.0, 110577.0, 111113.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.39365
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.33815
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 46 - Sample 46 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 46
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.098
varray: [2.109, 2.087, 2.09, 2.106]
inst: mean_all
operation: mean_all(residuals)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: residuals
SAMP_residuals_STD | 0.011105554166
varray: [2.109, 2.087, 2.09, 2.106]
inst: std_all
operation: std_all(residuals)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.97225
varray: [-0.959, -0.958, -1.168, -0.804]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.14944870915
varray: [-0.959, -0.958, -1.168, -0.804]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.63975
varray: [39.746, 39.67, 39.705, 39.438]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.138039547473
varray: [39.746, 39.67, 39.705, 39.438]
inst: std_all
operation: std_all(baseline_shift)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.328
varray: [28.6, 28.231, 28.186, 28.295]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.186767234814
varray: [28.6, 28.231, 28.186, 28.295]
inst: std_all
operation: std_all(slope_shift)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: slope_shift
SAMP_d18O_STD | 0.188254039355
varray: [-5.478, -5.785, -5.836, -5.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.469092297244
varray: [-56.1, -56.863, -57.09, -57.088]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 61.96075
varray: [62.568, 61.702, 61.208, 62.365]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.623365796837
varray: [62.568, 61.702, 61.208, 62.365]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.098
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.97225
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.63975
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.328
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.12525
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.12525
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.63364443945
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.8297
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.478, -5.785, -5.836, -5.904]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-56.1, -56.863, -57.09, -57.088]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.75075
varray: [-5.478, -5.785, -5.836, -5.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
dD_mean | -56.78525
varray: [-56.1, -56.863, -57.09, -57.088]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
d18O_stdev | 0.188254039355
varray: [-5.478, -5.785, -5.836, -5.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
dD_stdev | 0.469092297244
varray: [-56.1, -56.863, -57.09, -57.088]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.75075
varray: [-5.478, -5.785, -5.836, -5.904]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
measured_dD_mean | -56.78525
varray: [-56.1, -56.863, -57.09, -57.088]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.188254039355
varray: [-5.478, -5.785, -5.836, -5.904]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.469092297244
varray: [-56.1, -56.863, -57.09, -57.088]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.01779982021
actual_array: [ 1.93 -15.84]
varray: [-5.478, -5.785, -5.836, -5.904]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.75075
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -63.399253211
actual_array: [ 18.1 -119. ]
varray: [-56.1, -56.863, -57.09, -57.088]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.78525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [111699.0, 112234.0, 112769.0, 113304.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.4039
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.86215
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 47 - Sample 47 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 47
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.091
varray: [2.113, 2.127, 2.066, 2.058]
inst: mean_all
operation: mean_all(residuals)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: residuals
SAMP_residuals_STD | 0.034127212993
varray: [2.113, 2.127, 2.066, 2.058]
inst: std_all
operation: std_all(residuals)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.9705
varray: [-0.86, -1.077, -0.989, -0.956]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0896381615162
varray: [-0.86, -1.077, -0.989, -0.956]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 39.82525
varray: [39.942, 40.147, 39.625, 39.587]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.267093460546
varray: [39.942, 40.147, 39.625, 39.587]
inst: std_all
operation: std_all(baseline_shift)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.578
varray: [28.881, 29.277, 28.004, 28.15]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: slope_shift
SAMP_slope_shift_STD | 0.60361963741
varray: [28.881, 29.277, 28.004, 28.15]
inst: std_all
operation: std_all(slope_shift)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0615974837148
varray: [-8.477, -8.585, -8.605, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.524903165673
varray: [-65.688, -66.606, -66.665, -66.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 62.18
varray: [62.892, 62.376, 61.851, 61.601]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 0.574120196475
varray: [62.892, 62.376, 61.851, 61.601]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.091
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.9705
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 39.82525
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.578
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.13225
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.13225
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -5.9484988193
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -2.61045
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.477, -8.585, -8.605, -8.606]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-65.688, -66.606, -66.665, -66.869]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.56825
varray: [-8.477, -8.585, -8.605, -8.606]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
dD_mean | -66.457
varray: [-65.688, -66.606, -66.665, -66.869]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0615974837148
varray: [-8.477, -8.585, -8.605, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
dD_stdev | 0.524903165673
varray: [-65.688, -66.606, -66.665, -66.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.56825
varray: [-8.477, -8.585, -8.605, -8.606]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.457
varray: [-65.688, -66.606, -66.665, -66.869]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0615974837148
varray: [-8.477, -8.585, -8.605, -8.606]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.524903165673
varray: [-65.688, -66.606, -66.665, -66.869]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -9.95942065898
actual_array: [ 1.93 -15.84]
varray: [-8.477, -8.585, -8.605, -8.606]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.56825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.0258549185
actual_array: [ 18.1 -119. ]
varray: [-65.688, -66.606, -66.665, -66.869]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.457
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [113889.0, 114425.0, 114961.0, 115497.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.4109
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.67665
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 48 - Sample 48 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 48
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.164
varray: [2.248, 2.135, 2.132, 2.141]
inst: mean_all
operation: mean_all(residuals)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: residuals
SAMP_residuals_STD | 0.0561248608016
varray: [2.248, 2.135, 2.132, 2.141]
inst: std_all
operation: std_all(residuals)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.3045
varray: [-1.273, -1.383, -1.308, -1.254]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0569122131005
varray: [-1.273, -1.383, -1.308, -1.254]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 48.613
varray: [49.761, 48.482, 48.06, 48.149]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 0.786589685499
varray: [49.761, 48.482, 48.06, 48.149]
inst: std_all
operation: std_all(baseline_shift)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 27.8965
varray: [29.395, 27.513, 27.171, 27.507]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.0117040081
varray: [29.395, 27.513, 27.171, 27.507]
inst: std_all
operation: std_all(slope_shift)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0701237002256
varray: [-6.273, -6.125, -6.149, -6.129]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.158789168396
varray: [-67.278, -67.429, -67.539, -67.65]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.071
varray: [64.912, 61.867, 62.885, 62.62]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 1.3008835459
varray: [64.912, 61.867, 62.885, 62.62]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.164
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.3045
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 48.613
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 27.8965
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.05925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.05925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.66501742944
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.71945
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-6.273, -6.125, -6.149, -6.129]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-67.278, -67.429, -67.539, -67.65]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -6.169
varray: [-6.273, -6.125, -6.149, -6.129]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
dD_mean | -67.474
varray: [-67.278, -67.429, -67.539, -67.65]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0701237002256
varray: [-6.273, -6.125, -6.149, -6.129]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
dD_stdev | 0.158789168396
varray: [-67.278, -67.429, -67.539, -67.65]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
measured_d18O_mean | -6.169
varray: [-6.273, -6.125, -6.149, -6.129]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
measured_dD_mean | -67.474
varray: [-67.278, -67.429, -67.539, -67.65]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0701237002256
varray: [-6.273, -6.125, -6.149, -6.129]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.158789168396
varray: [-67.278, -67.429, -67.539, -67.65]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -7.45447521181
actual_array: [ 1.93 -15.84]
varray: [-6.273, -6.125, -6.149, -6.129]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.169
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -75.1432591046
actual_array: [ 18.1 -119. ]
varray: [-67.278, -67.429, -67.539, -67.65]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.474
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [116083.0, 116619.0, 117154.0, 117690.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3379
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 7.1111
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 49 - Sample 49 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 49
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.14225
varray: [2.308, 2.092, 2.1, 2.069]
inst: mean_all
operation: mean_all(residuals)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: residuals
SAMP_residuals_STD | 0.111278554388
varray: [2.308, 2.092, 2.1, 2.069]
inst: std_all
operation: std_all(residuals)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -0.89625
varray: [-0.891, -0.868, -0.883, -0.943]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.0325921769755
varray: [-0.891, -0.868, -0.883, -0.943]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 40.42575
varray: [42.607, 39.793, 39.839, 39.464]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.46372387537
varray: [42.607, 39.793, 39.839, 39.464]
inst: std_all
operation: std_all(baseline_shift)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.3865
varray: [32.302, 28.719, 28.741, 27.784]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.99418930897
varray: [32.302, 28.719, 28.741, 27.784]
inst: std_all
operation: std_all(slope_shift)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0587898800815
varray: [-8.77, -8.836, -8.893, -8.894]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.0746681547471
varray: [-66.871, -66.727, -66.816, -66.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.65425
varray: [68.298, 61.853, 62.173, 62.293]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.10139854206
varray: [68.298, 61.853, 62.173, 62.293]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.14225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 0.89625
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 40.42575
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.3865
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.081
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.081
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.64331496683
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1362
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.77, -8.836, -8.893, -8.894]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-66.871, -66.727, -66.816, -66.714]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.84825
varray: [-8.77, -8.836, -8.893, -8.894]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
dD_mean | -66.782
varray: [-66.871, -66.727, -66.816, -66.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0587898800815
varray: [-8.77, -8.836, -8.893, -8.894]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
dD_stdev | 0.0746681547471
varray: [-66.871, -66.727, -66.816, -66.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.84825
varray: [-8.77, -8.836, -8.893, -8.894]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
measured_dD_mean | -66.782
varray: [-66.871, -66.727, -66.816, -66.714]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0587898800815
varray: [-8.77, -8.836, -8.893, -8.894]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.0746681547471
varray: [-66.871, -66.727, -66.816, -66.714]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.2517556492
actual_array: [ 1.93 -15.84]
varray: [-8.77, -8.836, -8.893, -8.894]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.84825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -74.3829408187
actual_array: [ 18.1 -119. ]
varray: [-66.871, -66.727, -66.816, -66.714]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.782
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [118276.0, 118812.0, 119348.0, 119884.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.35965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | True
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | False
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | False
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | -1.07615
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 50 - Sample 50 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 50
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.16225
varray: [2.336, 2.123, 2.108, 2.082]
inst: mean_all
operation: mean_all(residuals)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: residuals
SAMP_residuals_STD | 0.1170651528
varray: [2.336, 2.123, 2.108, 2.082]
inst: std_all
operation: std_all(residuals)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.08875
varray: [-1.239, -0.984, -0.997, -1.135]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.121250085911
varray: [-1.239, -0.984, -0.997, -1.135]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 45.96
varray: [48.685, 45.299, 44.955, 44.901]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 1.82519880196
varray: [48.685, 45.299, 44.955, 44.901]
inst: std_all
operation: std_all(baseline_shift)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 28.58175
varray: [31.749, 27.79, 27.599, 27.189]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.12633414981
varray: [31.749, 27.79, 27.599, 27.189]
inst: std_all
operation: std_all(slope_shift)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0457456737481
varray: [-7.262, -7.155, -7.182, -7.189]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.381433589676
varray: [-73.104, -73.558, -73.828, -73.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.59975
varray: [67.891, 62.722, 62.128, 61.658]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 2.89376990274
varray: [67.891, 62.722, 62.128, 61.658]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.16225
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.08875
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 45.96
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 28.58175
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.061
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.061
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -2.7437310244
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.1907
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-7.262, -7.155, -7.182, -7.189]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-73.104, -73.558, -73.828, -73.971]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -7.197
varray: [-7.262, -7.155, -7.182, -7.189]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
dD_mean | -73.61525
varray: [-73.104, -73.558, -73.828, -73.971]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0457456737481
varray: [-7.262, -7.155, -7.182, -7.189]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
dD_stdev | 0.381433589676
varray: [-73.104, -73.558, -73.828, -73.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
measured_d18O_mean | -7.197
varray: [-7.262, -7.155, -7.182, -7.189]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
measured_dD_mean | -73.61525
varray: [-73.104, -73.558, -73.828, -73.971]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0457456737481
varray: [-7.262, -7.155, -7.182, -7.189]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.381433589676
varray: [-73.104, -73.558, -73.828, -73.971]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -8.52776224721
actual_array: [ 1.93 -15.84]
varray: [-7.262, -7.155, -7.182, -7.189]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.197
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -81.8908092104
actual_array: [ 18.1 -119. ]
varray: [-73.104, -73.558, -73.828, -73.971]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.61525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [120468.0, 121003.0, 121538.0, 122074.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.33965
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 4.4581
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 51 - Sample 51 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 51
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.12875
varray: [2.632, 2.006, 1.939, 1.938]
inst: mean_all
operation: mean_all(residuals)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: residuals
SAMP_residuals_STD | 0.337005811028
varray: [2.632, 2.006, 1.939, 1.938]
inst: std_all
operation: std_all(residuals)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.386
varray: [-1.6, -1.484, -1.208, -1.252]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.187118500778
varray: [-1.6, -1.484, -1.208, -1.252]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 67.22275
varray: [80.334, 64.152, 62.189, 62.216]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 8.78901903419
varray: [80.334, 64.152, 62.189, 62.216]
inst: std_all
operation: std_all(baseline_shift)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 26.176
varray: [34.513, 24.422, 22.879, 22.89]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: slope_shift
SAMP_slope_shift_STD | 5.60505991642
varray: [34.513, 24.422, 22.879, 22.89]
inst: std_all
operation: std_all(slope_shift)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: slope_shift
SAMP_d18O_STD | 0.121867072392
varray: [-3.695, -3.508, -3.451, -3.425]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 1.40537738111
varray: [-51.049, -48.887, -48.211, -47.954]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 60.5885
varray: [73.161, 57.031, 56.208, 55.954]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 8.39426258425
varray: [73.161, 57.031, 56.208, 55.954]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.12875
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.386
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 67.22275
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 26.176
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.0945
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.0945
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -4.25053412797
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -4.20195
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-3.695, -3.508, -3.451, -3.425]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-51.049, -48.887, -48.211, -47.954]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -3.51975
varray: [-3.695, -3.508, -3.451, -3.425]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
dD_mean | -49.02525
varray: [-51.049, -48.887, -48.211, -47.954]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
d18O_stdev | 0.121867072392
varray: [-3.695, -3.508, -3.451, -3.425]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
dD_stdev | 1.40537738111
varray: [-51.049, -48.887, -48.211, -47.954]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
measured_d18O_mean | -3.51975
varray: [-3.695, -3.508, -3.451, -3.425]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
measured_dD_mean | -49.02525
varray: [-51.049, -48.887, -48.211, -47.954]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.121867072392
varray: [-3.695, -3.508, -3.451, -3.425]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
measured_dD_stdev | 1.40537738111
varray: [-51.049, -48.887, -48.211, -47.954]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -4.68851638054
actual_array: [ 1.93 -15.84]
varray: [-3.695, -3.508, -3.451, -3.425]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.51975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -54.8731406409
actual_array: [ 18.1 -119. ]
varray: [-51.049, -48.887, -48.211, -47.954]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -49.02525
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [122660.0, 123195.0, 123730.0, 124265.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.37315
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 25.72085
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 52 - Sample 52 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 52
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 1.944
varray: [2.139, 2.056, 1.844, 1.737]
inst: mean_all
operation: mean_all(residuals)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: residuals
SAMP_residuals_STD | 0.185668162771
varray: [2.139, 2.056, 1.844, 1.737]
inst: std_all
operation: std_all(residuals)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.59175
varray: [-1.508, -1.709, -1.652, -1.498]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.105167723185
varray: [-1.508, -1.709, -1.652, -1.498]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 383.432
varray: [417.821, 407.499, 364.269, 344.139]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 34.9903997691
varray: [417.821, 407.499, 364.269, 344.139]
inst: std_all
operation: std_all(baseline_shift)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 11.3655
varray: [13.129, 12.532, 10.361, 9.44]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: slope_shift
SAMP_slope_shift_STD | 1.74997380933
varray: [13.129, 12.532, 10.361, 9.44]
inst: std_all
operation: std_all(slope_shift)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0551142752228
varray: [-5.377, -5.468, -5.492, -5.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.167510198694
varray: [-50.83, -51.212, -51.14, -51.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 49.56475
varray: [54.443, 51.58, 47.278, 44.958]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 4.25479834031
varray: [54.443, 51.58, 47.278, 44.958]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 1.944
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.59175
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 383.432
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 11.3655
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.27925
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.27925
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -12.5604407961
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -15.2257
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-5.377, -5.468, -5.492, -5.494]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-50.83, -51.212, -51.14, -51.012]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -5.45775
varray: [-5.377, -5.468, -5.492, -5.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
dD_mean | -51.0485
varray: [-50.83, -51.212, -51.14, -51.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0551142752228
varray: [-5.377, -5.468, -5.492, -5.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
dD_stdev | 0.167510198694
varray: [-50.83, -51.212, -51.14, -51.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
measured_d18O_mean | -5.45775
varray: [-5.377, -5.468, -5.492, -5.494]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
measured_dD_mean | -51.0485
varray: [-50.83, -51.212, -51.14, -51.012]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0551142752228
varray: [-5.377, -5.468, -5.492, -5.494]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.167510198694
varray: [-50.83, -51.212, -51.14, -51.012]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -6.71189213405
actual_array: [ 1.93 -15.84]
varray: [-5.377, -5.468, -5.492, -5.494]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.45775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -57.0961377102
actual_array: [ 18.1 -119. ]
varray: [-50.83, -51.212, -51.14, -51.012]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.0485
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [124851.0, 125385.0, 125920.0, 126456.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.5579
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | True
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | True
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | False
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | False
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | True
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | True
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | True
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 341.9301
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 53 - Sample 53 Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Sample Variables
Sample 53
Variable | Value | Operations
IS_STANDARD | False
operation: is_standard_sample()
inst: is_standard_sample
IS_NOT_STANDARD | True
operation: bool_xor(IS_STANDARD,TRUE)
inst: bool_xor
factors: ('IS_STANDARD', ',', 'TRUE')
SAMP_residuals_MEAN | 2.1405
varray: [2.374, 2.076, 2.07, 2.042]
inst: mean_all
operation: mean_all(residuals)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: residuals
SAMP_residuals_STD | 0.156370287033
varray: [2.374, 2.076, 2.07, 2.042]
inst: std_all
operation: std_all(residuals)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: residuals
SAMP_ch4_ppm_MEAN | -1.00825
varray: [-1.154, -0.978, -1.061, -0.84]
inst: mean_all
operation: mean_all(ch4_ppm)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: ch4_ppm
SAMP_ch4_ppm_STD | 0.133227562213
varray: [-1.154, -0.978, -1.061, -0.84]
inst: std_all
operation: std_all(ch4_ppm)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: ch4_ppm
SAMP_baseline_shift_MEAN | 42.724
varray: [48.555, 41.556, 40.803, 39.982]
inst: mean_all
operation: mean_all(baseline_shift)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: baseline_shift
SAMP_baseline_shift_STD | 3.94011801854
varray: [48.555, 41.556, 40.803, 39.982]
inst: std_all
operation: std_all(baseline_shift)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: baseline_shift
SAMP_slope_shift_MEAN | 29.32
varray: [33.693, 27.823, 28.031, 27.733]
inst: mean_all
operation: mean_all(slope_shift)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: slope_shift
SAMP_slope_shift_STD | 2.91800319854
varray: [33.693, 27.823, 28.031, 27.733]
inst: std_all
operation: std_all(slope_shift)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: slope_shift
SAMP_d18O_STD | 0.0821599050632
varray: [-8.628, -8.686, -8.767, -8.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
SAMP_dD_STD | 0.686702992567
varray: [-64.622, -65.48, -66.068, -66.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
SAMP_baseline_curvature_MEAN | 63.2745
varray: [68.351, 62.102, 61.007, 61.638]
inst: mean_all
operation: mean_all(baseline_curvature)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: baseline_curvature
SAMP_baseline_curvature_STD | 3.41395650236
varray: [68.351, 62.102, 61.007, 61.638]
inst: std_all
operation: std_all(baseline_curvature)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: baseline_curvature
SAMP_residuals_MEAN_ABS | 2.1405
operation: abs(SAMP_residuals_MEAN)
inst: abs
source_col: SAMP_residuals_MEAN
SAMP_ch4_ppm_MEAN_ABS | 1.00825
operation: abs(SAMP_ch4_ppm_MEAN)
inst: abs
source_col: SAMP_ch4_ppm_MEAN
SAMP_baseline_shift_MEAN_ABS | 42.724
operation: abs(SAMP_baseline_shift_MEAN)
inst: abs
source_col: SAMP_baseline_shift_MEAN
SAMP_slope_shift_MEAN_ABS | 29.32
operation: abs(SAMP_slope_shift_MEAN)
inst: abs
source_col: SAMP_slope_shift_MEAN
SAMP_residuals_DIFFERENTIAL | -0.08275
operation: evaluation(SAMP_residuals_MEAN-residuals_MEAN)
inst: evaluation
factors: SAMP_residuals_MEAN-residuals_MEAN
SAMP_residuals_DIFFERENTIAL_ABS | 0.08275
operation: abs(SAMP_residuals_DIFFERENTIAL)
inst: abs
source_col: SAMP_residuals_DIFFERENTIAL
SAMP_residuals_DIFFERENTIAL_PCT | -3.72202856179
operation: evaluation((SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100)
inst: evaluation
factors: (SAMP_residuals_DIFFERENTIAL/residuals_MEAN)*100
SAMP_baseline_curvature_DIFFERENTIAL | -1.51595
operation: evaluation(SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN-baseline_curvature_MEAN
SAMP_STD_DELTA_18_16O_CURR_MEAN | -14.2007
varray: [-8.628, -8.686, -8.767, -8.812]
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
SAMP_STD_DELTA_DH_CURR_MEAN | -107.39
varray: [-64.622, -65.48, -66.068, -66.08]
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
known_dD | 0.0
operation: known_standard(d(D_H)Mean)
inst: known_standard
source_col: d(D_H)Mean
known_d18O | 0.0
operation: known_standard(d(18_16)Mean)
inst: known_standard
source_col: d(18_16)Mean
d18O_mean | -8.72325
varray: [-8.628, -8.686, -8.767, -8.812]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
dD_mean | -65.5625
varray: [-64.622, -65.48, -66.068, -66.08]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
d18O_stdev | 0.0821599050632
varray: [-8.628, -8.686, -8.767, -8.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
dD_stdev | 0.686702992567
varray: [-64.622, -65.48, -66.068, -66.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
measured_d18O_mean | -8.72325
varray: [-8.628, -8.686, -8.767, -8.812]
inst: mean_all
operation: mean_all(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
measured_dD_mean | -65.5625
varray: [-64.622, -65.48, -66.068, -66.08]
inst: mean_all
operation: mean_all(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
measured_d18O_stdev | 0.0821599050632
varray: [-8.628, -8.686, -8.767, -8.812]
inst: std_all
operation: std_all(d(18_16)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
measured_dD_stdev | 0.686702992567
varray: [-64.622, -65.48, -66.068, -66.08]
inst: std_all
operation: std_all(d(D_H)Mean)
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
SAMP_DELTA_18_16O_CALIBRATE | -10.1212489571
actual_array: [ 1.93 -15.84]
varray: [-8.628, -8.686, -8.767, -8.812]
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.72325
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(18_16)Mean
SAMP_DELTA_DH_CALIBRATE | -73.0430446487
actual_array: [ 18.1 -119. ]
varray: [-64.622, -65.48, -66.068, -66.08]
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.5625
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: [127043.0, 127579.0, 128114.0, 128649.0]
source_col: d(D_H)Mean
SAMP_TEST_VALUE | 39.3614
operation: evaluation(baseline_shift_MEAN-SAMP_residuals_MEAN)
inst: evaluation
factors: baseline_shift_MEAN-SAMP_residuals_MEAN
residuals_GT_3SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>residuals_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>residuals_PLUS_3SIGMA
residuals_GE_2SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_2SIGMA
residuals_GE_15SIGMA | False
operation: evaluation(SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA)
inst: evaluation
factors: SAMP_residuals_MEAN>=residuals_PLUS_15SIGMA
residuals_GT_10PCT | False
operation: evaluation(SAMP_residuals_DIFFERENTIAL>residuals_10PCT)
inst: evaluation
factors: SAMP_residuals_DIFFERENTIAL>residuals_10PCT
ch4_ppm_GE_3SIGMA | False
operation: evaluation(SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_ch4_ppm_MEAN_ABS>=ch4_ppm_PLUS_3SIGMA
baseline_shift_GT_1SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_1SIGMA
baseline_shift_GT_3SIGMA | True
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_3SIGMA
baseline_shift_GT_18SIGMA | False
operation: evaluation(SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA)
inst: evaluation
factors: SAMP_baseline_shift_MEAN_ABS>baseline_shift_PLUS_18SIGMA
slope_shift_GT_5SIGMA | False
operation: evaluation(SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA)
inst: evaluation
factors: SAMP_slope_shift_MEAN_ABS>slope_shift_PLUS_5SIGMA
baseline_curvature_GT_3SIGMA | False
operation: evaluation(SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA)
inst: evaluation
factors: SAMP_baseline_curvature_MEAN>baseline_curvature_PLUS_3SIGMA
RED | False
operation: truth_evaluation(((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and residuals_GE_15SIGMA) or (IS_NOT_STANDARD and baseline_shift_GT_18SIGMA) or IS_NOT_STANDARD and baseline_curvature_GT_3SIGMA))
NOT_RED | True
operation: bool_xor(RED,TRUE)
inst: bool_xor
factors: ('RED', ',', 'TRUE')
YELLOW | True
operation: truth_evaluation(((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
inst: truth_evaluation
factors: (((IS_NOT_STANDARD and NOT_RED) and (ch4_ppm_GE_3SIGMA or baseline_shift_GT_3SIGMA or slope_shift_GT_5SIGMA)))
NOT_YELLOW | False
operation: bool_xor(YELLOW,TRUE)
inst: bool_xor
factors: ('YELLOW', ',', 'TRUE')
GREEN | False
operation: truth_evaluation((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
inst: truth_evaluation
factors: ((IS_NOT_STANDARD and NOT_RED and NOT_YELLOW))
NOT_GREEN | True
operation: bool_xor(GREEN,TRUE)
inst: bool_xor
factors: ('GREEN', ',', 'TRUE')
CYAN | False
operation: truth_evaluation((IS_STANDARD))
inst: truth_evaluation
factors: ((IS_STANDARD))
NOT_CYAN | True
operation: bool_xor(CYAN,TRUE)
inst: bool_xor
factors: ('CYAN', ',', 'TRUE')
METHANE | False
operation: truth_evaluation(((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and ch4_ppm_GE_3SIGMA))
OTHER1 | True
operation: truth_evaluation(((RED or YELLOW) and baseline_shift_GT_1SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and baseline_shift_GT_1SIGMA))
OTHER2 | False
operation: truth_evaluation(((RED or YELLOW) and slope_shift_GT_5SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and slope_shift_GT_5SIGMA))
OTHER3 | False
operation: truth_evaluation(((RED or YELLOW) and residuals_GE_15SIGMA))
inst: truth_evaluation
factors: (((RED or YELLOW) and residuals_GE_15SIGMA))
OTHER4 | False
operation: truth_evaluation(RED or YELLOW and baseline_curvature_GT_3SIGMA)
inst: truth_evaluation
factors: (RED or YELLOW and baseline_curvature_GT_3SIGMA)
OTHER5 | True
operation: truth_evaluation(RED or YELLOW or GREEN or CYAN)
inst: truth_evaluation
factors: (RED or YELLOW or GREEN or CYAN)
baseline_shift_DIFF | 1.2221
operation: math_evaluation((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))
inst: math_evaluation
factors: ((SAMP_baseline_shift_MEAN-baseline_shift_MEAN))

## 1 - Sample 1 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 1 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.755
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.851
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -11.1984511933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.755
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1498338287
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.851
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.058
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.957
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4707458784
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.058
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.7726695352
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.957
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.42
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2400100469
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0839279088
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.42
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.857
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.358
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2608911176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.857
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0158069063
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.358
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.846
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.007
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2494065287
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.846
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6301541342
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.007
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.895
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.268
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.300565152
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.895
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9169215801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.268
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.746
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2118206014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3433866882
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.746
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.822
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.708
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2243492438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.822
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.301635106
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.708
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.837
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.86
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2400100469
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.837
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4686414347
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.86
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 1 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 0
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 0
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.784
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.756
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1846752095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.784
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3543739467
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.756
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 2 - Sample 2 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 2 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.133
operation: col(d(18_16)Mean)
inst: col
dD_actual | 7.583
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.21325724727
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.133
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 7.3238067685
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 7.583
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.673
operation: col(d(18_16)Mean)
inst: col
dD_actual | 15.076
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.77704615692
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.673
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 15.5565595376
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 15.076
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.829
operation: col(d(18_16)Mean)
inst: col
dD_actual | 17.295
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.9399185086
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.829
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 17.9946321913
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 17.295
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.861
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.004
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.97332822176
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.861
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 18.7736288166
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 18.004
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.884
operation: col(d(18_16)Mean)
inst: col
dD_actual | 18.459
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.9973414531
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.884
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 19.2735490768
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 18.459
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.925
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.137
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.04014764809
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.0184852009
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 19.137
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.953
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.25
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.06938114711
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.953
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.1426412215
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 19.25
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.94
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.371
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.05580845113
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.94
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.275587049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 19.371
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 3.011
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.783
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.12993625222
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 3.011
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.7282620978
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 19.783
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 2 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 0
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | 2.986
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | 19.951
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | 2.8195
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | 17.3909
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.986
operation: col(d(18_16)Mean)
inst: col
dD_actual | 19.951
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 2.10383491381
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.986
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | 20.91284804
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: 19.951
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 3 - Sample 3 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 3 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.221
operation: col(d(18_16)Mean)
inst: col
dD_actual | -93.149
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -14.8171407504
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -13.221
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -103.353045218
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -93.149
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -13.964
operation: col(d(18_16)Mean)
inst: col
dD_actual | -104.054
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.5928725279
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -13.964
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -115.334650576
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -104.054
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.202
operation: col(d(18_16)Mean)
inst: col
dD_actual | -107.311
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8413572696
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.202
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -118.913200658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -107.311
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.22
operation: col(d(18_16)Mean)
inst: col
dD_actual | -108.263
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.8601502333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.22
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -119.959187664
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -108.263
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.277
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.013
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -15.9196612848
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.277
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -120.783232049
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -109.013
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.394
operation: col(d(18_16)Mean)
inst: col
dD_actual | -109.935
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0418155486
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.394
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -121.79625728
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -109.935
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.404
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.138
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0522560839
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.404
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.019298627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -110.138
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.449
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.462
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0992384931
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.449
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.375285801
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -110.462
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.456
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.687
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.1065468678
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.456
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.622499116
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -110.687
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 3 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -14.42
operation: col(d(18_16)Mean)
inst: col
dD_actual | -110.888
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -16.0689609405
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -14.42
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -122.843343012
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -110.888
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 4 - Sample 4 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 4 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.351
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.894
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.7766535646
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.351
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.0983531935
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -72.894
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.085
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.011
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4989353239
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.085
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8320007309
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -69.011
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.023
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.995
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4342040047
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.023
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7156952707
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.995
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.944
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.433
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3517237753
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.944
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0982113448
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.433
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 5
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -9.004
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.396
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4143669875
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -9.004
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0575584885
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.396
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 6
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.992
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.375
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.401838345
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.992
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0344852457
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.375
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 7
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.997
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.196
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.4070586127
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.997
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8378133192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.196
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 8
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.96
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.127
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3684286319
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.96
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7620012358
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.127
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 9
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.975
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.245
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3840894349
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.975
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8916508857
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.245
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 4 - 10
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.942
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.994
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3496356682
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.942
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.6158706982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.994
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 5 - Sample 5 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 5 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.409
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.25
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.57288745138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.409
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2060635883
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.25
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 5 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.136
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.653
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.28786083595
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.136
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.5501242578
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.653
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 5 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.06
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.229
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.20851276718
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.06
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0842644988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.229
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 5 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.049
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.214
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.19702817828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.049
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0677836111
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.214
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 6 - Sample 6 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 6 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.718
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.055
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.9836070669
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.718
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.2020052748
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -52.055
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 6 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.891
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.381
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.16422832869
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.891
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.4614640542
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.381
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 6 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.925
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.431
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.19972614893
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.5164003465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.431
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 6 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.914
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.18824156003
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.914
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.3449991144
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 7 - Sample 7 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 7 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.55
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.447
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.89631314556
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.55
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.5212384267
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -61.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 7 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.643
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.229
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.99341012444
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.643
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3804420388
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.229
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 7 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.621
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.591
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.97044094664
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.621
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7781807953
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.591
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 7 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.629
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.405
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.97879337493
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.629
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.5738177878
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.405
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 8 - Sample 8 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 8 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.804
operation: col(d(18_16)Mean)
inst: col
dD_actual | -72.461
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.1615027438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.804
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.6226049019
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -72.461
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 8 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.786
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.45
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.14270978014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.786
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.7092447642
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.45
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 8 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.814
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.663
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.17194327916
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.814
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.9432733696
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.663
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 8 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.827
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.865
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.18551597514
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.827
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1652159906
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.865
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 9 - Sample 9 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 9 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.792
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.431
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0168134922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.792
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.2934655865
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -69.431
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 9 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.716
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.178
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.93746542344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.716
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -77.0154879473
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -69.178
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 9 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.708
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.811
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.92911299515
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.708
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6122555615
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.811
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 9 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -69.065
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.91554029917
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.8913319266
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -69.065
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 10 - Sample 10 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 10 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.487
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.168
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.78648423638
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.487
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7083231488
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.168
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 10 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.448
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.487
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.74576614846
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.448
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9600908472
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.487
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 10 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.455
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.264
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.75307452321
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.455
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.7150749834
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.264
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 10 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.225
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.80214503942
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.6722246754
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.225
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 11 - Sample 11 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 11 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.424
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.081
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6766553272
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.424
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6127340002
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.081
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 11 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.373
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.077
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62340859684
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.373
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6083390968
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.077
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 11 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.37
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.162
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62027643623
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.37
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.7017307937
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.162
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 11 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.358
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.949
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.6077477938
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.358
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4677021884
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.949
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 12 - Sample 12 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 12 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.142
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.83
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.42628576632
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.142
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.349695346
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -55.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 12 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.166
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.896
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.4513430512
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.166
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -61.3234854052
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -54.896
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 12 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.152
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.408
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.43672630169
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.152
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.787307192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -54.408
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 12 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.117
operation: col(d(18_16)Mean)
inst: col
dD_actual | -54.053
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.40018442792
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.117
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -60.3972595165
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -54.053
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 13 - Sample 13 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 13 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.735
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.085
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1335165862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.735
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.5184030569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.085
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 13 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.853
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.291
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2567149035
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.853
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.843466428
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.291
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 13 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.911
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.534
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3172700086
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.911
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1104568087
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.534
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 13 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.925
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.588
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3318867581
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1697880044
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.588
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 14 - Sample 14 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 14 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.825
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.274
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0512672589
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.825
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7515453086
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -45.274
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 14 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.569
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.04
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.78398955359
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.569
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.2969917672
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -43.04
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 14 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.504
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.213
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.71612607372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.504
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.388345492
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -42.213
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 14 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.525
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.173
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.73805119799
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.525
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.3443964581
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -42.173
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 15 - Sample 15 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 15 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.415
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.61
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79941945453
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.415
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.0952341264
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.61
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 15 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.61
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.034
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0030098941
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.61
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.7585455787
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.034
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 15 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.577
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.558
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.96855612743
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.577
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.3342779223
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.558
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 15 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.832
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0238909649
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.6353288043
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.832
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 16 - Sample 16 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 16 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.885
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.094
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.0698569347
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.885
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.1333882028
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -61.094
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 16 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.65
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.82450435365
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.65
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.473053969
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 16 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.558
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.415
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.7284514283
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.558
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.387353353
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.415
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 16 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.549
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.184
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.71905494648
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.549
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.1335476824
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.184
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 17 - Sample 17 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 17 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.353
operation: col(d(18_16)Mean)
inst: col
dD_actual | -85.58
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -12.8668487444
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -11.353
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -95.0367892843
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -85.58
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 17 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.759
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.369
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.2907344802
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -11.759
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.1011356706
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -88.369
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 17 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.753
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.666
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.284470159
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -11.753
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.4274572471
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -88.666
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 17 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -11.829
operation: col(d(18_16)Mean)
inst: col
dD_actual | -88.918
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -13.3638182278
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -11.829
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -98.7043361604
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -88.918
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 18 - Sample 18 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 18 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.047
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.222
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.0794151067555
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.047
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.1880407979
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -50.222
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 18 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.694
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.046
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.754917744797
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.694
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -51.5997616622
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -46.046
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 18 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.807
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.872895794409
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.807
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7823096323
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -45.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 18 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.87
operation: col(d(18_16)Mean)
inst: col
dD_actual | -44.877
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.938671167201
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.87
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.3153511475
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -44.877
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 19 - Sample 19 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 19 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.084
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.23357005206
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.084
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.6939507569
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -41.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 19 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.257
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.662
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.41419131385
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.257
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.6842217038
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -40.662
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 19 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.304
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.495
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.46326183006
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.304
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.5007344874
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -40.495
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 19 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.342
operation: col(d(18_16)Mean)
inst: col
dD_actual | -40.631
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.50293586444
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.342
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -45.6501612026
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -40.631
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 20 - Sample 20 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 20 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.702
operation: col(d(18_16)Mean)
inst: col
dD_actual | -61.642
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.92284867393
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.702
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -68.7354899668
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -61.642
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 20 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.839
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.869
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.06588400841
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.839
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1823524273
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 20 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.847
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.506
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0742364367
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.847
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.8822407917
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.506
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 20 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.805
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.411
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.03038618818
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.805
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7778618362
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.411
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 21 - Sample 21 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 21 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.53
operation: col(d(18_16)Mean)
inst: col
dD_actual | -76.829
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.91948561121
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.53
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -85.4218394001
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -76.829
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 21 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.004
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0896663376
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.7128422699
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -78.004
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 21 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.676
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.249
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0719174275
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.676
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -86.9820301024
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -78.249
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 21 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.731
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.129340372
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.731
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.1830969323
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -78.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 22 - Sample 22 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 22 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.922
operation: col(d(18_16)Mean)
inst: col
dD_actual | -68.275
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3287545975
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.922
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -76.0233385077
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -68.275
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 22 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.925
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.216
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3318867581
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.925
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.8597878361
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.216
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 22 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.9
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.736
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3057854197
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.9
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3323994297
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.736
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 22 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.959
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.866
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3673845783
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.959
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4752337898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.866
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 23 - Sample 23 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 23 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.132
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.199
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.50395230373
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.132
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.2360128834
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.199
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 23 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.036
operation: col(d(18_16)Mean)
inst: col
dD_actual | -70.941
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.40372316424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.036
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -78.9525416149
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -70.941
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 23 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.018
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.448
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.38493020059
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.018
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.5095956192
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.448
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 23 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.046
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.086
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.4141636996
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.046
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1118568627
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.086
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 24 - Sample 24 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 24 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.724
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.298
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.85771077896
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.724
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2588024289
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.298
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 24 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.502
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.312
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.62593089388
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.502
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1754587441
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.312
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 24 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.39
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.5089968978
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.39
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.1106339191
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 24 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.394
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.014
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.51317311195
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.394
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.8480384418
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.014
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 25 - Sample 25 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 25 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.156
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.753
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.39684897945
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.156
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.0549002291
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.753
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 25 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.315
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.18
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.56285349173
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.315
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5240561656
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.18
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 25 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.235
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.163
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.47932920882
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.235
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.5053778263
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.163
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 25 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.279
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.361
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.52526756442
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.279
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.7229255439
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.361
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 26 - Sample 26 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 26 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.562
operation: col(d(18_16)Mean)
inst: col
dD_actual | -78.302
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.95289532438
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.562
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.0402625722
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -78.302
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 26 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.63
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.146
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0238909649
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.63
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -87.9675871868
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -79.146
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 26 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.698
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.573
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0948866053
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.698
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.4367431233
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -79.573
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 26 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -79.503
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0740055346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -88.3598323141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -79.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 27 - Sample 27 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 27 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.769
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.815
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.99280026087
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.769
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.1230212316
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.815
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 27 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.602
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.614
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.81844332029
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.602
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8034514898
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.614
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 27 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.589
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.252
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.80487062432
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.589
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4057127333
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.252
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 27 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.58
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.157
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.79547414249
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.58
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3013337778
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.157
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 28 - Sample 28 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 28 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.054
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.839
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.29035551874
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.054
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.0506648053
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.839
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 28 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.98
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.637
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.21309555704
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.98
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.8287221842
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.637
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 28 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.048
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.531
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.28409119752
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.048
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7122572445
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.531
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 28 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.999
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.309
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.23293257424
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.999
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.4683401065
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.309
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 29 - Sample 29 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 29 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.559
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.1
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.94976316377
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.559
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.1272390246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.1
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 29 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.678
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.573
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0740055346
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.678
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.64693635
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.573
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 29 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.707
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.955
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1042830872
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.707
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -80.0666496235
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.955
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 29 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.702
operation: col(d(18_16)Mean)
inst: col
dD_actual | -71.825
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0990628195
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.702
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -79.9238152634
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -71.825
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 30 - Sample 30 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 30 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.627
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.02705073971
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3241056925
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -57.627
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 30 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.713
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.447
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.89027972644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.713
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.0276091934
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.447
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 30 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.644
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.246
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.81824003243
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.644
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.8067652982
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.246
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 30 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.609
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.952
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.78169815866
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.609
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4837398993
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -55.952
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 31 - Sample 31 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 31 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.666
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.926
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0614768922
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.666
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4424314939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.926
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 31 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.848
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.072
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2514946358
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.848
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6028454675
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.072
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 31 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.924
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.3308427046
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.924
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3082274611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 31 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.385
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.257758957
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.9467466575
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.385
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 32 - Sample 32 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 32 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.714
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.297
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.06753792552
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.714
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.5411397097
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.297
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 32 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.62
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.743
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.9693968931
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.62
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.0311714373
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.743
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 32 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.638
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.997
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.98818985676
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.638
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.3102478023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.997
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 32 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.595
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.898
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.94329555469
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.595
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2014739435
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.898
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 33 - Sample 33 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 33 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.95
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.862
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.18177395095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.95
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -55.7924994931
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -49.862
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 33 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.854
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.844
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.08154481146
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.854
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.5752707345
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.844
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 33 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.775
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.09
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.99906458208
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.775
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.7468314462
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.09
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 33 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.758
operation: col(d(18_16)Mean)
inst: col
dD_actual | -46.802
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.98131567197
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.758
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -52.4303984023
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -46.802
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 34 - Sample 34 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 34 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -0.222
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.24548883092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -0.222
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.1525626117
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -52.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 34 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.002
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -1.01162083877
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 0.002
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.8273397611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 34 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.017
operation: col(d(18_16)Mean)
inst: col
dD_actual | -52.215
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.995960035722
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 0.017
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.3778014103
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -52.215
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 34 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 0.09
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.981
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -0.919744127566
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 0.09
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -58.1206995622
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.981
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 35 - Sample 35 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 35 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.994
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.314
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.35987291571
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.994
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8687371224
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.314
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 35 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.327
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.141
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.70754274333
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.327
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.7773833976
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.141
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 35 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.347
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.432
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.72842381406
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.347
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.097112619
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.432
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 35 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.376
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.437
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.75870136661
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.376
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.1026062482
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.437
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 36 - Sample 36 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 36 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.897
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.253
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.03833204075
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.897
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.2093597658
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.253
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 36 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.111
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.73346840813
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.9546148489
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.111
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 36 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.611
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.229
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.73973272935
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.611
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -66.0842644988
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.229
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 36 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.596
operation: col(d(18_16)Mean)
inst: col
dD_actual | -59.064
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.7240719263
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.596
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -65.9029747341
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -59.064
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 37 - Sample 37 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 37 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.71
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.196
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.77162260138
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.71
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.1594739259
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -50.196
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 37 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.006
operation: col(d(18_16)Mean)
inst: col
dD_actual | -49.055
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.08066244815
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.006
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.9058277349
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -49.055
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 37 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.012
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.95
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.08692676937
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.012
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.790461521
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -48.95
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 37 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 2.068
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.549
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 1.14539376741
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 2.068
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.3498724564
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -48.549
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 38 - Sample 38 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 38 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.359
operation: col(d(18_16)Mean)
inst: col
dD_actual | -60.585
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.56473831095
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.359
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -67.5741367469
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -60.585
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 38 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.858
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.235
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.0857210256
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.858
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.3870343939
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.235
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 38 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.932
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.581
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.1629809873
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.932
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -69.7671935368
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.581
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 38 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.015
operation: col(d(18_16)Mean)
inst: col
dD_actual | -62.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.24963743082
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.015
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.1847093586
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -62.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 39 - Sample 39 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 39 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.888
operation: col(d(18_16)Mean)
inst: col
dD_actual | -45.261
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.02893555892
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.888
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -50.7372618726
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -45.261
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 39 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.655
operation: col(d(18_16)Mean)
inst: col
dD_actual | -43.142
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.78567108495
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.655
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -48.4090618035
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -43.142
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 39 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.668
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.757
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.79924378092
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.668
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.9860523526
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -42.757
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 39 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -2.591
operation: col(d(18_16)Mean)
inst: col
dD_actual | -42.407
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -3.71885165862
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -2.591
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -47.6014983062
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -42.407
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 40 - Sample 40 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 40 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.458
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.493
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.8443137566
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.458
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -70.769231509
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.493
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 40 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.764
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.418
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1637941387
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.764
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.8842787638
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.418
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 40 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.81
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.924
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2118206014
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.81
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.4402340422
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.924
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 40 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.844
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.01
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2473184216
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.844
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.534724465
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.01
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 41 - Sample 41 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 41 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.685
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.83
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.90509976381
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.685
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.534405506
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 41 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.503
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.681
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.71508202019
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.503
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3706953548
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.681
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 41 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.482
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.576
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.69315689592
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.482
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2553291409
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.576
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 41 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -4.431
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.585
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -5.63991016557
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -4.431
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2652176735
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.585
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 42 - Sample 42 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 42 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.069
operation: col(d(18_16)Mean)
inst: col
dD_actual | -41.145
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.2785984301
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 5.069
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -46.2149062877
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -41.145
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 42 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.463
operation: col(d(18_16)Mean)
inst: col
dD_actual | -38.515
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.68995552344
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 5.463
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -43.325257311
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -38.515
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 42 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.565
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.897
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.79644898415
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 5.565
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.6462447378
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -37.897
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 42 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 5.638
operation: col(d(18_16)Mean)
inst: col
dD_actual | -37.541
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 4.8726648923
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 5.638
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -42.2550983364
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -37.541
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 43 - Sample 43 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 43 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.693
operation: col(d(18_16)Mean)
inst: col
dD_actual | -63.942
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.04561280126
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.693
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -71.2625594141
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -63.942
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 43 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.252
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.324
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.6292387281
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.252
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.8797243809
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.324
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 43 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.414
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.937
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.79837540099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.414
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.5532433249
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.937
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 43 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.398
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.719
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.78167054441
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.398
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3137210903
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.719
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 44 - Sample 44 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 44 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.683
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.614
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.85895812035
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.683
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -64.3098222565
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -57.614
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 44 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.333
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.503
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.49353938262
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.333
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.0891378408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.503
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 44 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.247
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.961
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.40375077849
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.247
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.4936284319
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -55.961
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 44 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.215
operation: col(d(18_16)Mean)
inst: col
dD_actual | -55.988
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.37034106532
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.215
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.5232940298
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -55.988
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 45 - Sample 45 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 45 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.454
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.185908567467
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.245493501
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -48.454
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 45 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.335
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.705
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.380102525235
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.335
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.4225478419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.705
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 45 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.432
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.438
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.481375718264
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.432
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.1291880408
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.438
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 45 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | 1.465
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.428
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | 0.515829484965
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: 1.465
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.1182007823
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.428
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 46 - Sample 46 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 46 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.478
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.1
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.73303421816
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.478
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -62.6463513246
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.1
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 46 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.785
operation: col(d(18_16)Mean)
inst: col
dD_actual | -56.863
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.05355865383
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.785
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.4846791456
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -56.863
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 46 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.836
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.09
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.10680538419
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.836
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.7340899128
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -57.09
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 46 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.904
operation: col(d(18_16)Mean)
inst: col
dD_actual | -57.088
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.17780102466
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.904
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -63.7318924611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -57.088
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 47 - Sample 47 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 47 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.477
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.688
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.86415077379
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.477
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.1809347424
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.688
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 47 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.585
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.606
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.97690855572
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.585
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.1895650697
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.606
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 47 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.605
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.665
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99778962644
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.605
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.2543898946
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.665
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 47 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.606
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.869
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -9.99883367998
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.606
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4785299673
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.869
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 48 - Sample 48 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 48 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.273
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.278
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.56305677959
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.273
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.9279088386
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.278
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 48 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.125
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.429
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.40853685621
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.125
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.0938164415
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.429
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 48 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.149
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.539
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.43359414108
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.149
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.2146762846
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.539
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 48 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -6.129
operation: col(d(18_16)Mean)
inst: col
dD_actual | -67.65
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -7.41271307035
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -6.129
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -75.3366348536
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -67.65
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 49 - Sample 49 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 49 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.77
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.871
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1700584599
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.77
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.480727419
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.871
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 49 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.836
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.727
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2389659933
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.836
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3225108971
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.727
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 49 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.893
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.816
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2984770449
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.893
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.4202974975
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.816
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 49 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.894
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.714
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2995210985
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.894
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -74.3082274611
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.714
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 50 - Sample 50 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 50 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.262
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.104
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.59562572708
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.262
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.3290856213
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.104
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 50 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.155
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.558
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.48391199868
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.155
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -81.8279071557
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.558
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 50 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.182
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.828
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.51210144417
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.182
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.1245631343
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.828
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 50 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -7.189
operation: col(d(18_16)Mean)
inst: col
dD_actual | -73.971
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -8.51940981892
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -7.189
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -82.2816809303
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -73.971
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 51 - Sample 51 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 51 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.695
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.049
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.87148676279
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.695
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0966870731
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.049
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 51 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.508
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.887
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.67624875148
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.508
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -54.7212417926
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -48.887
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 51 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.451
operation: col(d(18_16)Mean)
inst: col
dD_actual | -48.211
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.61673769991
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.451
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.9785031203
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -48.211
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 51 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -3.425
operation: col(d(18_16)Mean)
inst: col
dD_actual | -47.954
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -4.58959230796
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -3.425
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -53.6961305777
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -47.954
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 52 - Sample 52 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 52 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.377
operation: col(d(18_16)Mean)
inst: col
dD_actual | -50.83
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.62758481099
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.377
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -56.8560661127
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -50.83
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 52 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.468
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.212
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.7225936828
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.468
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.2757793861
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.212
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 52 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.492
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.14
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74765096767
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.492
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.1966711251
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.14
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 52 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -5.494
operation: col(d(18_16)Mean)
inst: col
dD_actual | -51.012
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -6.74973907475
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -5.494
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -57.0560342168
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -51.012
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean

## 53 - Sample 53 Detail Vars
ChemCorrect(tm) - Tue, May 20, 2025, 09:41:39
Detail Variables
Detail 53 - 1
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.628
operation: col(d(18_16)Mean)
inst: col
dD_actual | -64.622
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0218028578
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.628
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.0096929899
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -64.622
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 53 - 2
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.686
operation: col(d(18_16)Mean)
inst: col
dD_actual | -65.48
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.0823579629
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.686
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -72.9523997663
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -65.48
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 53 - 3
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.767
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.068
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.1669262993
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.767
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.5984505641
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.068
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean
Detail 53 - 4
Variable | Value | Operations
DTL_TEST | 3
operation: math_evaluation((1 + 2))
inst: math_evaluation
factors: ((1 + 2))
DTL_RED | True
operation: bool_and(TRUE,TRUE)
inst: bool_and
factors: ('TRUE', ',', 'TRUE')
previous_standard_d18O_mean | 2.8195
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
previous_standard_dD_mean | 17.3909
varray: []
inst: prior_standards_mean
operation: prior_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_endval | -14.42
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_endval | -110.888
varray: []
inst: current_standards_endval
operation: current_standards_endval(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
current_standard_d18O_mean | -14.2007
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(18_16)Mean)
tarray: []
source_col: d(18_16)Mean
current_standard_dD_mean | -107.39
varray: []
inst: current_standards_mean
operation: current_standards_mean(d(D_H)Mean)
tarray: []
source_col: d(D_H)Mean
d18O_actual | -8.812
operation: col(d(18_16)Mean)
inst: col
dD_actual | -66.08
operation: col(d(D_H)Mean)
inst: col
DTL_DELTA_18_16O_CALIBRATE | -10.2139087085
actual_array: [ 1.93 -15.84]
varray: []
fit_rsq: 1.0
fit_intercept: 0.970935790658
measure_array: [ 2.8195 -14.2007]
inst: fit_type1
actual_mean: -8.812
operation: fit_type1(d(18_16)Mean)
fit_slope: 0.957805289814
tarray: []
source_col: d(18_16)Mean
DTL_DELTA_DH_CALIBRATE | -73.6116352743
actual_array: [ 18.1 -119. ]
varray: []
fit_rsq: 1.0
fit_intercept: 0.917272793581
measure_array: [ 17.3909 -107.39 ]
inst: fit_type1
actual_mean: -66.08
operation: fit_type1(d(D_H)Mean)
fit_slope: 0.910145149526
tarray: []
source_col: d(D_H)Mean